## README
CTDI k-factors are used by RTI Piranha and RTI Ocean software to calculate CTDI(vol) from measurements of CTDI100(center).
k = CTDIw/CTDI100(center)
k-factors for some common scanners are built into RTI Ocean, but for newer scanners that aren't on the list, need to either select a similar scanner or use an appropriate k-factor.
See RTI Application Note AN-40
Spreadsheet tabs contain k-factors derived from Siemens Dosimetry and Imaging performance reports for the scanner.

## Sheet9
Scanner | kV | Coll NxT | Width (mm) | Acq NxT | k (16 cm) | k (32 cm) | Naeotom Alpha | 1 | Siemens | 1
Confidence 64 | 70 | 2x1 | 2 | 2x1 | 1.0699 | 1.7414 | Definition AS | 2 | Neurologica | 2
Confidence 64 | 70 | 1x5 | 5 | 1x5 | 1.0661 | 1.75 | Somatom Flash A tube | 3
Confidence 64 | 70 | 10x0.6 | 6 | 20x0.6 | 1.0645 | 1.7246 | Somatom Force | 4
Confidence 64 | 70 | 6x1.2 | 7.2 | 6x1.2 | 1.0678 | 1.7333 | Somatom Drive | 5
Confidence 64 | 70 | 1x10 | 10 | 1x10 | 1.0661 | 1.75 | Definition Edge | 6
Confidence 64 | 70 | 20x0.6 | 12 | 40x0.6 | 1.0632 | 1.7429 | go.Top | 7
Confidence 64 | 70 | 12x1.2 | 14.4 | 12x1.2 | 1.0699 | 1.7313 | Confidence 64 | 8
Confidence 64 | 70 | 32x0.6 | 19.2 | 64x0.6 | 1.0617 | 1.7121 | Confidence 20 | 9
Confidence 64 | 70 | 16x1.2 | 19.2 | 16x1.2 | 1.0698 | 1.7344 | X.cite | 10
Confidence 64 | 80 | 2x1 | 2 | 2x1 | 1.0438 | 1.6117 | Biograph Vision | 11
Confidence 64 | 80 | 1x5 | 5 | 1x5 | 1.0439 | 1.6196 | Intevo Bold | 12
Confidence 64 | 80 | 10x0.6 | 6 | 20x0.6 | 1.0423 | 1.6066 | Symbia Pro.specta | 13
Confidence 64 | 80 | 6x1.2 | 7.2 | 6x1.2 | 1.0458 | 1.609 | Pro.Pulse A tube | 14
Confidence 64 | 80 | 1x10 | 10 | 1x10 | 1.0439 | 1.6196 | Pro.Pulse B tube | 15
Confidence 64 | 80 | 20x0.6 | 12 | 40x0.6 | 1.0414 | 1.6129 | BodyTom | 16
Confidence 64 | 80 | 12x1.2 | 14.4 | 12x1.2 | 1.0457 | 1.6186 | Somatom Flash B tube | 17
Confidence 64 | 80 | 32x0.6 | 19.2 | 64x0.6 | 1.0406 | 1.6034
Confidence 64 | 80 | 16x1.2 | 19.2 | 16x1.2 | 1.0457 | 1.6195
Confidence 64 | 100 | 2x1 | 2 | 2x1 | 1.0204 | 1.4889
Confidence 64 | 100 | 1x5 | 5 | 1x5 | 1.0214 | 1.4901
Confidence 64 | 100 | 10x0.6 | 6 | 20x0.6 | 1.0161 | 1.4794
Confidence 64 | 100 | 6x1.2 | 7.2 | 6x1.2 | 1.0211 | 1.4897
Confidence 64 | 100 | 1x10 | 10 | 1x10 | 1.0214 | 1.4901
Confidence 64 | 100 | 20x0.6 | 12 | 40x0.6 | 1.0168 | 1.4816
Confidence 64 | 100 | 12x1.2 | 14.4 | 12x1.2 | 1.0216 | 1.4903
Confidence 64 | 100 | 32x0.6 | 19.2 | 64x0.6 | 1.016 | 1.4764
Confidence 64 | 100 | 16x1.2 | 19.2 | 16x1.2 | 1.0216 | 1.4939
Confidence 64 | 120 | 2x1 | 2 | 2x1 | 1.0121 | 1.4328
Confidence 64 | 120 | 1x5 | 5 | 1x5 | 1.012 | 1.4332
Confidence 64 | 120 | 10x0.6 | 6 | 20x0.6 | 1.0085 | 1.4148
Confidence 64 | 120 | 6x1.2 | 7.2 | 6x1.2 | 1.0115 | 1.4318
Confidence 64 | 120 | 1x10 | 10 | 1x10 | 1.012 | 1.4332
Confidence 64 | 120 | 20x0.6 | 12 | 40x0.6 | 1.0083 | 1.4153
Confidence 64 | 120 | 12x1.2 | 14.4 | 12x1.2 | 1.0117 | 1.4331
Confidence 64 | 120 | 32x0.6 | 19.2 | 64x0.6 | 1.0083 | 1.4156
Confidence 64 | 120 | 16x1.2 | 19.2 | 16x1.2 | 1.0116 | 1.4324
Confidence 64 | 140 | 2x1 | 2 | 2x1 | 1.0075 | 1.3961
Confidence 64 | 140 | 1x5 | 5 | 1x5 | 1.0074 | 1.3967
Confidence 64 | 140 | 10x0.6 | 6 | 20x0.6 | 1.0048 | 1.3882
Confidence 64 | 140 | 6x1.2 | 7.2 | 6x1.2 | 1.0073 | 1.3967
Confidence 64 | 140 | 1x10 | 10 | 1x10 | 1.0074 | 1.3967
Confidence 64 | 140 | 20x0.6 | 12 | 40x0.6 | 1.0043 | 1.389
Confidence 64 | 140 | 12x1.2 | 14.4 | 12x1.2 | 1.0074 | 1.3949
Confidence 64 | 140 | 32x0.6 | 19.2 | 64x0.6 | 1.0046 | 1.3887
Confidence 64 | 140 | 16x1.2 | 19.2 | 16x1.2 | 1.0073 | 1.3968
Confidence 20 | 70 | 2x1 | 2 | 2x1 | 1.0699 | 2.0147
Confidence 20 | 70 | 1x5 | 5 | 1x5 | 1.0661 | 2
Confidence 20 | 70 | 10x0.6 | 6 | 10x0.6 | 1.0706 | 2.0119
Confidence 20 | 70 | 6x1.2 | 7.2 | 6x1.2 | 1.0678 | 2
Confidence 20 | 70 | 1x10 | 10 | 1x10 | 1.0661 | 2
Confidence 20 | 70 | 20x0.6 | 12 | 20x0.6 | 1.0661 | 2
Confidence 20 | 70 | 12x1.2 | 14.4 | 12x1.2 | 1.0699 | 1.9873
Confidence 20 | 70 | 16x1.2 | 19.2 | 16x1.2 | 1.0698 | 2.12
Confidence 20 | 80 | 2x1 | 2 | 2x1 | 1.0438 | 1.8475
Confidence 20 | 80 | 1x5 | 5 | 1x5 | 1.0439 | 1.8491
Confidence 20 | 80 | 10x0.6 | 6 | 10x0.6 | 1.046 | 1.8493
Confidence 20 | 80 | 6x1.2 | 7.2 | 6x1.2 | 1.0458 | 1.8487
Confidence 20 | 80 | 1x10 | 10 | 1x10 | 1.0439 | 1.8491
Confidence 20 | 80 | 20x0.6 | 12 | 20x0.6 | 1.0469 | 1.8531
Confidence 20 | 80 | 12x1.2 | 14.4 | 12x1.2 | 1.0457 | 1.8456
Confidence 20 | 80 | 16x1.2 | 19.2 | 16x1.2 | 1.0457 | 1.8462
Confidence 20 | 100 | 2x1 | 2 | 2x1 | 1.0204 | 1.6719
Confidence 20 | 100 | 1x5 | 5 | 1x5 | 1.0214 | 1.6769
Confidence 20 | 100 | 10x0.6 | 6 | 10x0.6 | 1.021 | 1.6719
Confidence 20 | 100 | 6x1.2 | 7.2 | 6x1.2 | 1.0211 | 1.6727
Confidence 20 | 100 | 1x10 | 10 | 1x10 | 1.0214 | 1.6769
Confidence 20 | 100 | 20x0.6 | 12 | 20x0.6 | 1.0214 | 1.6742
Confidence 20 | 100 | 12x1.2 | 14.4 | 12x1.2 | 1.0216 | 1.6735
Confidence 20 | 100 | 16x1.2 | 19.2 | 16x1.2 | 1.0216 | 1.6702
Confidence 20 | 120 | 2x1 | 2 | 2x1 | 1.0121 | 1.59
Confidence 20 | 120 | 1x5 | 5 | 1x5 | 1.012 | 1.5894
Confidence 20 | 120 | 10x0.6 | 6 | 10x0.6 | 1.0119 | 1.5902
Confidence 20 | 120 | 6x1.2 | 7.2 | 6x1.2 | 1.0115 | 1.5899
Confidence 20 | 120 | 1x10 | 10 | 1x10 | 1.012 | 1.5894
Confidence 20 | 120 | 20x0.6 | 12 | 20x0.6 | 1.0122 | 1.5911
Confidence 20 | 120 | 12x1.2 | 14.4 | 12x1.2 | 1.0117 | 1.5895
Confidence 20 | 120 | 16x1.2 | 19.2 | 16x1.2 | 1.0116 | 1.5906
Confidence 20 | 140 | 2x1 | 2 | 2x1 | 1.0075 | 1.5392
Confidence 20 | 140 | 1x5 | 5 | 1x5 | 1.0074 | 1.5397
Confidence 20 | 140 | 10x0.6 | 6 | 10x0.6 | 1.0076 | 1.5387
Confidence 20 | 140 | 6x1.2 | 7.2 | 6x1.2 | 1.0073 | 1.54
Confidence 20 | 140 | 1x10 | 10 | 1x10 | 1.0074 | 1.5397
Confidence 20 | 140 | 20x0.6 | 12 | 20x0.6 | 1.0074 | 1.5383
Confidence 20 | 140 | 12x1.2 | 14.4 | 12x1.2 | 1.0074 | 1.5391
Confidence 20 | 140 | 16x1.2 | 19.2 | 16x1.2 | 1.0073 | 1.5383
go.Top | 70 | 6x0.6 | 3.6 | 6x0.6
go.Top | 70 | 1x5 | 5 | 1x5 | 1.0811 | 1.8889
go.Top | 70 | 1x10 | 10 | 1x10 | 1.0541 | 2.125
go.Top | 70 | 3x5 | 15 | 3x5 | 1.0811 | 2.125
go.Top | 70 | 64x0.6 | 38.4 | 64x0.6 | 1.0789 | 1.8889
go.Top | 80 | 6x0.6 | 3.6 | 6x0.6
go.Top | 80 | 1x5 | 5 | 1x5 | 1.0526 | 1.9286
go.Top | 80 | 1x10 | 10 | 1x10 | 1.0526 | 1.8571
go.Top | 80 | 3x5 | 15 | 3x5 | 1.0526 | 1.8571
go.Top | 80 | 64x0.6 | 38.4 | 64x0.6 | 1.0508 | 1.8
go.Top | 90 | 6x0.6 | 3.6 | 6x0.6
go.Top | 90 | 1x5 | 5 | 1x5 | 1.0366 | 1.7727
go.Top | 90 | 1x10 | 10 | 1x10 | 1.037 | 1.7273
go.Top | 90 | 3x5 | 15 | 3x5 | 1.0366 | 1.7273
go.Top | 90 | 64x0.6 | 38.4 | 64x0.6 | 1.0476 | 1.7391
go.Top | 100 | 6x0.6 | 3.6 | 6x0.6 | 1 | 1.4
go.Top | 100 | 1x5 | 5 | 1x5 | 1.036 | 1.6875
go.Top | 100 | 1x10 | 10 | 1x10 | 1.0273 | 1.6774
go.Top | 100 | 3x5 | 15 | 3x5 | 1.0364 | 1.7097
go.Top | 100 | 64x0.6 | 38.4 | 64x0.6 | 1.0351 | 1.6875
go.Top | 110 | 6x0.6 | 3.6 | 6x0.6
go.Top | 110 | 1x5 | 5 | 1x5 | 1.0352 | 1.6667
go.Top | 110 | 1x10 | 10 | 1x10 | 1.0284 | 1.6429
go.Top | 110 | 3x5 | 15 | 3x5 | 1.0282 | 1.6429
go.Top | 110 | 64x0.6 | 38.4 | 64x0.6 | 1.0342 | 1.6512
go.Top | 120 | 6x0.6 | 3.6 | 6x0.6
go.Top | 120 | 1x5 | 5 | 1x5 | 1.0225 | 1.6481
go.Top | 120 | 1x10 | 10 | 1x10 | 1.0227 | 1.6226
go.Top | 120 | 3x5 | 15 | 3x5 | 1.0226 | 1.6111
go.Top | 120 | 64x0.6 | 38.4 | 64x0.6 | 1.0273 | 1.6071
go.Top | 130 | 6x0.6 | 3.6 | 6x0.6
go.Top | 130 | 1x5 | 5 | 1x5 | 1.0282 | 1.6119
go.Top | 130 | 1x10 | 10 | 1x10 | 1.0238 | 1.5909
go.Top | 130 | 3x5 | 15 | 3x5 | 1.0236 | 1.5909
go.Top | 130 | 64x0.6 | 38.4 | 64x0.6 | 1.0228 | 1.6029
go.Top | 140 | 6x0.6 | 3.6 | 6x0.6 | 1.0156 | 1.48
go.Top | 140 | 1x5 | 5 | 1x5 | 1.024 | 1.5926
go.Top | 140 | 1x10 | 10 | 1x10 | 1.0243 | 1.5823
go.Top | 140 | 3x5 | 15 | 3x5 | 1.0201 | 1.5625
go.Top | 140 | 64x0.6 | 38.4 | 64x0.6 | 1.0194 | 1.5663
Intevo Bold | 80 | 4x0.6 | 2.4 | 4x0.6 | 1.0684 | 1.8564
Intevo Bold | 80 | 12x0.6 | 7.2 | 12x0.6 | 1.0649 | 1.8632
Intevo Bold | 80 | 16x0.6 | 9.6 | 16x0.6 | 1.0828 | 1.8571
Intevo Bold | 80 | 1x10 | 10 | 1x10 | 1.0631 | 1.8613
Intevo Bold | 80 | 12x1.2 | 14.4 | 12x1.2 | 1.0679 | 1.8533
Intevo Bold | 80 | 2x8 | 16 | 2x8 | 1.07 | 1.8519
Intevo Bold | 80 | 16x1.2 | 19.2 | 16x1.2 | 1.0701 | 1.9079
Intevo Bold | 110 | 4x0.6 | 2.4 | 4x0.6 | 1.0361 | 1.6469
Intevo Bold | 110 | 12x0.6 | 7.2 | 12x0.6 | 1.0386 | 1.6602
Intevo Bold | 110 | 16x0.6 | 9.6 | 16x0.6 | 1.0392 | 1.6615
Intevo Bold | 110 | 1x10 | 10 | 1x10 | 1.0386 | 1.663
Intevo Bold | 110 | 12x1.2 | 14.4 | 12x1.2 | 1.0396 | 1.6519
Intevo Bold | 110 | 2x8 | 16 | 2x8 | 1.04 | 1.6538
Intevo Bold | 110 | 16x1.2 | 19.2 | 16x1.2 | 1.0396 | 1.6553
Intevo Bold | 130 | 4x0.6 | 2.4 | 4x0.6 | 1.0287 | 1.5887
Intevo Bold | 130 | 12x0.6 | 7.2 | 12x0.6 | 1.0305 | 1.5948
Intevo Bold | 130 | 16x0.6 | 9.6 | 16x0.6 | 1.031 | 1.5924
Intevo Bold | 130 | 1x10 | 10 | 1x10 | 1.0305 | 1.5934
Intevo Bold | 130 | 12x1.2 | 14.4 | 12x1.2 | 1.0309 | 1.5911
Intevo Bold | 130 | 2x8 | 16 | 2x8 | 1.0305 | 1.5915
Intevo Bold | 130 | 16x1.2 | 19.2 | 16x1.2 | 1.0308 | 1.5928
Definition Edge | 70 | 2x1 | 2 | 2x1 | 1.0538 | 1.8525
Definition Edge | 70 | 6x0.6 | 3.6 | 6x0.6 | 1.0498 | 1.8485
Definition Edge | 70 | 8x0.6 | 4.8 | 8x0.6 | 1.049 | 1.8462
Definition Edge | 70 | 8x0.6 | 4.8 | 16x0.6 | 1.0437 | 1.8411
Definition Edge | 70 | 1x5 | 5 | 1x5 | 1.0513 | 1.8364
Definition Edge | 70 | 10x0.6 | 6 | 20x0.6 | 1.0442 | 1.8514
Definition Edge | 70 | 6x1.2 | 7.2 | 6x1.2 | 1.0506 | 1.8481
Definition Edge | 70 | 1x10 | 10 | 1x10 | 1.0513 | 1.8364
Definition Edge | 70 | 20x0.6 | 12 | 40x0.6 | 1.0464 | 1.8667
Definition Edge | 70 | 12x1.2 | 14.4 | 12x1.2 | 1.05 | 1.8571
Definition Edge | 70 | 32x1.2 | 38.4 | 32x1.2 | 1.0496 | 1.8689
Definition Edge | 70 | 64x0.6 | 38.4 | 28x0.6 | 1.0449 | 1.7164
Definition Edge | 80 | 2x1 | 2 | 2x1 | 1.0431 | 1.7383
Definition Edge | 80 | 6x0.6 | 3.6 | 6x0.6 | 1.0444 | 1.7399
Definition Edge | 80 | 8x0.6 | 4.8 | 8x0.6 | 1.0434 | 1.7421
Definition Edge | 80 | 8x0.6 | 4.8 | 16x0.6 | 1.0304 | 1.7377
Definition Edge | 80 | 1x5 | 5 | 1x5 | 1.0427 | 1.7396
Definition Edge | 80 | 10x0.6 | 6 | 20x0.6 | 1.0299 | 1.7402
Definition Edge | 80 | 6x1.2 | 7.2 | 6x1.2 | 1.0427 | 1.7391
Definition Edge | 80 | 1x10 | 10 | 1x10 | 1.0427 | 1.7396
Definition Edge | 80 | 20x0.6 | 12 | 40x0.6 | 1.0293 | 1.7442
Definition Edge | 80 | 12x1.2 | 14.4 | 12x1.2 | 1.0437 | 1.7398
Definition Edge | 80 | 32x1.2 | 38.4 | 32x1.2 | 1.0429 | 1.7315
Definition Edge | 80 | 64x0.6 | 38.4 | 28x0.6 | 1.0284 | 1.7383
Definition Edge | 100 | 2x1 | 2 | 2x1 | 1.0249 | 1.6049
Definition Edge | 100 | 6x0.6 | 3.6 | 6x0.6 | 1.0249 | 1.6041
Definition Edge | 100 | 8x0.6 | 4.8 | 8x0.6 | 1.0255 | 1.605
Definition Edge | 100 | 8x0.6 | 4.8 | 16x0.6 | 1.0163 | 1.611
Definition Edge | 100 | 1x5 | 5 | 1x5 | 1.0251 | 1.6055
Definition Edge | 100 | 10x0.6 | 6 | 20x0.6 | 1.0157 | 1.6082
Definition Edge | 100 | 6x1.2 | 7.2 | 6x1.2 | 1.0248 | 1.6019
Definition Edge | 100 | 1x10 | 10 | 1x10 | 1.0251 | 1.6055
Definition Edge | 100 | 20x0.6 | 12 | 40x0.6 | 1.0164 | 1.6115
Definition Edge | 100 | 12x1.2 | 14.4 | 12x1.2 | 1.0257 | 1.6036
Definition Edge | 100 | 32x1.2 | 38.4 | 32x1.2 | 1.0248 | 1.6066
Definition Edge | 100 | 64x0.6 | 38.4 | 28x0.6 | 1.0163 | 1.6082
Definition Edge | 120 | 2x1 | 2 | 2x1 | 1.0188 | 1.5444
Definition Edge | 120 | 6x0.6 | 3.6 | 6x0.6 | 1.0192 | 1.544
Definition Edge | 120 | 8x0.6 | 4.8 | 8x0.6 | 1.0194 | 1.5447
Definition Edge | 120 | 8x0.6 | 4.8 | 16x0.6 | 1.0123 | 1.54
Definition Edge | 120 | 1x5 | 5 | 1x5 | 1.0193 | 1.5443
Definition Edge | 120 | 10x0.6 | 6 | 20x0.6 | 1.0124 | 1.5404
Definition Edge | 120 | 6x1.2 | 7.2 | 6x1.2 | 1.019 | 1.5453
Definition Edge | 120 | 1x10 | 10 | 1x10 | 1.0193 | 1.5443
Definition Edge | 120 | 20x0.6 | 12 | 40x0.6 | 1.0128 | 1.5396
Definition Edge | 120 | 12x1.2 | 14.4 | 12x1.2 | 1.0195 | 1.5436
Definition Edge | 120 | 32x1.2 | 38.4 | 32x1.2 | 1.0187 | 1.5442
Definition Edge | 120 | 64x0.6 | 38.4 | 28x0.6 | 1.0127 | 1.5423
Definition Edge | 140 | 2x1 | 2 | 2x1 | 1.0178 | 1.5129
Definition Edge | 140 | 6x0.6 | 3.6 | 6x0.6 | 1.0177 | 1.5122
Definition Edge | 140 | 8x0.6 | 4.8 | 8x0.6 | 1.0176 | 1.5123
Definition Edge | 140 | 8x0.6 | 4.8 | 16x0.6 | 1.0133 | 1.5203
Definition Edge | 140 | 1x5 | 5 | 1x5 | 1.0176 | 1.5127
Definition Edge | 140 | 10x0.6 | 6 | 20x0.6 | 1.0132 | 1.5218
Definition Edge | 140 | 6x1.2 | 7.2 | 6x1.2 | 1.0172 | 1.5124
Definition Edge | 140 | 1x10 | 10 | 1x10 | 1.0176 | 1.5127
Definition Edge | 140 | 20x0.6 | 12 | 40x0.6 | 1.0133 | 1.5201
Definition Edge | 140 | 12x1.2 | 14.4 | 12x1.2 | 1.0172 | 1.5119
Definition Edge | 140 | 32x1.2 | 38.4 | 32x1.2 | 1.0177 | 1.5121
Definition Edge | 140 | 64x0.6 | 38.4 | 28x0.6 | 1.0133 | 1.5207
Somatom Force | 70 | 2x1 | 2 | 2x1 | 1.0751 | 1.8793
Somatom Force | 70 | 6x0.6 | 3.6 | 6x0.6 | 1.076 | 1.866
Somatom Force | 70 | 1x5 | 5 | 1x5 | 1.0779 | 1.875
Somatom Force | 70 | 1x10 | 10 | 1x10 | 1.0779 | 1.875
Somatom Force | 70 | 12x1.2 | 14.4 | 12x1.2 | 1.0776 | 1.875
Somatom Force | 70 | 32x0.6 | 19.2 | 32x0.6 | 1.0764 | 1.8438
Somatom Force | 70 | 32x0.6 | 19.2 | 64x0.6 | 1.0743 | 1.8676
Somatom Force | 70 | 32x0.6 | 19.2 | 64x0.6 | 1.0774 | 1.8462
Somatom Force | 70 | 48x0.6 | 28.8 | 96x0.6 | 1.0754 | 1.8714
Somatom Force | 70 | 64x0.6 | 38.4 | 128x0.6 | 1.0759 | 1.8507
Somatom Force | 70 | 96x0.6 | 57.6 | 192x0.6 | 1.0772 | 1.8571
Somatom Force | 70 | 24x1.2 | 28.8 | 24x1.2 | 1.0767 | 1.8696
Somatom Force | 70 | 48x1.2 | 57.6 | 48x1.2 | 1.0733 | 1.8571
Somatom Force | 80 | 2x1 | 2 | 2x1 | 1.0495 | 1.7282
Somatom Force | 80 | 6x0.6 | 3.6 | 6x0.6 | 1.0505 | 1.7267
Somatom Force | 80 | 1x5 | 5 | 1x5 | 1.0513 | 1.72
Somatom Force | 80 | 1x10 | 10 | 1x10 | 1.0513 | 1.72
Somatom Force | 80 | 12x1.2 | 14.4 | 12x1.2 | 1.0522 | 1.7324
Somatom Force | 80 | 32x0.6 | 19.2 | 32x0.6 | 1.0523 | 1.7321
Somatom Force | 80 | 32x0.6 | 19.2 | 64x0.6 | 1.053 | 1.719
Somatom Force | 80 | 32x0.6 | 19.2 | 64x0.6 | 1.0521 | 1.7299
Somatom Force | 80 | 48x0.6 | 28.8 | 96x0.6 | 1.0513 | 1.7339
Somatom Force | 80 | 64x0.6 | 38.4 | 128x0.6 | 1.0519 | 1.7288
Somatom Force | 80 | 96x0.6 | 57.6 | 192x0.6 | 1.0529 | 1.7297
Somatom Force | 80 | 24x1.2 | 28.8 | 24x1.2 | 1.0501 | 1.7295
Somatom Force | 80 | 48x1.2 | 57.6 | 48x1.2 | 1.0529 | 1.7297
Somatom Force | 90 | 2x1 | 2 | 2x1 | 1.0253 | 1.6543
Somatom Force | 90 | 6x0.6 | 3.6 | 6x0.6 | 1.0253 | 1.658
Somatom Force | 90 | 1x5 | 5 | 1x5 | 1.0244 | 1.6603
Somatom Force | 90 | 1x10 | 10 | 1x10 | 1.0244 | 1.6603
Somatom Force | 90 | 12x1.2 | 14.4 | 12x1.2 | 1.0245 | 1.6547
Somatom Force | 90 | 32x0.6 | 19.2 | 32x0.6 | 1.0248 | 1.6591
Somatom Force | 90 | 32x0.6 | 19.2 | 64x0.6 | 1.0245 | 1.6561
Somatom Force | 90 | 32x0.6 | 19.2 | 64x0.6 | 1.0241 | 1.6558
Somatom Force | 90 | 48x0.6 | 28.8 | 96x0.6 | 1.0237 | 1.6564
Somatom Force | 90 | 64x0.6 | 38.4 | 128x0.6 | 1.025 | 1.6595
Somatom Force | 90 | 96x0.6 | 57.6 | 192x0.6 | 1.025 | 1.7
Somatom Force | 90 | 24x1.2 | 28.8 | 24x1.2 | 1.0242 | 1.6562
Somatom Force | 90 | 48x1.2 | 57.6 | 48x1.2 | 1.0234 | 1.6609
Somatom Force | 100 | 2x1 | 2 | 2x1 | 1.0186 | 1.6034
Somatom Force | 100 | 6x0.6 | 3.6 | 6x0.6 | 1.0186 | 1.6062
Somatom Force | 100 | 1x5 | 5 | 1x5 | 1.0179 | 1.6027
Somatom Force | 100 | 1x10 | 10 | 1x10 | 1.0179 | 1.6027
Somatom Force | 100 | 12x1.2 | 14.4 | 12x1.2 | 1.018 | 1.6081
Somatom Force | 100 | 32x0.6 | 19.2 | 32x0.6 | 1.0182 | 1.6071
Somatom Force | 100 | 32x0.6 | 19.2 | 64x0.6 | 1.019 | 1.6052
Somatom Force | 100 | 32x0.6 | 19.2 | 64x0.6 | 1.0186 | 1.6071
Somatom Force | 100 | 48x0.6 | 28.8 | 96x0.6 | 1.0184 | 1.6036
Somatom Force | 100 | 64x0.6 | 38.4 | 128x0.6 | 1.0184 | 1.6075
Somatom Force | 100 | 96x0.6 | 57.6 | 192x0.6 | 1.0184 | 1.604
Somatom Force | 100 | 24x1.2 | 28.8 | 24x1.2 | 1.0188 | 1.6036
Somatom Force | 100 | 48x1.2 | 57.6 | 48x1.2 | 1.0183 | 1.604
Somatom Force | 110 | 2x1 | 2 | 2x1 | 1.0153 | 1.5563
Somatom Force | 110 | 6x0.6 | 3.6 | 6x0.6 | 1.0149 | 1.5549
Somatom Force | 110 | 1x5 | 5 | 1x5 | 1.0149 | 1.5548
Somatom Force | 110 | 1x10 | 10 | 1x10 | 1.0149 | 1.5548
Somatom Force | 110 | 12x1.2 | 14.4 | 12x1.2 | 1.0146 | 1.5548
Somatom Force | 110 | 32x0.6 | 19.2 | 32x0.6 | 1.0149 | 1.5575
Somatom Force | 110 | 32x0.6 | 19.2 | 64x0.6 | 1.0147 | 1.5549
Somatom Force | 110 | 32x0.6 | 19.2 | 64x0.6 | 1.0144 | 1.5556
Somatom Force | 110 | 48x0.6 | 28.8 | 96x0.6 | 1.0151 | 1.5559
Somatom Force | 110 | 64x0.6 | 38.4 | 128x0.6 | 1.015 | 1.5546
Somatom Force | 110 | 96x0.6 | 57.6 | 192x0.6 | 1.0151 | 1.5536
Somatom Force | 110 | 24x1.2 | 28.8 | 24x1.2 | 1.0145 | 1.5556
Somatom Force | 110 | 48x1.2 | 57.6 | 48x1.2 | 1.0151 | 1.5536
Somatom Force | 120 | 2x1 | 2 | 2x1 | 1.0123 | 1.5423
Somatom Force | 120 | 6x0.6 | 3.6 | 6x0.6 | 1.0129 | 1.5456
Somatom Force | 120 | 1x5 | 5 | 1x5 | 1.0119 | 1.5438
Somatom Force | 120 | 1x10 | 10 | 1x10 | 1.0119 | 1.5438
Somatom Force | 120 | 12x1.2 | 14.4 | 12x1.2 | 1.0122 | 1.5461
Somatom Force | 120 | 32x0.6 | 19.2 | 32x0.6 | 1.0126 | 1.5434
Somatom Force | 120 | 32x0.6 | 19.2 | 64x0.6 | 1.0124 | 1.5447
Somatom Force | 120 | 32x0.6 | 19.2 | 64x0.6 | 1.0127 | 1.5449
Somatom Force | 120 | 48x0.6 | 28.8 | 96x0.6 | 1.0127 | 1.5443
Somatom Force | 120 | 64x0.6 | 38.4 | 128x0.6 | 1.0127 | 1.5457
Somatom Force | 120 | 96x0.6 | 57.6 | 192x0.6 | 1.0128 | 1.545
Somatom Force | 120 | 24x1.2 | 28.8 | 24x1.2 | 1.0123 | 1.5462
Somatom Force | 120 | 48x1.2 | 57.6 | 48x1.2 | 1.0121 | 1.545
Somatom Force | 130 | 2x1 | 2 | 2x1 | 1.0121 | 1.5252
Somatom Force | 130 | 6x0.6 | 3.6 | 6x0.6 | 1.0122 | 1.5254
Somatom Force | 130 | 1x5 | 5 | 1x5 | 1.0118 | 1.525
Somatom Force | 130 | 1x10 | 10 | 1x10 | 1.0118 | 1.525
Somatom Force | 130 | 12x1.2 | 14.4 | 12x1.2 | 1.0115 | 1.5263
Somatom Force | 130 | 32x0.6 | 19.2 | 32x0.6 | 1.0117 | 1.5268
Somatom Force | 130 | 32x0.6 | 19.2 | 64x0.6 | 1.012 | 1.525
Somatom Force | 130 | 32x0.6 | 19.2 | 64x0.6 | 1.012 | 1.5273
Somatom Force | 130 | 48x0.6 | 28.8 | 96x0.6 | 1.0116 | 1.5275
Somatom Force | 130 | 64x0.6 | 38.4 | 128x0.6 | 1.0122 | 1.5237
Somatom Force | 130 | 96x0.6 | 57.6 | 192x0.6 | 1.0118 | 1.5271
Somatom Force | 130 | 24x1.2 | 28.8 | 24x1.2 | 1.0118 | 1.5246
Somatom Force | 130 | 48x1.2 | 57.6 | 48x1.2 | 1.0118 | 1.5243
Somatom Force | 140 | 2x1 | 2 | 2x1 | 1.0108 | 1.5134
Somatom Force | 140 | 6x0.6 | 3.6 | 6x0.6 | 1.0107 | 1.5125
Somatom Force | 140 | 1x5 | 5 | 1x5 | 1.0106 | 1.5121
Somatom Force | 140 | 1x10 | 10 | 1x10 | 1.0106 | 1.5121
Somatom Force | 140 | 12x1.2 | 14.4 | 12x1.2 | 1.0106 | 1.5115
Somatom Force | 140 | 32x0.6 | 19.2 | 32x0.6 | 1.0104 | 1.5123
Somatom Force | 140 | 32x0.6 | 19.2 | 64x0.6 | 1.0111 | 1.5114
Somatom Force | 140 | 32x0.6 | 19.2 | 64x0.6 | 1.0105 | 1.5113
Somatom Force | 140 | 48x0.6 | 28.8 | 96x0.6 | 1.0107 | 1.5104
Somatom Force | 140 | 64x0.6 | 38.4 | 128x0.6 | 1.0108 | 1.5102
Somatom Force | 140 | 96x0.6 | 57.6 | 192x0.6 | 1.0105 | 1.5116
Somatom Force | 140 | 24x1.2 | 28.8 | 24x1.2 | 1.0109 | 1.5113
Somatom Force | 140 | 48x1.2 | 57.6 | 48x1.2 | 1.011 | 1.5108
Somatom Force | 150 | 2x1 | 2 | 2x1 | 1.0112 | 1.5007
Somatom Force | 150 | 6x0.6 | 3.6 | 6x0.6 | 1.0112 | 1.4996
Somatom Force | 150 | 1x5 | 5 | 1x5 | 1.0111 | 1.5007
Somatom Force | 150 | 1x10 | 10 | 1x10 | 1.0111 | 1.5007
Somatom Force | 150 | 12x1.2 | 14.4 | 12x1.2 | 1.0111 | 1.4995
Somatom Force | 150 | 32x0.6 | 19.2 | 32x0.6 | 1.0111 | 1.4994
Somatom Force | 150 | 32x0.6 | 19.2 | 64x0.6 | 1.0107 | 1.5042
Somatom Force | 150 | 32x0.6 | 19.2 | 64x0.6 | 1.0112 | 1.4995
Somatom Force | 150 | 48x0.6 | 28.8 | 96x0.6 | 1.0112 | 1.5
Somatom Force | 150 | 64x0.6 | 38.4 | 128x0.6 | 1.011 | 1.4982
Somatom Force | 150 | 96x0.6 | 57.6 | 192x0.6 | 1.0108 | 1.326
Somatom Force | 150 | 24x1.2 | 28.8 | 24x1.2 | 1.011 | 1.4994
Somatom Force | 150 | 48x1.2 | 57.6 | 48x1.2 | 1.0108 | 1.4993
Somatom Force | Sn100 | 2x1 | 2 | 2x1 | 0.97468 | 1.3929
Somatom Force | Sn100 | 6x0.6 | 3.6 | 6x0.6 | 0.98473 | 1.4043
Somatom Force | Sn100 | 1x5 | 5 | 1x5 | 0.98684 | 1.4074
Somatom Force | Sn100 | 1x10 | 10 | 1x10 | 0.98684 | 1.4074
Somatom Force | Sn100 | 12x1.2 | 14.4 | 12x1.2 | 0.99074 | 1.4211
Somatom Force | Sn100 | 32x0.6 | 19.2 | 32x0.6 | 0.97674 | 1.4333
Somatom Force | Sn100 | 32x0.6 | 19.2 | 64x0.6 | 0.97826 | 1.3939
Somatom Force | Sn100 | 32x0.6 | 19.2 | 64x0.6 | 0.98095 | 1.4054
Somatom Force | Sn100 | 48x0.6 | 28.8 | 96x0.6 | 0.97895 | 1.3824
Somatom Force | Sn100 | 64x0.6 | 38.4 | 128x0.6 | 0.98889 | 1.4062
Somatom Force | Sn100 | 96x0.6 | 57.6 | 192x0.6 | 0.97647 | 1.4
Somatom Force | Sn100 | 24x1.2 | 28.8 | 24x1.2 | 0.98925 | 1.4242
Somatom Force | Sn100 | 48x1.2 | 57.6 | 48x1.2 | 0.97647 | 1.4
Somatom Force | Sn150 | 2x1 | 2 | 2x1 | 1.004 | 1.3776
Somatom Force | Sn150 | 6x0.6 | 3.6 | 6x0.6 | 1.0048 | 1.3884
Somatom Force | Sn150 | 1x5 | 5 | 1x5 | 1.0041 | 1.381
Somatom Force | Sn150 | 1x10 | 10 | 1x10 | 1.0041 | 1.381
Somatom Force | Sn150 | 12x1.2 | 14.4 | 12x1.2 | 1.0044 | 1.3815
Somatom Force | Sn150 | 32x0.6 | 19.2 | 32x0.6 | 1.0055 | 1.3785
Somatom Force | Sn150 | 32x0.6 | 19.2 | 64x0.6 | 1.0051 | 1.3799
Somatom Force | Sn150 | 32x0.6 | 19.2 | 64x0.6 | 1.0045 | 1.3793
Somatom Force | Sn150 | 48x0.6 | 28.8 | 96x0.6 | 1.005 | 1.3755
Somatom Force | Sn150 | 64x0.6 | 38.4 | 128x0.6 | 1.0053 | 1.3778
Somatom Force | Sn150 | 96x0.6 | 57.6 | 192x0.6 | 1.0037 | 1.3839
Somatom Force | Sn150 | 24x1.2 | 28.8 | 24x1.2 | 1.0051 | 1.3836
Somatom Force | Sn150 | 48x1.2 | 57.6 | 48x1.2 | 1.0037 | 1.3839
Somatom Force | 80/Sn150 | 32x0.6 | 19.2 | 64x0.6 |  | 1.6994
Somatom Force | 80/Sn150 | 64x0.6 | 38.4 | 128x0.6 |  | 1.6975
Somatom Force | 80/Sn150 | 96x0.6 | 57.6 | 192x0.6 |  | 1.7378
Somatom Force | 90/Sn150 | 32x0.6 | 19.2 | 64x0.6 |  | 1.6792
Somatom Force | 90/Sn150 | 64x0.6 | 38.4 | 128x0.6 |  | 1.6571
Somatom Force | 90/Sn150 | 96x0.6 | 57.6 | 192x0.6 |  | 1.6726
Somatom Force | 100/Sn150 | 32x0.6 | 19.2 | 64x0.6 |  | 1.6912
Somatom Force | 100/Sn150 | 64x0.6 | 38.4 | 128x0.6 |  | 1.6095
Somatom Force | 100/Sn150 | 96x0.6 | 57.6 | 192x0.6 |  | 1.6803
Somatom Flash (B) | 70 | 2x1 | 2 | 2x1 |  | 2.4286
Somatom Flash (B) | 70 | 6x0.6 | 3.6 | 6x0.6 |  | 2.25
Somatom Flash (B) | 70 | 8x0.6 | 2.4 | 16x0.6 |  | 2.4167
Somatom Flash (B) | 70 | 1x5 | 5 | 1x5 |  | 2.5
Somatom Flash (B) | 70 | 10x0.6 | 6 | 20x0.6 |  | 2.2222
Somatom Flash (B) | 70 | 1x10 | 10 | 1x10 |  | 2.5
Somatom Flash (B) | 70 | 12x1.2 | 14.4 | 12x1.2 |  | 2.5
Somatom Flash (B) | 70 | 20x0.6 | 12 | 40x0.6 |  | 2.3333
Somatom Flash (B) | 70 | 32x0.6 | 19.2 | 64x0.6 |  | 2.375
Somatom Flash (B) | 70 | 32x1.2 | 38.4 | 32x1.2 |  | 2.4286
Somatom Flash (B) | 70 | 64x0.6 | 38.4 | 128x0.6 |  | 2.4286
Somatom Flash (B) | 80 | 2x1 | 2 | 2x1 |  | 2.1667
Somatom Flash (B) | 80 | 6x0.6 | 3.6 | 6x0.6 |  | 2.15
Somatom Flash (B) | 80 | 8x0.6 | 2.4 | 16x0.6 |  | 2.1429
Somatom Flash (B) | 80 | 1x5 | 5 | 1x5 |  | 2.1818
Somatom Flash (B) | 80 | 10x0.6 | 6 | 20x0.6 |  | 2.0667
Somatom Flash (B) | 80 | 1x10 | 10 | 1x10 |  | 2.1818
Somatom Flash (B) | 80 | 12x1.2 | 14.4 | 12x1.2 |  | 2.1429
Somatom Flash (B) | 80 | 20x0.6 | 12 | 40x0.6 |  | 2.1333
Somatom Flash (B) | 80 | 32x0.6 | 19.2 | 64x0.6 |  | 2.0714
Somatom Flash (B) | 80 | 32x1.2 | 38.4 | 32x1.2 |  | 2.1667
Somatom Flash (B) | 80 | 64x0.6 | 38.4 | 128x0.6 |  | 2.1667
Somatom Flash (B) | 100 | 2x1 | 2 | 2x1 |  | 1.9259
Somatom Flash (B) | 100 | 6x0.6 | 3.6 | 6x0.6 |  | 1.9091
Somatom Flash (B) | 100 | 8x0.6 | 2.4 | 16x0.6 |  | 1.8936
Somatom Flash (B) | 100 | 1x5 | 5 | 1x5 |  | 1.84
Somatom Flash (B) | 100 | 10x0.6 | 6 | 20x0.6 |  | 1.8485
Somatom Flash (B) | 100 | 1x10 | 10 | 1x10 |  | 1.84
Somatom Flash (B) | 100 | 12x1.2 | 14.4 | 12x1.2 |  | 1.875
Somatom Flash (B) | 100 | 20x0.6 | 12 | 40x0.6 |  | 1.9091
Somatom Flash (B) | 100 | 32x0.6 | 19.2 | 64x0.6 |  | 1.871
Somatom Flash (B) | 100 | 32x1.2 | 38.4 | 32x1.2 |  | 1.8571
Somatom Flash (B) | 100 | 64x0.6 | 38.4 | 128x0.6 |  | 1.8889
Somatom Flash (B) | 120 | 2x1 | 2 | 2x1 |  | 1.766
Somatom Flash (B) | 120 | 6x0.6 | 3.6 | 6x0.6 |  | 1.7632
Somatom Flash (B) | 120 | 8x0.6 | 2.4 | 16x0.6 |  | 1.775
Somatom Flash (B) | 120 | 1x5 | 5 | 1x5 |  | 1.7619
Somatom Flash (B) | 120 | 10x0.6 | 6 | 20x0.6 |  | 1.7679
Somatom Flash (B) | 120 | 1x10 | 10 | 1x10 |  | 1.7619
Somatom Flash (B) | 120 | 12x1.2 | 14.4 | 12x1.2 |  | 1.7778
Somatom Flash (B) | 120 | 20x0.6 | 12 | 40x0.6 |  | 1.7544
Somatom Flash (B) | 120 | 32x0.6 | 19.2 | 64x0.6 |  | 1.7736
Somatom Flash (B) | 120 | 32x1.2 | 38.4 | 32x1.2 |  | 1.766
Somatom Flash (B) | 120 | 64x0.6 | 38.4 | 128x0.6 |  | 1.7826
Somatom Flash (B) | 140 | 2x1 | 2 | 2x1 |  | 1.6986
Somatom Flash (B) | 140 | 6x0.6 | 3.6 | 6x0.6 |  | 1.7034
Somatom Flash (B) | 140 | 8x0.6 | 2.4 | 16x0.6 |  | 1.6825
Somatom Flash (B) | 140 | 1x5 | 5 | 1x5 |  | 1.6818
Somatom Flash (B) | 140 | 10x0.6 | 6 | 20x0.6 |  | 1.6897
Somatom Flash (B) | 140 | 1x10 | 10 | 1x10 |  | 1.6818
Somatom Flash (B) | 140 | 12x1.2 | 14.4 | 12x1.2 |  | 1.7024
Somatom Flash (B) | 140 | 20x0.6 | 12 | 40x0.6 |  | 1.6854
Somatom Flash (B) | 140 | 32x0.6 | 19.2 | 64x0.6 |  | 1.6867
Somatom Flash (B) | 140 | 32x1.2 | 38.4 | 32x1.2 |  | 1.7123
Somatom Flash (B) | 140 | 64x0.6 | 38.4 | 128x0.6 |  | 1.6944
Somatom Flash (B) | Sn140 | 2x1 | 2 | 2x1 |  | 1.48
Somatom Flash (B) | Sn140 | 6x0.6 | 3.6 | 6x0.6 |  | 1.4634
Somatom Flash (B) | Sn140 | 8x0.6 | 2.4 | 16x0.6 |  | 1.4651
Somatom Flash (B) | Sn140 | 1x5 | 5 | 1x5 |  | 1.4348
Somatom Flash (B) | Sn140 | 10x0.6 | 6 | 20x0.6 |  | 1.4667
Somatom Flash (B) | Sn140 | 1x10 | 10 | 1x10 |  | 1.4348
Somatom Flash (B) | Sn140 | 12x1.2 | 14.4 | 12x1.2 |  | 1.4483
Somatom Flash (B) | Sn140 | 20x0.6 | 12 | 40x0.6 |  | 1.4516
Somatom Flash (B) | Sn140 | 32x0.6 | 19.2 | 64x0.6 |  | 1.4643
Somatom Flash (B) | Sn140 | 32x1.2 | 38.4 | 32x1.2 |  | 1.48
Somatom Flash (B) | Sn140 | 64x0.6 | 38.4 | 128x0.6 |  | 1.44
Somatom Flash (A) | 70 | 2x1 | 2 | 2x1 | 1.1379 | 2.1667
Somatom Flash (A) | 70 | 6x0.6 | 3.6 | 6x0.6 | 1.1277 | 2.1
Somatom Flash (A) | 70 | 8x0.6 | 2.4 | 16x0.6 | 1.12 | 2
Somatom Flash (A) | 70 | 1x5 | 5 | 1x5 | 1.1154 | 1.8333
Somatom Flash (A) | 70 | 10x0.6 | 6 | 20x0.6 | 1.1471 | 1.875
Somatom Flash (A) | 70 | 1x10 | 10 | 1x10 | 1.1154 | 1.8333
Somatom Flash (A) | 70 | 12x1.2 | 14.4 | 12x1.2 | 1.1515 | 2.1429
Somatom Flash (A) | 70 | 20x0.6 | 12 | 40x0.6 | 1.1429 | 1.875
Somatom Flash (A) | 70 | 32x0.6 | 19.2 | 64x0.6 | 1.1212 | 2
Somatom Flash (A) | 70 | 32x1.2 | 38.4 | 32x1.2 | 1.1379 | 2.1667
Somatom Flash (A) | 70 | 64x0.6 | 38.4 | 128x0.6 | 1.1429 | 2
Somatom Flash (A) | 80 | 2x1 | 2 | 2x1 | 1.087 | 1.8182
Somatom Flash (A) | 80 | 6x0.6 | 3.6 | 6x0.6 | 1.1081 | 1.7778
Somatom Flash (A) | 80 | 8x0.6 | 2.4 | 16x0.6 | 1.0886 | 1.7895
Somatom Flash (A) | 80 | 1x5 | 5 | 1x5 | 1.0976 | 1.8
Somatom Flash (A) | 80 | 10x0.6 | 6 | 20x0.6 | 1.1111 | 1.8462
Somatom Flash (A) | 80 | 1x10 | 10 | 1x10 | 1.0976 | 1.8
Somatom Flash (A) | 80 | 12x1.2 | 14.4 | 12x1.2 | 1.0943 | 1.7692
Somatom Flash (A) | 80 | 20x0.6 | 12 | 40x0.6 | 1.0893 | 1.8462
Somatom Flash (A) | 80 | 32x0.6 | 19.2 | 64x0.6 | 1.0962 | 1.8333
Somatom Flash (A) | 80 | 32x1.2 | 38.4 | 32x1.2 | 1.1087 | 1.8182
Somatom Flash (A) | 80 | 64x0.6 | 38.4 | 128x0.6 | 1.1111 | 1.8182
Somatom Flash (A) | 100 | 2x1 | 2 | 2x1 | 1.0652 | 1.68
Somatom Flash (A) | 100 | 6x0.6 | 3.6 | 6x0.6 | 1.0671 | 1.6585
Somatom Flash (A) | 100 | 8x0.6 | 2.4 | 16x0.6 | 1.0633 | 1.6744
Somatom Flash (A) | 100 | 1x5 | 5 | 1x5 | 1.0732 | 1.6522
Somatom Flash (A) | 100 | 10x0.6 | 6 | 20x0.6 | 1.0642 | 1.6667
Somatom Flash (A) | 100 | 1x10 | 10 | 1x10 | 1.0732 | 1.6522
Somatom Flash (A) | 100 | 12x1.2 | 14.4 | 12x1.2 | 1.066 | 1.6552
Somatom Flash (A) | 100 | 20x0.6 | 12 | 40x0.6 | 1.0625 | 1.6452
Somatom Flash (A) | 100 | 32x0.6 | 19.2 | 64x0.6 | 1.0577 | 1.6207
Somatom Flash (A) | 100 | 32x1.2 | 38.4 | 32x1.2 | 1.0652 | 1.68
Somatom Flash (A) | 100 | 64x0.6 | 38.4 | 128x0.6 | 1.0549 | 1.64
Somatom Flash (A) | 120 | 2x1 | 2 | 2x1 | 1.0473 | 1.5682
Somatom Flash (A) | 120 | 6x0.6 | 3.6 | 6x0.6 | 1.0502 | 1.5634
Somatom Flash (A) | 120 | 8x0.6 | 2.4 | 16x0.6 | 1.0472 | 1.56
Somatom Flash (A) | 120 | 1x5 | 5 | 1x5 | 1.0451 | 1.5897
Somatom Flash (A) | 120 | 10x0.6 | 6 | 20x0.6 | 1.0455 | 1.5577
Somatom Flash (A) | 120 | 1x10 | 10 | 1x10 | 1.0451 | 1.5897
Somatom Flash (A) | 120 | 12x1.2 | 14.4 | 12x1.2 | 1.0471 | 1.58
Somatom Flash (A) | 120 | 20x0.6 | 12 | 40x0.6 | 1.0503 | 1.566
Somatom Flash (A) | 120 | 32x0.6 | 19.2 | 64x0.6 | 1.0479 | 1.5714
Somatom Flash (A) | 120 | 32x1.2 | 38.4 | 32x1.2 | 1.0541 | 1.5682
Somatom Flash (A) | 120 | 64x0.6 | 38.4 | 128x0.6 | 1.0479 | 1.5581
Somatom Flash (A) | 140 | 2x1 | 2 | 2x1 | 1.0409 | 1.5385
Somatom Flash (A) | 140 | 6x0.6 | 3.6 | 6x0.6 | 1.0421 | 1.5283
Somatom Flash (A) | 140 | 8x0.6 | 2.4 | 16x0.6 | 1.0397 | 1.5268
Somatom Flash (A) | 140 | 1x5 | 5 | 1x5 | 1.0406 | 1.5254
Somatom Flash (A) | 140 | 10x0.6 | 6 | 20x0.6 | 1.0382 | 1.5128
Somatom Flash (A) | 140 | 1x10 | 10 | 1x10 | 1.0406 | 1.5254
Somatom Flash (A) | 140 | 12x1.2 | 14.4 | 12x1.2 | 1.0395 | 1.5333
Somatom Flash (A) | 140 | 20x0.6 | 12 | 40x0.6 | 1.0412 | 1.5316
Somatom Flash (A) | 140 | 32x0.6 | 19.2 | 64x0.6 | 1.0402 | 1.5135
Somatom Flash (A) | 140 | 32x1.2 | 38.4 | 32x1.2 | 1.0407 | 1.5152
Somatom Flash (A) | 140 | 64x0.6 | 38.4 | 128x0.6 | 1.0367 | 1.5077
Somatom Flash (A) | 80/Sn140 | 32x0.6 | 19.2 | 64x0.6 |  | 1.7183
Somatom Flash (A) | 100/Sn140 | 32x0.6 | 19.2 | 64x0.6 |  | 1.6565
Somatom Flash (A) | 140/80 | 32x0.6 | 19.2 | 64x0.6 |  | 1.8719
Somatom Flash (B) | Sn140 | 2x1 | 2 | 2x1 | 1
Somatom Flash (B) | Sn140 | 6x0.6 | 3.6 | 6x0.6 | 1.0094
Somatom Flash (B) | Sn140 | 8x0.6 | 2.4 | 16x0.6 | 1.0088
Somatom Flash (B) | Sn140 | 1x5 | 5 | 1x5 | 1
Somatom Flash (B) | Sn140 | 10x0.6 | 6 | 20x0.6 | 1.0128
Somatom Flash (B) | Sn140 | 1x10 | 10 | 1x10 | 1
Somatom Flash (B) | Sn140 | 12x1.2 | 14.4 | 12x1.2 | 1.0133
Somatom Flash (B) | Sn140 | 20x0.6 | 12 | 40x0.6 | 1
Somatom Flash (B) | Sn140 | 32x0.6 | 19.2 | 64x0.6 | 1.0135
Somatom Flash (B) | Sn140 | 32x1.2 | 38.4 | 32x1.2 | 1
Somatom Flash (B) | Sn140 | 64x0.6 | 38.4 | 128x0.6 | 1
Definition AS | 70 | 2x1 | 2 | 2x1 | 1.0538 | 1.8525
Definition AS | 70 | 6x0.6 | 3.6 | 6x0.6 | 1.0498 | 1.8485
Definition AS | 70 | 8x0.6 | 4.8 | 8x0.6 | 1.049 | 1.8462
Definition AS | 70 | 8x0.6 | 4.8 | 16x0.6 | 1.0437 | 1.8411
Definition AS | 70 | 1x5 | 5 | 1x5 | 1.0513 | 1.8364
Definition AS | 70 | 10x0.6 | 6 | 20x0.6 | 1.0442 | 1.8514
Definition AS | 70 | 6x1.2 | 7.2 | 6x1.2 | 1.0152 | 1.8481
Definition AS | 70 | 1x10 | 10 | 1x10 | 1.0513 | 1.8364
Definition AS | 70 | 20x0.6 | 12 | 40x0.6 | 1.0464 | 1.8667
Definition AS | 70 | 12x1.2 | 14.4 | 12x1.2 | 1.05 | 1.8571
Definition AS | 70 | 32x0.6 | 38.4 | 32x1.2 | 1.0465 | 1.8689
Definition AS | 70 | 64x0.6 | 38.4 | 128x0.6 | 1.0488 | 1.8548
Definition AS | 80 | 2x1 | 2 | 2x1 | 1.0431 | 1.7383
Definition AS | 80 | 6x0.6 | 3.6 | 6x0.6 | 1.0444 | 1.7399
Definition AS | 80 | 8x0.6 | 4.8 | 8x0.6 | 1.0434 | 1.7421
Definition AS | 80 | 8x0.6 | 4.8 | 16x0.6 | 1.0304 | 1.7377
Definition AS | 80 | 1x5 | 5 | 1x5 | 1.0427 | 1.7396
Definition AS | 80 | 10x0.6 | 6 | 20x0.6 | 1.0299 | 1.7402
Definition AS | 80 | 6x1.2 | 7.2 | 6x1.2 | 1.0427 | 1.7391
Definition AS | 80 | 1x10 | 10 | 1x10 | 1.0427 | 1.7396
Definition AS | 80 | 20x0.6 | 12 | 40x0.6 | 1.0293 | 1.7442
Definition AS | 80 | 12x1.2 | 14.4 | 12x1.2 | 1.0437 | 1.7398
Definition AS | 80 | 32x0.6 | 38.4 | 32x1.2 | 1.0294 | 1.7315
Definition AS | 80 | 64x0.6 | 38.4 | 128x0.6 | 1.0435 | 1.7383
Definition AS | 100 | 2x1 | 2 | 2x1 | 1.0249 | 1.6049
Definition AS | 100 | 6x0.6 | 3.6 | 6x0.6 | 1.0249 | 1.6041
Definition AS | 100 | 8x0.6 | 4.8 | 8x0.6 | 1.0255 | 1.605
Definition AS | 100 | 8x0.6 | 4.8 | 16x0.6 | 1.0163 | 1.611
Definition AS | 100 | 1x5 | 5 | 1x5 | 1.0251 | 1.6055
Definition AS | 100 | 10x0.6 | 6 | 20x0.6 | 1.0157 | 1.6082
Definition AS | 100 | 6x1.2 | 7.2 | 6x1.2 | 1.0248 | 1.6019
Definition AS | 100 | 1x10 | 10 | 1x10 | 1.0251 | 1.6055
Definition AS | 100 | 20x0.6 | 12 | 40x0.6 | 1.0164 | 1.6115
Definition AS | 100 | 12x1.2 | 14.4 | 12x1.2 | 1.0257 | 1.6036
Definition AS | 100 | 32x0.6 | 38.4 | 32x1.2 | 1.0248 | 1.6066
Definition AS | 100 | 64x0.6 | 38.4 | 128x0.6 | 1.0163 | 1.6082
Definition AS | 120 | 2x1 | 2 | 2x1 | 1.0188 | 1.5444
Definition AS | 120 | 6x0.6 | 3.6 | 6x0.6 | 1.0192 | 1.544
Definition AS | 120 | 8x0.6 | 4.8 | 8x0.6 | 1.0194 | 1.5447
Definition AS | 120 | 8x0.6 | 4.8 | 16x0.6 | 1.0123 | 1.54
Definition AS | 120 | 1x5 | 5 | 1x5 | 1.0193 | 1.5443
Definition AS | 120 | 10x0.6 | 6 | 20x0.6 | 1.0124 | 1.602
Definition AS | 120 | 6x1.2 | 7.2 | 6x1.2 | 1.019 | 1.5453
Definition AS | 120 | 1x10 | 10 | 1x10 | 1.0193 | 1.5443
Definition AS | 120 | 20x0.6 | 12 | 40x0.6 | 1.0128 | 1.5396
Definition AS | 120 | 12x1.2 | 14.4 | 12x1.2 | 1.0195 | 1.5436
Definition AS | 120 | 32x0.6 | 38.4 | 32x1.2 | 1.0187 | 1.5442
Definition AS | 120 | 64x0.6 | 38.4 | 128x0.6 | 1.0127 | 1.5423
Definition AS | 140 | 2x1 | 2 | 2x1 | 1.0178 | 1.5129
Definition AS | 140 | 6x0.6 | 3.6 | 6x0.6 | 1.0177 | 1.5122
Definition AS | 140 | 8x0.6 | 4.8 | 8x0.6 | 1.0176 | 1.5123
Definition AS | 140 | 8x0.6 | 4.8 | 16x0.6 | 1.0122 | 1.5203
Definition AS | 140 | 1x5 | 5 | 1x5 | 1.0176 | 1.5127
Definition AS | 140 | 10x0.6 | 6 | 20x0.6 | 1.0132 | 1.5218
Definition AS | 140 | 6x1.2 | 7.2 | 6x1.2 | 1.0172 | 1.5124
Definition AS | 140 | 1x10 | 10 | 1x10 | 1.0176 | 1.5127
Definition AS | 140 | 20x0.6 | 12 | 40x0.6 | 1.0133 | 1.5201
Definition AS | 140 | 12x1.2 | 14.4 | 12x1.2 | 1.0172 | 1.5119
Definition AS | 140 | 32x0.6 | 38.4 | 32x1.2 | 1.0177 | 1.5121
Definition AS | 140 | 64x0.6 | 38.4 | 128x0.6 | 1.0133 | 1.5207
Naeotom Alpha | 90 | 1x5 | 5 | 1x5 | 1.0435 | 1.7527
Naeotom Alpha | 90 | 36x0.2 | 7.2 | 36x0.2 | 1.0434 | 1.7452
Naeotom Alpha | 90 | 36x0.4 | 14.4 | 36x0.4 | 1.0437 | 1.7481
Naeotom Alpha | 90 | 48x0.4 | 19.2 | 48x0.4 | 1.0428 | 1.7521
Naeotom Alpha | 90 | 120x0.2 | 24 | 120x0.2 | 1.0429 | 1.7488
Naeotom Alpha | 90 | 72x0.4 | 28.8 | 72x0.4 | 1.0447 | 1.7544
Naeotom Alpha | 90 | 96x0.4 | 38.4 | 96x0.4 | 1.0435 | 1.75
Naeotom Alpha | 90 | 144x0.4 | 57.6 | 144x0.4 | 1.0436 | 1.75
Naeotom Alpha | 120 | 1x5 | 5 | 1x5 | 1.0279 | 1.6175
Naeotom Alpha | 120 | 36x0.2 | 7.2 | 36x0.2 | 1.0277 | 1.6144
Naeotom Alpha | 120 | 36x0.4 | 14.4 | 36x0.4 | 1.0274 | 1.6165
Naeotom Alpha | 120 | 48x0.4 | 19.2 | 48x0.4 | 1.0276 | 1.6165
Naeotom Alpha | 120 | 120x0.2 | 24 | 120x0.2 | 1.0281 | 1.6157
Naeotom Alpha | 120 | 72x0.4 | 28.8 | 72x0.4 | 1.0282 | 1.615
Naeotom Alpha | 120 | 96x0.4 | 38.4 | 96x0.4 | 1.0281 | 1.6178
Naeotom Alpha | 120 | 144x0.4 | 57.6 | 144x0.4 | 1.0274 | 1.6155
Naeotom Alpha | 140 | 1x5 | 5 | 1x5 | 1.022 | 1.571
Naeotom Alpha | 140 | 36x0.2 | 7.2 | 36x0.2 | 1.0222 | 1.5702
Naeotom Alpha | 140 | 36x0.4 | 14.4 | 36x0.4 | 1.022 | 1.5708
Naeotom Alpha | 140 | 48x0.4 | 19.2 | 48x0.4 | 1.0186 | 1.5708
Naeotom Alpha | 140 | 120x0.2 | 24 | 120x0.2 | 1.0219 | 1.5699
Naeotom Alpha | 140 | 72x0.4 | 28.8 | 72x0.4 | 1.0217 | 1.5707
Naeotom Alpha | 140 | 96x0.4 | 38.4 | 96x0.4 | 1.0218 | 1.5709
Naeotom Alpha | 140 | 144x0.4 | 57.6 | 144x0.4 | 1.0218 | 1.571
Naeotom Alpha | Sn100 | 1x5 | 5 | 1x5 | 1 | 1.4717
Naeotom Alpha | Sn100 | 36x0.2 | 7.2 | 36x0.2 | 1 | 1.5
Naeotom Alpha | Sn100 | 36x0.4 | 14.4 | 36x0.4 | 1.0094 | 1.4795
Naeotom Alpha | Sn100 | 48x0.4 | 19.2 | 48x0.4 | 1 | 1.4925
Naeotom Alpha | Sn100 | 120x0.2 | 24 | 120x0.2 | 1 | 1.4746
Naeotom Alpha | Sn100 | 72x0.4 | 28.8 | 72x0.4 | 1 | 1.4769
Naeotom Alpha | Sn100 | 96x0.4 | 38.4 | 96x0.4 | 1 | 1.4754
Naeotom Alpha | Sn100 | 144x0.4 | 57.6 | 144x0.4 | 1 | 1.4655
Naeotom Alpha | Sn140 | 1x5 | 5 | 1x5 | 1.01 | 1.4425
Naeotom Alpha | Sn140 | 36x0.2 | 7.2 | 36x0.2 | 1.0099 | 1.4421
Naeotom Alpha | Sn140 | 36x0.4 | 14.4 | 36x0.4 | 1.0096 | 1.4423
Naeotom Alpha | Sn140 | 48x0.4 | 19.2 | 48x0.4 | 1.0104 | 1.4394
Naeotom Alpha | Sn140 | 120x0.2 | 24 | 120x0.2 | 1.0104 | 1.4422
Naeotom Alpha | Sn140 | 72x0.4 | 28.8 | 72x0.4 | 1.0095 | 1.4404
Naeotom Alpha | Sn140 | 96x0.4 | 38.4 | 96x0.4 | 1.01 | 1.4427
Naeotom Alpha | Sn140 | 144x0.4 | 57.6 | 144x0.4 | 1.0106 | 1.4413
Somatom Drive | 70 | 2x1 | 2 | 2x1 | 1.0467 | 1.8103
Somatom Drive | 70 | 6x0.6 | 3.6 | 6x0.6 | 1.0432 | 1.8085
Somatom Drive | 70 | 8x0.6 | 4.8 | 16x0.6 | 1.0387 | 1.8081
Somatom Drive | 70 | 1x5 | 5 | 1x5 | 1.0433 | 1.8077
Somatom Drive | 70 | 10x0.6 | 6 | 20x0.6 | 1.0395 | 1.7971
Somatom Drive | 70 | 1x10 | 10 | 1x10 | 1.0433 | 1.8077
Somatom Drive | 70 | 12x1.2 | 14.4 | 12x1.2 | 1.0439 | 1.8333
Somatom Drive | 70 | 20x0.6 | 12 | 40x0.6 | 1.0387 | 1.8143
Somatom Drive | 70 | 32x0.6 | 19.2 | 64x0.6 | 1.042 | 1.8281
Somatom Drive | 70 | 12x1.2 | 14.4 | 12x1.2 | 1.0425 | 1.8103
Somatom Drive | 70 | 64x0.6 | 38.4 | 128x0.6 | 1.0397 | 1.807
Somatom Drive | 80 | 2x1 | 2 | 2x1 | 1.0291 | 1.7212
Somatom Drive | 80 | 6x0.6 | 3.6 | 6x0.6 | 1.0315 | 1.716
Somatom Drive | 80 | 8x0.6 | 4.8 | 16x0.6 | 1.0256 | 1.6983
Somatom Drive | 80 | 1x5 | 5 | 1x5 | 1.0298 | 1.7021
Somatom Drive | 80 | 10x0.6 | 6 | 20x0.6 | 1.0267 | 1.7016
Somatom Drive | 80 | 1x10 | 10 | 1x10 | 1.0298 | 1.7021
Somatom Drive | 80 | 12x1.2 | 14.4 | 12x1.2 | 1.0295 | 1.7167
Somatom Drive | 80 | 20x0.6 | 12 | 40x0.6 | 1.0262 | 1.7063
Somatom Drive | 80 | 32x0.6 | 19.2 | 64x0.6 | 1.0284 | 1.7155
Somatom Drive | 80 | 12x1.2 | 14.4 | 12x1.2 | 1.0315 | 1.7143
Somatom Drive | 80 | 64x0.6 | 38.4 | 128x0.6 | 1.0272 | 1.699
Somatom Drive | 90 | 2x1 | 2 | 2x1 | 1.0231 | 1.6446
Somatom Drive | 90 | 6x0.6 | 3.6 | 6x0.6 | 1.0194 | 1.6493
Somatom Drive | 90 | 8x0.6 | 4.8 | 16x0.6 | 1.0192 | 1.6338
Somatom Drive | 90 | 1x5 | 5 | 1x5 | 1.0221 | 1.6443
Somatom Drive | 90 | 10x0.6 | 6 | 20x0.6 | 1.0181 | 1.6345
Somatom Drive | 90 | 1x10 | 10 | 1x10 | 1.0221 | 1.6443
Somatom Drive | 90 | 12x1.2 | 14.4 | 12x1.2 | 1.023 | 1.644
Somatom Drive | 90 | 20x0.6 | 12 | 40x0.6 | 1.0177 | 1.64
Somatom Drive | 90 | 32x0.6 | 19.2 | 64x0.6 | 1.0207 | 1.6378
Somatom Drive | 90 | 12x1.2 | 14.4 | 12x1.2 | 1.0213 | 1.6506
Somatom Drive | 90 | 64x0.6 | 38.4 | 128x0.6 | 1.0168 | 1.638
Somatom Drive | 100 | 2x1 | 2 | 2x1 | 1.0168 | 1.6008
Somatom Drive | 100 | 6x0.6 | 3.6 | 6x0.6 | 1.0222 | 1.6026
Somatom Drive | 100 | 8x0.6 | 4.8 | 16x0.6 | 1.014 | 1.5921
Somatom Drive | 100 | 1x5 | 5 | 1x5 | 1.0174 | 1.6056
Somatom Drive | 100 | 10x0.6 | 6 | 20x0.6 | 1.0091 | 1.5957
Somatom Drive | 100 | 1x10 | 10 | 1x10 | 1.0174 | 1.6056
Somatom Drive | 100 | 12x1.2 | 14.4 | 12x1.2 | 1.0177 | 1.6081
Somatom Drive | 100 | 20x0.6 | 12 | 40x0.6 | 1.0099 | 1.5903
Somatom Drive | 100 | 32x0.6 | 19.2 | 64x0.6 | 1.0151 | 1.5962
Somatom Drive | 100 | 12x1.2 | 14.4 | 12x1.2 | 1.0179 | 1.6025
Somatom Drive | 100 | 64x0.6 | 38.4 | 128x0.6 | 1.0122 | 1.594
Somatom Drive | 110 | 2x1 | 2 | 2x1 | 1.0179 | 1.5732
Somatom Drive | 110 | 6x0.6 | 3.6 | 6x0.6 | 1.0166 | 1.5744
Somatom Drive | 110 | 8x0.6 | 4.8 | 16x0.6 | 1.0104 | 1.5605
Somatom Drive | 110 | 1x5 | 5 | 1x5 | 1.0099 | 1.5748
Somatom Drive | 110 | 10x0.6 | 6 | 20x0.6 | 1.0075 | 1.559
Somatom Drive | 110 | 1x10 | 10 | 1x10 | 1.0099 | 1.5748
Somatom Drive | 110 | 12x1.2 | 14.4 | 12x1.2 | 1.0155 | 1.5756
Somatom Drive | 110 | 20x0.6 | 12 | 40x0.6 | 1.0148 | 1.5617
Somatom Drive | 110 | 32x0.6 | 19.2 | 64x0.6 | 1.008 | 1.5656
Somatom Drive | 110 | 12x1.2 | 14.4 | 12x1.2 | 1.0088 | 1.5745
Somatom Drive | 110 | 64x0.6 | 38.4 | 128x0.6 | 1.0091 | 1.5604
Somatom Drive | 120 | 2x1 | 2 | 2x1 | 1.0143 | 1.5474
Somatom Drive | 120 | 6x0.6 | 3.6 | 6x0.6 | 1.0132 | 1.5497
Somatom Drive | 120 | 8x0.6 | 4.8 | 16x0.6 | 1.0084 | 1.5353
Somatom Drive | 120 | 1x5 | 5 | 1x5 | 1.0079 | 1.5462
Somatom Drive | 120 | 10x0.6 | 6 | 20x0.6 | 1.006 | 1.5369
Somatom Drive | 120 | 1x10 | 10 | 1x10 | 1.0079 | 1.5462
Somatom Drive | 120 | 12x1.2 | 14.4 | 12x1.2 | 1.0124 | 1.5473
Somatom Drive | 120 | 20x0.6 | 12 | 40x0.6 | 1.0118 | 1.5362
Somatom Drive | 120 | 32x0.6 | 19.2 | 64x0.6 | 1.0128 | 1.5426
Somatom Drive | 120 | 12x1.2 | 14.4 | 12x1.2 | 1.0071 | 1.5472
Somatom Drive | 120 | 64x0.6 | 38.4 | 128x0.6 | 1.0072 | 1.5361
Somatom Drive | 130 | 2x1 | 2 | 2x1 | 1.0178 | 1.5312
Somatom Drive | 130 | 6x0.6 | 3.6 | 6x0.6 | 1.0146 | 1.5286
Somatom Drive | 130 | 8x0.6 | 4.8 | 16x0.6 | 1.0103 | 1.5215
Somatom Drive | 130 | 1x5 | 5 | 1x5 | 1.0132 | 1.5284
Somatom Drive | 130 | 10x0.6 | 6 | 20x0.6 | 1.01 | 1.5183
Somatom Drive | 130 | 1x10 | 10 | 1x10 | 1.0132 | 1.5284
Somatom Drive | 130 | 12x1.2 | 14.4 | 12x1.2 | 1.0154 | 1.5304
Somatom Drive | 130 | 20x0.6 | 12 | 40x0.6 | 1.0098 | 1.5195
Somatom Drive | 130 | 32x0.6 | 19.2 | 64x0.6 | 1.0106 | 1.5237
Somatom Drive | 130 | 12x1.2 | 14.4 | 12x1.2 | 1.0176 | 1.5311
Somatom Drive | 130 | 64x0.6 | 38.4 | 128x0.6 | 1.006 | 1.5172
Somatom Drive | 140 | 2x1 | 2 | 2x1 | 1.0099 | 1.5162
Somatom Drive | 140 | 6x0.6 | 3.6 | 6x0.6 | 1.0122 | 1.5143
Somatom Drive | 140 | 8x0.6 | 4.8 | 16x0.6 | 1.0086 | 1.5045
Somatom Drive | 140 | 1x5 | 5 | 1x5 | 1.011 | 1.5172
Somatom Drive | 140 | 10x0.6 | 6 | 20x0.6 | 1.0083 | 1.5026
Somatom Drive | 140 | 1x10 | 10 | 1x10 | 1.011 | 1.5172
Somatom Drive | 140 | 12x1.2 | 14.4 | 12x1.2 | 1.0085 | 1.5127
Somatom Drive | 140 | 20x0.6 | 12 | 40x0.6 | 1.0081 | 1.5013
Somatom Drive | 140 | 32x0.6 | 19.2 | 64x0.6 | 1.0133 | 1.5055
Somatom Drive | 140 | 12x1.2 | 14.4 | 12x1.2 | 1.0245 | 1.5169
Somatom Drive | 140 | 64x0.6 | 38.4 | 128x0.6 | 1.01 | 1.5047
Somatom Drive | Sn100 | 2x1 | 2 | 2x1 | 1 | 1.4259
Somatom Drive | Sn100 | 6x0.6 | 3.6 | 6x0.6 | 0.996 | 1.4368
Somatom Drive | Sn100 | 8x0.6 | 4.8 | 16x0.6 | 0.99245 | 1.4348
Somatom Drive | Sn100 | 1x5 | 5 | 1x5 | 1 | 1.4375
Somatom Drive | Sn100 | 10x0.6 | 6 | 20x0.6 | 0.98913 | 1.4219
Somatom Drive | Sn100 | 1x10 | 10 | 1x10 | 1 | 1.4375
Somatom Drive | Sn100 | 12x1.2 | 14.4 | 12x1.2 | 1 | 1.4355
Somatom Drive | Sn100 | 20x0.6 | 12 | 40x0.6 | 0.99465 | 1.4308
Somatom Drive | Sn100 | 32x0.6 | 19.2 | 64x0.6 | 0.99419 | 1.4333
Somatom Drive | Sn100 | 12x1.2 | 14.4 | 12x1.2 | 0.99355 | 1.4444
Somatom Drive | Sn100 | 64x0.6 | 38.4 | 128x0.6 | 0.99342 | 1.434
Somatom Drive | Sn140 | 2x1 | 2 | 2x1 | 1.0092 | 1.4163
Somatom Drive | Sn140 | 6x0.6 | 3.6 | 6x0.6 | 1.0095 | 1.4156
Somatom Drive | Sn140 | 8x0.6 | 4.8 | 16x0.6 | 1 | 1.4109
Somatom Drive | Sn140 | 1x5 | 5 | 1x5 | 1.0086 | 1.4182
Somatom Drive | Sn140 | 10x0.6 | 6 | 20x0.6 | 1.0065 | 1.411
Somatom Drive | Sn140 | 1x10 | 10 | 1x10 | 1.0086 | 1.4182
Somatom Drive | Sn140 | 12x1.2 | 14.4 | 12x1.2 | 1.0094 | 1.4184
Somatom Drive | Sn140 | 20x0.6 | 12 | 40x0.6 | 1.0076 | 1.4141
Somatom Drive | Sn140 | 32x0.6 | 19.2 | 64x0.6 | 1.0069 | 1.4176
Somatom Drive | Sn140 | 12x1.2 | 14.4 | 12x1.2 | 1.0092 | 1.4187
Somatom Drive | Sn140 | 64x0.6 | 38.4 | 128x0.6 | 1.0062 | 1.4132
Symbia Pro.specta Q3 | 80 | 1x5 | 5 | 1x5 | 1.1107 | 1.9412
Symbia Pro.specta Q4 | 80 | 1x10 | 10 | 1x10 | 1.1063 | 1.9412
Symbia Pro.specta Q5 | 80 | 3x3 | 9 | 3x3 | 1.1057 | 1.947
Symbia Pro.specta Q6 | 80 | 16x0.7 | 11.2 | 16x0.7 | 1.1097 | 1.9444
Symbia Pro.specta Q7 | 110 | 1x5 | 5 | 1x5 | 1.0728 | 1.7136
Symbia Pro.specta Q8 | 110 | 1x10 | 10 | 1x10 | 1.0695 | 1.7204
Symbia Pro.specta Q9 | 110 | 3x3 | 9 | 3x3 | 1.0706 | 1.7211
Symbia Pro.specta Q10 | 110 | 16x0.7 | 11.2 | 16x0.7 | 1.0701 | 1.7116
Symbia Pro.specta Q11 | Sn110 | 16x0.7 (Sn) | 11.2 | 16x0.7 (Sn) | 1.0301 | 1.5138
Symbia Pro.specta Q12 | 130 | 1x5 | 5 | 1x5 | 1.0617 | 1.6516
Symbia Pro.specta Q13 | 130 | 1x10 | 10 | 1x10 | 1.0617 | 1.6541
Symbia Pro.specta Q14 | 130 | 3x3 | 9 | 3x3 | 1.0617 | 1.6535
Symbia Pro.specta Q14 | 130 | 16x0.7 | 11.2 | 16x0.7 | 1.0602 | 1.6474
Symbia Pro.specta Q14 | Sn130 | 16x0.7 (Sn) | 11.2 | 16x0.7 (Sn) | 1.0324 | 1.4888
X.cite | 70 | 1x10 | 10 | 1x10 | 1.0759 | 1.0071
X.cite | 80 | 1x10 | 10 | 1x10 | 1.0554 | 0.99772
X.cite | 90 | 1x10 | 10 | 1x10 | 1.0451 | 0.98882
X.cite | 100 | 1x10 | 10 | 1x10 | 1.0372 | 0.98565
X.cite | 110 | 1x10 | 10 | 1x10 | 1.0327 | 0.98223
X.cite | 120 | 1x10 | 10 | 1x10 | 1.0265 | 0.98106
X.cite | 130 | 1x10 | 10 | 1x10 | 1.0262 | 0.9811
X.cite | 140 | 1x10 | 10 | 1x10 | 1.0248 | 0.98129
X.cite | 150 | 1x10 | 10 | 1x10 | 1.0227 | 0.98158
X.cite | 70 | 64x0.6 |  | 32x0.6 | 1.0738
X.cite | 80 | 64x0.6 |  | 32x0.6 | 1.0538
X.cite | 90 | 64x0.6 |  | 32x0.6 | 1.0432
X.cite | 100 | 64x0.6 |  | 32x0.6 | 1.034
X.cite | 110 | 64x0.6 |  | 32x0.6 | 1.0298
X.cite | 120 | 64x0.6 |  | 32x0.6 | 1.0242
X.cite | 130 | 64x0.6 |  | 32x0.6 | 1.0231
X.cite | 140 | 64x0.6 |  | 32x0.6 | 1.0221
X.cite | 150 | 64x0.6 |  | 32x0.6 | 1.02

## Sheet18
Model | scanner_id | manufacturer_id | phantom_diameter | shaped_filter | kv | spectral_filter | coll_N | coll_T | coll_width | ctdi100_center | ctdi_w | scanner_id | manufacturer_id
Confidence 64 | 8 | 1 | 16 | Standard | 70 |  | 2 | 1 | 2 | 2.86 | 3.06 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Standard | 70 |  | 1 | 5 | 5 | 2.57 | 2.74 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Standard | 70 |  | 10 | 0.6 | 6 | 3.41 | 3.63 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Standard | 70 |  | 6 | 1.2 | 7.2 | 3.69 | 3.94 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Standard | 70 |  | 1 | 10 | 10 | 2.57 | 2.74 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Standard | 70 |  | 20 | 0.6 | 12 | 3.48 | 3.7 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Standard | 70 |  | 12 | 1.2 | 14.4 | 3.29 | 3.52 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Standard | 70 |  | 32 | 0.6 | 19.2 | 3.24 | 3.44 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Standard | 70 |  | 16 | 1.2 | 19.2 | 3.15 | 3.37 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Standard | 80 |  | 2 | 1 | 2 | 4.57 | 4.77 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Standard | 80 |  | 1 | 5 | 5 | 4.1 | 4.28 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Standard | 80 |  | 10 | 0.6 | 6 | 5.44 | 5.67 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Standard | 80 |  | 6 | 1.2 | 7.2 | 5.89 | 6.16 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Standard | 80 |  | 1 | 10 | 10 | 4.1 | 4.28 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Standard | 80 |  | 20 | 0.6 | 12 | 5.55 | 5.78 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Standard | 80 |  | 12 | 1.2 | 14.4 | 5.25 | 5.49 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Standard | 80 |  | 32 | 0.6 | 19.2 | 5.17 | 5.38 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Standard | 80 |  | 16 | 1.2 | 19.2 | 5.03 | 5.26 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Standard | 100 |  | 2 | 1 | 2 | 8.84 | 9.02 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Standard | 100 |  | 1 | 5 | 5 | 7.93 | 8.1 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Standard | 100 |  | 10 | 0.6 | 6 | 10.53 | 10.7 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Standard | 100 |  | 6 | 1.2 | 7.2 | 11.4 | 11.64 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Standard | 100 |  | 1 | 10 | 10 | 7.93 | 8.1 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Standard | 100 |  | 20 | 0.6 | 12 | 10.73 | 10.91 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Standard | 100 |  | 12 | 1.2 | 14.4 | 10.17 | 10.39 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Standard | 100 |  | 32 | 0.6 | 19.2 | 10 | 10.16 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Standard | 100 |  | 16 | 1.2 | 19.2 | 9.73 | 9.94 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Standard | 120 |  | 2 | 1 | 2 | 14.89 | 15.07 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Standard | 120 |  | 1 | 5 | 5 | 13.36 | 13.52 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Standard | 120 |  | 10 | 0.6 | 6 | 17.69 | 17.84 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Standard | 120 |  | 6 | 1.2 | 7.2 | 19.21 | 19.43 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Standard | 120 |  | 1 | 10 | 10 | 13.36 | 13.52 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Standard | 120 |  | 20 | 0.6 | 12 | 18.03 | 18.18 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Standard | 120 |  | 12 | 1.2 | 14.4 | 17.14 | 17.34 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Standard | 120 |  | 32 | 0.6 | 19.2 | 16.8 | 16.94 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Standard | 120 |  | 16 | 1.2 | 19.2 | 16.4 | 16.59 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Standard | 140 |  | 2 | 1 | 2 | 21.22 | 21.38 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Standard | 140 |  | 1 | 5 | 5 | 19.04 | 19.18 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Standard | 140 |  | 10 | 0.6 | 6 | 25.23 | 25.35 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Standard | 140 |  | 6 | 1.2 | 7.2 | 27.37 | 27.57 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Standard | 140 |  | 1 | 10 | 10 | 19.04 | 19.18 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Standard | 140 |  | 20 | 0.6 | 12 | 25.72 | 25.83 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Standard | 140 |  | 12 | 1.2 | 14.4 | 24.42 | 24.6 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Standard | 140 |  | 32 | 0.6 | 19.2 | 23.96 | 24.07 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Standard | 140 |  | 16 | 1.2 | 19.2 | 23.37 | 23.54 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Narrow | 70 |  | 2 | 1 | 2 | 2.54 | 2.55 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Narrow | 70 |  | 1 | 5 | 5 | 2.28 | 2.28 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Narrow | 70 |  | 10 | 0.6 | 6 | 3.02 | 3.01 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Narrow | 70 |  | 6 | 1.2 | 7.2 | 3.28 | 3.28 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Narrow | 70 |  | 1 | 10 | 10 | 2.28 | 2.28 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Narrow | 70 |  | 20 | 0.6 | 12 | 3.08 | 3.07 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Narrow | 70 |  | 12 | 1.2 | 14.4 | 2.93 | 2.93 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Narrow | 70 |  | 32 | 0.6 | 19.2 | 2.87 | 2.86 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Narrow | 70 |  | 16 | 1.2 | 19.2 | 2.8 | 2.8 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Narrow | 80 |  | 2 | 1 | 2 | 4.11 | 4.05 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Narrow | 80 |  | 1 | 5 | 5 | 3.69 | 3.64 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Narrow | 80 |  | 10 | 0.6 | 6 | 4.91 | 4.8 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Narrow | 80 |  | 6 | 1.2 | 7.2 | 5.3 | 5.23 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Narrow | 80 |  | 1 | 10 | 10 | 3.69 | 3.64 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Narrow | 80 |  | 20 | 0.6 | 12 | 5 | 4.89 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Narrow | 80 |  | 12 | 1.2 | 14.4 | 4.73 | 4.66 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Narrow | 80 |  | 32 | 0.6 | 19.2 | 4.66 | 4.56 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Narrow | 80 |  | 16 | 1.2 | 19.2 | 4.53 | 5.26 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Narrow | 100 |  | 2 | 1 | 2 | 8.04 | 7.82 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Narrow | 100 |  | 1 | 5 | 5 | 7.21 | 7.02 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Narrow | 100 |  | 10 | 0.6 | 6 | 9.58 | 9.26 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Narrow | 100 |  | 6 | 1.2 | 7.2 | 10.37 | 10.09 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Narrow | 100 |  | 1 | 10 | 10 | 7.21 | 7.02 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Narrow | 100 |  | 20 | 0.6 | 12 | 9.76 | 9.43 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Narrow | 100 |  | 12 | 1.2 | 14.4 | 9.25 | 9 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Narrow | 100 |  | 32 | 0.6 | 19.2 | 9.09 | 8.79 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Narrow | 100 |  | 16 | 1.2 | 19.2 | 8.85 | 8.61 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Narrow | 120 |  | 2 | 1 | 2 | 11.68 | 13.28 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Narrow | 120 |  | 1 | 5 | 5 | 12.28 | 11.92 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Narrow | 120 |  | 10 | 0.6 | 6 | 16.3 | 15.71 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Narrow | 120 |  | 6 | 1.2 | 7.2 | 17.65 | 17.13 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Narrow | 120 |  | 1 | 10 | 10 | 12.28 | 11.92 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Narrow | 120 |  | 20 | 0.6 | 12 | 16.61 | 16.01 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Narrow | 120 |  | 12 | 1.2 | 14.4 | 15.75 | 15.28 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Narrow | 120 |  | 32 | 0.6 | 19.2 | 15.48 | 14.92 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Narrow | 120 |  | 16 | 1.2 | 19.2 | 15.07 | 14.63 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Narrow | 140 |  | 2 | 1 | 2 | 19.6 | 19.01 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Narrow | 140 |  | 1 | 5 | 5 | 17.58 | 17.05 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Narrow | 140 |  | 10 | 0.6 | 6 | 23.34 | 22.49 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Narrow | 140 |  | 6 | 1.2 | 7.2 | 25.28 | 24.51 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Narrow | 140 |  | 1 | 10 | 10 | 17.58 | 17.05 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Narrow | 140 |  | 20 | 0.6 | 12 | 23.79 | 22.92 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Narrow | 140 |  | 12 | 1.2 | 14.4 | 22.56 | 21.87 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Narrow | 140 |  | 32 | 0.6 | 19.2 | 22.16 | 21.35 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 16 | Narrow | 140 |  | 16 | 1.2 | 19.2 | 21.58 | 20.93 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Standard | 70 |  | 2 | 1 | 2 | 0.68 | 1.37 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Standard | 70 |  | 1 | 5 | 5 | 0.61 | 1.22 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Standard | 70 |  | 10 | 0.6 | 6 | 0.82 | 1.61 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Standard | 70 |  | 6 | 1.2 | 7.2 | 0.88 | 1.76 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Standard | 70 |  | 1 | 10 | 10 | 0.61 | 1.22 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Standard | 70 |  | 20 | 0.6 | 12 | 0.83 | 1.64 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Standard | 70 |  | 12 | 1.2 | 14.4 | 0.79 | 1.57 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Standard | 70 |  | 32 | 0.6 | 19.2 | 0.78 | 1.52 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Standard | 70 |  | 16 | 1.2 | 19.2 | 0.75 | 1.59 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Standard | 80 |  | 2 | 1 | 2 | 1.18 | 2.18 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Standard | 80 |  | 1 | 5 | 5 | 1.06 | 1.96 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Standard | 80 |  | 10 | 0.6 | 6 | 1.4 | 2.58 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Standard | 80 |  | 6 | 1.2 | 7.2 | 1.52 | 2.81 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Standard | 80 |  | 1 | 10 | 10 | 1.06 | 1.96 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Standard | 80 |  | 20 | 0.6 | 12 | 1.43 | 2.62 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Standard | 80 |  | 12 | 1.2 | 14.4 | 1.36 | 2.51 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Standard | 80 |  | 32 | 0.6 | 19.2 | 1.33 | 2.45 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Standard | 80 |  | 16 | 1.2 | 19.2 | 1.3 | 2.4 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Standard | 100 |  | 2 | 1 | 2 | 2.56 | 4.28 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Standard | 100 |  | 1 | 5 | 5 | 2.29 | 3.84 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Standard | 100 |  | 10 | 0.6 | 6 | 3.04 | 5.05 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Standard | 100 |  | 6 | 1.2 | 7.2 | 3.3 | 5.52 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Standard | 100 |  | 1 | 10 | 10 | 2.29 | 3.84 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Standard | 100 |  | 20 | 0.6 | 12 | 3.1 | 5.15 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Standard | 100 |  | 12 | 1.2 | 14.4 | 2.94 | 4.92 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Standard | 100 |  | 32 | 0.6 | 19.2 | 2.89 | 4.8 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Standard | 100 |  | 16 | 1.2 | 19.2 | 2.82 | 4.71 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Standard | 120 |  | 2 | 1 | 2 | 4.61 | 7.33 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Standard | 120 |  | 1 | 5 | 5 | 4.14 | 6.58 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Standard | 120 |  | 10 | 0.6 | 6 | 5.48 | 8.68 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Standard | 120 |  | 6 | 1.2 | 7.2 | 5.95 | 9.46 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Standard | 120 |  | 1 | 10 | 10 | 4.14 | 6.58 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Standard | 120 |  | 20 | 0.6 | 12 | 5.59 | 8.85 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Standard | 120 |  | 12 | 1.2 | 14.4 | 5.31 | 8.44 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Standard | 120 |  | 32 | 0.6 | 19.2 | 5.21 | 8.25 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Standard | 120 |  | 16 | 1.2 | 19.2 | 5.08 | 8.08 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Standard | 140 |  | 2 | 1 | 2 | 6.88 | 10.59 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Standard | 140 |  | 1 | 5 | 5 | 6.17 | 9.5 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Standard | 140 |  | 10 | 0.6 | 6 | 8.2 | 12.51 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Standard | 140 |  | 6 | 1.2 | 7.2 | 8.87 | 13.66 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Standard | 140 |  | 1 | 10 | 10 | 6.17 | 9.5 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Standard | 140 |  | 20 | 0.6 | 12 | 8.36 | 12.75 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Standard | 140 |  | 12 | 1.2 | 14.4 | 7.92 | 12.19 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Standard | 140 |  | 32 | 0.6 | 19.2 | 7.79 | 11.88 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Standard | 140 |  | 16 | 1.2 | 19.2 | 7.58 | 11.66 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Narrow | 70 |  | 2 | 1 | 2 | 0.58 | 1.01 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Narrow | 70 |  | 1 | 5 | 5 | 0.52 | 0.91 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Narrow | 70 |  | 10 | 0.6 | 6 | 0.69 | 1.19 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Narrow | 70 |  | 6 | 1.2 | 7.2 | 0.75 | 1.3 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Narrow | 70 |  | 1 | 10 | 10 | 0.52 | 0.91 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Narrow | 70 |  | 20 | 0.6 | 12 | 0.7 | 1.22 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Narrow | 70 |  | 12 | 1.2 | 14.4 | 0.67 | 1.16 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Narrow | 70 |  | 32 | 0.6 | 19.2 | 0.66 | 1.13 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Narrow | 70 |  | 16 | 1.2 | 19.2 | 0.64 | 1.11 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Narrow | 80 |  | 2 | 1 | 2 | 1.03 | 1.66 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Narrow | 80 |  | 1 | 5 | 5 | 0.92 | 1.49 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Narrow | 80 |  | 10 | 0.6 | 6 | 1.22 | 1.96 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Narrow | 80 |  | 6 | 1.2 | 7.2 | 1.33 | 2.14 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Narrow | 80 |  | 1 | 10 | 10 | 0.92 | 1.49 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Narrow | 80 |  | 20 | 0.6 | 12 | 1.24 | 2 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Narrow | 80 |  | 12 | 1.2 | 14.4 | 1.18 | 1.91 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Narrow | 80 |  | 32 | 0.6 | 19.2 | 1.16 | 1.86 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Narrow | 80 |  | 16 | 1.2 | 19.2 | 1.13 | 1.83 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Narrow | 100 |  | 2 | 1 | 2 | 2.25 | 3.35 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Narrow | 100 |  | 1 | 5 | 5 | 2.02 | 3.01 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Narrow | 100 |  | 10 | 0.6 | 6 | 2.67 | 3.95 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Narrow | 100 |  | 6 | 1.2 | 7.2 | 2.9 | 4.32 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Narrow | 100 |  | 1 | 10 | 10 | 2.02 | 3.01 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Narrow | 100 |  | 20 | 0.6 | 12 | 2.72 | 4.03 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Narrow | 100 |  | 12 | 1.2 | 14.4 | 2.59 | 3.86 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Narrow | 100 |  | 32 | 0.6 | 19.2 | 2.54 | 3.75 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Narrow | 100 |  | 16 | 1.2 | 19.2 | 2.47 | 3.69 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Narrow | 120 |  | 2 | 1 | 2 | 4.09 | 5.86 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Narrow | 120 |  | 1 | 5 | 5 | 3.67 | 5.26 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Narrow | 120 |  | 10 | 0.6 | 6 | 4.87 | 6.89 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Narrow | 120 |  | 6 | 1.2 | 7.2 | 5.28 | 7.56 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Narrow | 120 |  | 1 | 10 | 10 | 3.67 | 5.26 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Narrow | 120 |  | 20 | 0.6 | 12 | 4.96 | 7.02 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Narrow | 120 |  | 12 | 1.2 | 14.4 | 4.71 | 6.75 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Narrow | 120 |  | 32 | 0.6 | 19.2 | 4.62 | 6.54 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Narrow | 120 |  | 16 | 1.2 | 19.2 | 4.51 | 6.46 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Narrow | 140 |  | 2 | 1 | 2 | 6.16 | 8.6 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Narrow | 140 |  | 1 | 5 | 5 | 5.52 | 7.71 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Narrow | 140 |  | 10 | 0.6 | 6 | 7.29 | 10.12 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Narrow | 140 |  | 6 | 1.2 | 7.2 | 7.94 | 11.09 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Narrow | 140 |  | 1 | 10 | 10 | 5.52 | 7.71 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Narrow | 140 |  | 20 | 0.6 | 12 | 7.43 | 10.32 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Narrow | 140 |  | 12 | 1.2 | 14.4 | 7.09 | 9.89 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Narrow | 140 |  | 32 | 0.6 | 19.2 | 6.92 | 9.61 | Confidence 64 | Siemens
Confidence 64 | 8 | 1 | 32 | Narrow | 140 |  | 16 | 1.2 | 19.2 | 6.78 | 9.47 | Confidence 64 | Siemens
Confidence 20 | 9 | 1 | 16 | Standard | 70 |  | 2 | 1 | 2 | 2.86 | 3.06 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Standard | 70 |  | 1 | 5 | 5 | 2.57 | 2.74 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Standard | 70 |  | 10 | 0.6 | 6 | 3.54 | 3.79 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Standard | 70 |  | 6 | 1.2 | 7.2 | 3.69 | 3.94 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Standard | 70 |  | 1 | 10 | 10 | 2.57 | 2.74 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Standard | 70 |  | 20 | 0.6 | 12 | 3.48 | 3.71 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Standard | 70 |  | 12 | 1.2 | 14.4 | 3.29 | 3.52 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Standard | 70 |  | 16 | 1.2 | 19.2 | 3.15 | 3.37 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Standard | 80 |  | 2 | 1 | 2 | 4.57 | 4.77 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Standard | 80 |  | 1 | 5 | 5 | 4.1 | 4.28 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Standard | 80 |  | 10 | 0.6 | 6 | 5.65 | 5.91 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Standard | 80 |  | 6 | 1.2 | 7.2 | 5.89 | 6.16 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Standard | 80 |  | 1 | 10 | 10 | 4.1 | 4.28 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Standard | 80 |  | 20 | 0.6 | 12 | 5.54 | 5.8 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Standard | 80 |  | 12 | 1.2 | 14.4 | 5.25 | 5.49 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Standard | 80 |  | 16 | 1.2 | 19.2 | 5.03 | 5.26 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Standard | 100 |  | 2 | 1 | 2 | 8.84 | 9.02 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Standard | 100 |  | 1 | 5 | 5 | 7.93 | 8.1 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Standard | 100 |  | 10 | 0.6 | 6 | 10.94 | 11.17 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Standard | 100 |  | 6 | 1.2 | 7.2 | 11.4 | 11.64 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Standard | 100 |  | 1 | 10 | 10 | 7.93 | 8.1 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Standard | 100 |  | 20 | 0.6 | 12 | 10.73 | 10.96 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Standard | 100 |  | 12 | 1.2 | 14.4 | 10.17 | 10.39 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Standard | 100 |  | 16 | 1.2 | 19.2 | 9.73 | 9.94 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Standard | 120 |  | 2 | 1 | 2 | 14.89 | 15.07 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Standard | 120 |  | 1 | 5 | 5 | 13.36 | 13.52 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Standard | 120 |  | 10 | 0.6 | 6 | 18.44 | 18.66 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Standard | 120 |  | 6 | 1.2 | 7.2 | 19.21 | 19.43 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Standard | 120 |  | 1 | 10 | 10 | 13.36 | 13.52 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Standard | 120 |  | 20 | 0.6 | 12 | 18.08 | 18.3 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Standard | 120 |  | 12 | 1.2 | 14.4 | 17.14 | 17.34 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Standard | 120 |  | 16 | 1.2 | 19.2 | 16.4 | 16.59 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Standard | 140 |  | 2 | 1 | 2 | 21.22 | 21.38 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Standard | 140 |  | 1 | 5 | 5 | 19.04 | 19.18 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Standard | 140 |  | 10 | 0.6 | 6 | 26.27 | 26.47 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Standard | 140 |  | 6 | 1.2 | 7.2 | 27.37 | 27.57 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Standard | 140 |  | 1 | 10 | 10 | 19.04 | 19.18 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Standard | 140 |  | 20 | 0.6 | 12 | 25.77 | 15.96 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Standard | 140 |  | 12 | 1.2 | 14.4 | 24.42 | 24.6 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Standard | 140 |  | 16 | 1.2 | 19.2 | 23.37 | 23.54 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Narrow | 70 |  | 2 | 1 | 2 | 2.54 | 2.55 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Narrow | 70 |  | 1 | 5 | 5 | 2.28 | 2.28 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Narrow | 70 |  | 10 | 0.6 | 6 | 3.15 | 3.15 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Narrow | 70 |  | 6 | 1.2 | 7.2 | 3.28 | 3.28 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Narrow | 70 |  | 1 | 10 | 10 | 2.28 | 2.28 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Narrow | 70 |  | 20 | 0.6 | 12 | 3.09 | 3.09 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Narrow | 70 |  | 12 | 1.2 | 14.4 | 2.93 | 2.93 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Narrow | 70 |  | 16 | 1.2 | 19.2 | 2.8 | 2.8 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Narrow | 80 |  | 2 | 1 | 2 | 4.11 | 4.05 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Narrow | 80 |  | 1 | 5 | 5 | 3.69 | 3.64 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Narrow | 80 |  | 10 | 0.6 | 6 | 5.09 | 5.02 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Narrow | 80 |  | 6 | 1.2 | 7.2 | 5.3 | 5.23 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Narrow | 80 |  | 1 | 10 | 10 | 3.69 | 3.64 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Narrow | 80 |  | 20 | 0.6 | 12 | 4.99 | 4.92 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Narrow | 80 |  | 12 | 1.2 | 14.4 | 4.73 | 4.66 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Narrow | 80 |  | 16 | 1.2 | 19.2 | 4.53 | 5.26 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Narrow | 100 |  | 2 | 1 | 2 | 8.04 | 7.82 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Narrow | 100 |  | 1 | 5 | 5 | 7.21 | 7.02 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Narrow | 100 |  | 10 | 0.6 | 6 | 9.95 | 9.68 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Narrow | 100 |  | 6 | 1.2 | 7.2 | 10.37 | 10.09 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Narrow | 100 |  | 1 | 10 | 10 | 7.21 | 7.02 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Narrow | 100 |  | 20 | 0.6 | 12 | 9.76 | 9.5 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Narrow | 100 |  | 12 | 1.2 | 14.4 | 9.25 | 9 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Narrow | 100 |  | 16 | 1.2 | 19.2 | 8.85 | 8.61 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Narrow | 120 |  | 2 | 1 | 2 | 11.68 | 13.28 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Narrow | 120 |  | 1 | 5 | 5 | 12.28 | 11.92 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Narrow | 120 |  | 10 | 0.6 | 6 | 16.94 | 16.44 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Narrow | 120 |  | 6 | 1.2 | 7.2 | 17.65 | 17.13 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Narrow | 120 |  | 1 | 10 | 10 | 12.28 | 11.92 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Narrow | 120 |  | 20 | 0.6 | 12 | 16.62 | 16.13 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Narrow | 120 |  | 12 | 1.2 | 14.4 | 15.75 | 15.28 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Narrow | 120 |  | 16 | 1.2 | 19.2 | 15.07 | 14.63 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Narrow | 140 |  | 2 | 1 | 2 | 19.6 | 19.01 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Narrow | 140 |  | 1 | 5 | 5 | 17.58 | 17.05 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Narrow | 140 |  | 10 | 0.6 | 6 | 24.27 | 23.53 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Narrow | 140 |  | 6 | 1.2 | 7.2 | 25.28 | 24.51 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Narrow | 140 |  | 1 | 10 | 10 | 17.58 | 17.05 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Narrow | 140 |  | 20 | 0.6 | 12 | 23.8 | 23.08 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Narrow | 140 |  | 12 | 1.2 | 14.4 | 22.56 | 21.87 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 16 | Narrow | 140 |  | 16 | 1.2 | 19.2 | 21.58 | 20.93 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Standard | 70 |  | 2 | 1 | 2 | 0.68 | 1.37 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Standard | 70 |  | 1 | 5 | 5 | 0.61 | 1.22 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Standard | 70 |  | 10 | 0.6 | 6 | 0.84 | 1.69 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Standard | 70 |  | 6 | 1.2 | 7.2 | 0.88 | 1.76 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Standard | 70 |  | 1 | 10 | 10 | 0.61 | 1.22 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Standard | 70 |  | 20 | 0.6 | 12 | 0.83 | 1.66 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Standard | 70 |  | 12 | 1.2 | 14.4 | 0.79 | 1.57 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Standard | 70 |  | 16 | 1.2 | 19.2 | 0.75 | 1.59 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Standard | 80 |  | 2 | 1 | 2 | 1.18 | 2.18 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Standard | 80 |  | 1 | 5 | 5 | 1.06 | 1.96 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Standard | 80 |  | 10 | 0.6 | 6 | 1.46 | 2.7 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Standard | 80 |  | 6 | 1.2 | 7.2 | 1.52 | 2.81 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Standard | 80 |  | 1 | 10 | 10 | 1.06 | 1.96 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Standard | 80 |  | 20 | 0.6 | 12 | 1.43 | 2.65 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Standard | 80 |  | 12 | 1.2 | 14.4 | 1.36 | 2.51 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Standard | 80 |  | 16 | 1.2 | 19.2 | 1.3 | 2.4 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Standard | 100 |  | 2 | 1 | 2 | 2.56 | 4.28 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Standard | 100 |  | 1 | 5 | 5 | 2.29 | 3.84 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Standard | 100 |  | 10 | 0.6 | 6 | 3.17 | 5.3 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Standard | 100 |  | 6 | 1.2 | 7.2 | 3.3 | 5.52 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Standard | 100 |  | 1 | 10 | 10 | 2.29 | 3.84 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Standard | 100 |  | 20 | 0.6 | 12 | 3.1 | 5.19 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Standard | 100 |  | 12 | 1.2 | 14.4 | 2.94 | 4.92 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Standard | 100 |  | 16 | 1.2 | 19.2 | 2.82 | 4.71 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Standard | 120 |  | 2 | 1 | 2 | 4.61 | 7.33 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Standard | 120 |  | 1 | 5 | 5 | 4.14 | 6.58 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Standard | 120 |  | 10 | 0.6 | 6 | 5.71 | 9.08 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Standard | 120 |  | 6 | 1.2 | 7.2 | 5.95 | 9.46 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Standard | 120 |  | 1 | 10 | 10 | 4.14 | 6.58 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Standard | 120 |  | 20 | 0.6 | 12 | 5.6 | 8.91 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Standard | 120 |  | 12 | 1.2 | 14.4 | 5.31 | 8.44 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Standard | 120 |  | 16 | 1.2 | 19.2 | 5.08 | 8.08 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Standard | 140 |  | 2 | 1 | 2 | 6.88 | 10.59 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Standard | 140 |  | 1 | 5 | 5 | 6.17 | 9.5 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Standard | 140 |  | 10 | 0.6 | 6 | 8.52 | 13.11 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Standard | 140 |  | 6 | 1.2 | 7.2 | 8.87 | 13.66 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Standard | 140 |  | 1 | 10 | 10 | 6.17 | 9.5 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Standard | 140 |  | 20 | 0.6 | 12 | 8.36 | 12.86 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Standard | 140 |  | 12 | 1.2 | 14.4 | 7.92 | 12.19 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Standard | 140 |  | 16 | 1.2 | 19.2 | 7.58 | 11.66 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Narrow | 70 |  | 2 | 1 | 2 | 0.58 | 1.01 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Narrow | 70 |  | 1 | 5 | 5 | 0.52 | 0.91 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Narrow | 70 |  | 10 | 0.6 | 6 | 0.72 | 1.25 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Narrow | 70 |  | 6 | 1.2 | 7.2 | 0.75 | 1.3 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Narrow | 70 |  | 1 | 10 | 10 | 0.52 | 0.91 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Narrow | 70 |  | 20 | 0.6 | 12 | 0.7 | 1.23 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Narrow | 70 |  | 12 | 1.2 | 14.4 | 0.67 | 1.16 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Narrow | 70 |  | 16 | 1.2 | 19.2 | 0.64 | 1.11 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Narrow | 80 |  | 2 | 1 | 2 | 1.03 | 1.66 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Narrow | 80 |  | 1 | 5 | 5 | 0.92 | 1.49 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Narrow | 80 |  | 10 | 0.6 | 6 | 1.27 | 2.06 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Narrow | 80 |  | 6 | 1.2 | 7.2 | 1.33 | 2.14 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Narrow | 80 |  | 1 | 10 | 10 | 0.92 | 1.49 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Narrow | 80 |  | 20 | 0.6 | 12 | 1.25 | 2.02 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Narrow | 80 |  | 12 | 1.2 | 14.4 | 1.18 | 1.91 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Narrow | 80 |  | 16 | 1.2 | 19.2 | 1.13 | 1.83 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Narrow | 100 |  | 2 | 1 | 2 | 2.25 | 3.35 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Narrow | 100 |  | 1 | 5 | 5 | 2.02 | 3.01 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Narrow | 100 |  | 10 | 0.6 | 6 | 2.78 | 4.15 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Narrow | 100 |  | 6 | 1.2 | 7.2 | 2.9 | 4.32 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Narrow | 100 |  | 1 | 10 | 10 | 2.02 | 3.01 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Narrow | 100 |  | 20 | 0.6 | 12 | 2.73 | 4.07 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Narrow | 100 |  | 12 | 1.2 | 14.4 | 2.59 | 3.86 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Narrow | 100 |  | 16 | 1.2 | 19.2 | 2.47 | 3.69 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Narrow | 120 |  | 2 | 1 | 2 | 4.09 | 5.86 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Narrow | 120 |  | 1 | 5 | 5 | 3.67 | 5.26 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Narrow | 120 |  | 10 | 0.6 | 6 | 5.07 | 7.26 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Narrow | 120 |  | 6 | 1.2 | 7.2 | 5.28 | 7.56 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Narrow | 120 |  | 1 | 10 | 10 | 3.67 | 5.26 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Narrow | 120 |  | 20 | 0.6 | 12 | 4.97 | 7.12 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Narrow | 120 |  | 12 | 1.2 | 14.4 | 4.71 | 6.75 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Narrow | 120 |  | 16 | 1.2 | 19.2 | 4.51 | 6.46 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Narrow | 140 |  | 2 | 1 | 2 | 6.16 | 8.6 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Narrow | 140 |  | 1 | 5 | 5 | 5.52 | 7.71 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Narrow | 140 |  | 10 | 0.6 | 6 | 7.62 | 10.64 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Narrow | 140 |  | 6 | 1.2 | 7.2 | 7.94 | 11.09 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Narrow | 140 |  | 1 | 10 | 10 | 5.52 | 7.71 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Narrow | 140 |  | 20 | 0.6 | 12 | 7.48 | 10.44 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Narrow | 140 |  | 12 | 1.2 | 14.4 | 7.09 | 9.89 | Confidence 20 | Siemens
Confidence 20 | 9 | 1 | 32 | Narrow | 140 |  | 16 | 1.2 | 19.2 | 6.78 | 9.47 | Confidence 20 | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 70 |  | 1 | 5 | 5 | 4.07 | 4.39 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 70 |  | 1 | 10 | 10 | 4.01 | 4.32 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 70 |  | 3 | 3 | 9 | 4.45 | 4.8 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 70 |  | 3 | 5 | 15 | 4.04 | 4.35 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 70 |  | 32 | 0.6 | 19.2 | 4.73 | 5.1 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 70 |  | 64 | 0.6 | 38.4 | 4.18 | 4.52 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 80 |  | 1 | 5 | 5 | 6.44 | 6.81 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 80 |  | 1 | 10 | 10 | 6.34 | 6.69 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 80 |  | 3 | 3 | 9 | 7.04 | 7.43 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 80 |  | 3 | 5 | 15 | 6.39 | 6.74 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 80 |  | 32 | 0.6 | 19.2 | 7.49 | 7.9 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 80 |  | 64 | 0.6 | 38.4 | 6.61 | 7 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 90 |  | 1 | 5 | 5 | 9.29 | 9.69 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 90 |  | 1 | 10 | 10 | 9.13 | 9.53 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 90 |  | 3 | 3 | 9 | 10.14 | 10.59 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 90 |  | 3 | 5 | 15 | 9.2 | 9.6 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 90 |  | 32 | 0.6 | 19.2 | 10.77 | 11.25 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 90 |  | 64 | 0.6 | 38.4 | 9.51 | 9.96 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 100 |  | 1 | 5 | 5 | 12.52 | 12.96 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 100 |  | 1 | 10 | 10 | 12.31 | 12.75 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 100 |  | 3 | 3 | 9 | 13.68 | 14.17 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 100 |  | 3 | 5 | 15 | 12.42 | 12.85 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 100 |  | 32 | 0.6 | 19.2 | 14.53 | 15.04 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 100 |  | 64 | 0.6 | 38.4 | 12.81 | 13.3 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 110 |  | 1 | 5 | 5 | 15.99 | 16.57 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 110 |  | 1 | 10 | 10 | 15.79 | 16.3 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 110 |  | 3 | 3 | 9 | 17.54 | 18.11 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 110 |  | 3 | 5 | 15 | 15.9 | 16.4 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 110 |  | 32 | 0.6 | 19.2 | 18.64 | 19.21 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 110 |  | 64 | 0.6 | 38.4 | 16.42 | 16.98 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 120 |  | 1 | 5 | 5 | 19.92 | 20.5 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 120 |  | 1 | 10 | 10 | 19.6 | 20.13 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 120 |  | 3 | 3 | 9 | 21.77 | 22.37 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 120 |  | 3 | 5 | 15 | 19.71 | 20.26 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 120 |  | 32 | 0.6 | 19.2 | 23.08 | 23.73 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 120 |  | 64 | 0.6 | 38.4 | 20.34 | 20.96 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 130 |  | 1 | 5 | 5 | 24.1 | 24.72 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 130 |  | 1 | 10 | 10 | 23.69 | 24.26 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 130 |  | 3 | 3 | 9 | 26.32 | 26.95 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 130 |  | 3 | 5 | 15 | 23.81 | 24.41 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 130 |  | 32 | 0.6 | 19.2 | 27.89 | 28.57 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 130 |  | 64 | 0.6 | 38.4 | 24.52 | 25.22 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 140 |  | 1 | 5 | 5 | 28.32 | 29.08 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 140 |  | 1 | 10 | 10 | 27.94 | 28.6 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 140 |  | 3 | 3 | 9 | 31.04 | 31.78 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 140 |  | 3 | 5 | 15 | 28.13 | 28.77 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 140 |  | 32 | 0.6 | 19.2 | 23.89 | 33.69 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 140 |  | 64 | 0.6 | 38.4 | 29 | 29.75 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 100 | Sn | 1 | 5 | 5 |  |  | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 100 | Sn | 1 | 10 | 10 |  |  | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 100 | Sn | 3 | 3 | 9 | 2.11 | 2.13 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 100 | Sn | 3 | 5 | 15 | 1.92 | 1.93 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 100 | Sn | 32 | 0.6 | 19.2 |  |  | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 100 | Sn | 64 | 0.6 | 38.4 | 2 | 2.02 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 110 | Sn | 1 | 5 | 5 |  |  | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 110 | Sn | 1 | 10 | 10 |  |  | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 110 | Sn | 3 | 3 | 9 | 3.26 | 3.3 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 110 | Sn | 3 | 5 | 15 | 2.97 | 3 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 110 | Sn | 32 | 0.6 | 19.2 |  |  | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 110 | Sn | 64 | 0.6 | 38.4 | 3.1 | 3.14 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 120 | Sn | 1 | 5 | 5 |  |  | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 120 | Sn | 1 | 10 | 10 |  |  | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 120 | Sn | 3 | 3 | 9 | 4.72 | 4.77 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 120 | Sn | 3 | 5 | 15 | 4.29 | 4.33 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 120 | Sn | 32 | 0.6 | 19.2 |  |  | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 120 | Sn | 64 | 0.6 | 38.4 | 4.47 | 4.53 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 130 | Sn | 1 | 5 | 5 |  |  | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 130 | Sn | 1 | 10 | 10 |  |  | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 130 | Sn | 3 | 3 | 9 | 6.47 | 6.55 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 130 | Sn | 3 | 5 | 15 | 5.87 | 5.95 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 130 | Sn | 32 | 0.6 | 19.2 |  |  | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 130 | Sn | 64 | 0.6 | 38.4 | 6.11 | 6.21 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 140 | Sn | 1 | 5 | 5 |  |  | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 140 | Sn | 1 | 10 | 10 |  |  | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 140 | Sn | 3 | 3 | 9 | 8.5 | 8.63 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 140 | Sn | 3 | 5 | 15 | 7.73 | 7.83 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 140 | Sn | 32 | 0.6 | 19.2 |  |  | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Standard | 140 | Sn | 64 | 0.6 | 38.4 | 8.04 | 8.19 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 70 |  | 1 | 5 | 5 | 3.64 | 3.68 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 70 |  | 1 | 10 | 10 | 3.61 | 3.63 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 70 |  | 3 | 3 | 9 | 4.01 | 4.03 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 70 |  | 3 | 5 | 15 | 3.63 | 3.65 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 70 |  | 32 | 0.6 | 19.2 | 4.24 | 4.27 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 70 |  | 64 | 0.6 | 38.4 | 3.74 | 3.77 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 80 |  | 1 | 5 | 5 | 5.83 | 5.8 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 80 |  | 1 | 10 | 10 | 5.76 | 5.72 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 80 |  | 3 | 3 | 9 | 6.39 | 6.36 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 80 |  | 3 | 5 | 15 | 5.8 | 5.76 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 80 |  | 32 | 0.6 | 19.2 | 6.79 | 6.74 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 80 |  | 64 | 0.6 | 38.4 | 5.98 | 5.95 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 90 |  | 1 | 5 | 5 | 8.5 | 8.39 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 90 |  | 1 | 10 | 10 | 8.35 | 8.25 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 90 |  | 3 | 3 | 9 | 9.28 | 9.17 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 90 |  | 3 | 5 | 15 | 8.41 | 8.3 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 90 |  | 32 | 0.6 | 19.2 | 9.84 | 9.72 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 90 |  | 64 | 0.6 | 38.4 | 8.66 | 8.58 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 100 |  | 1 | 5 | 5 | 11.51 | 11.33 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 100 |  | 1 | 10 | 10 | 11.35 | 11.16 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 100 |  | 3 | 3 | 9 | 12.61 | 12.4 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 100 |  | 3 | 5 | 15 | 11.4 | 11.2 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 100 |  | 32 | 0.6 | 19.2 | 13.35 | 13.12 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 100 |  | 64 | 0.6 | 38.4 | 11.73 | 11.57 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 110 |  | 1 | 5 | 5 | 16.9 | 16.62 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 110 |  | 1 | 10 | 10 | 16.65 | 14.38 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 110 |  | 3 | 3 | 9 | 16.28 | 15.98 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 110 |  | 3 | 5 | 15 | 14.72 | 14.44 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 110 |  | 32 | 0.6 | 19.2 | 17.19 | 16.88 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 110 |  | 64 | 0.6 | 38.4 | 15.1 | 14.88 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 120 |  | 1 | 5 | 5 | 18.54 | 18.19 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 120 |  | 1 | 10 | 10 | 18.2 | 17.86 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 120 |  | 3 | 3 | 9 | 20.22 | 19.84 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 120 |  | 3 | 5 | 15 | 18.27 | 17.92 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 120 |  | 32 | 0.6 | 19.2 | 21.41 | 20.99 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 120 |  | 64 | 0.6 | 38.4 | 18.8 | 18.49 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 130 |  | 1 | 5 | 5 | 22.44 | 22.01 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 130 |  | 1 | 10 | 10 | 22.05 | 21.61 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 130 |  | 3 | 3 | 9 | 24.49 | 24.01 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 130 |  | 3 | 5 | 15 | 22.15 | 21.7 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 130 |  | 32 | 0.6 | 19.2 | 25.92 | 25.41 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 130 |  | 64 | 0.6 | 38.4 | 22.75 | 22.37 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 140 |  | 1 | 5 | 5 | 26.59 | 26.09 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 140 |  | 1 | 10 | 10 | 26.16 | 25.66 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 140 |  | 3 | 3 | 9 | 29.06 | 28.53 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 140 |  | 3 | 5 | 15 | 26.26 | 25.74 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 140 |  | 32 | 0.6 | 19.2 | 30.73 | 30.1 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 140 |  | 64 | 0.6 | 38.4 | 26.99 | 26.53 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 100 | Sn | 1 | 5 | 5 |  |  | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 100 | Sn | 1 | 10 | 10 |  |  | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 100 | Sn | 3 | 3 | 9 | 1.99 | 1.94 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 100 | Sn | 3 | 5 | 15 | 1.8 | 1.76 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 100 | Sn | 32 | 0.6 | 19.2 |  |  | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 100 | Sn | 64 | 0.6 | 38.4 | 1.88 | 1.84 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 110 | Sn | 1 | 5 | 5 |  |  | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 110 | Sn | 1 | 10 | 10 |  |  | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 110 | Sn | 3 | 3 | 9 | 3.1 | 3.03 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 110 | Sn | 3 | 5 | 15 | 2.81 | 2.74 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 110 | Sn | 32 | 0.6 | 19.2 |  |  | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 110 | Sn | 64 | 0.6 | 38.4 | 2.92 | 2.86 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 120 | Sn | 1 | 5 | 5 |  |  | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 120 | Sn | 1 | 10 | 10 |  |  | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 120 | Sn | 3 | 3 | 9 | 4.48 | 4.39 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 120 | Sn | 3 | 5 | 15 | 4.06 | 3.98 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 120 | Sn | 32 | 0.6 | 19.2 |  |  | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 120 | Sn | 64 | 0.6 | 38.4 | 4.23 | 4.15 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 130 | Sn | 1 | 5 | 5 |  |  | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 130 | Sn | 1 | 10 | 10 |  |  | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 130 | Sn | 3 | 3 | 9 | 6.15 | 6.04 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 130 | Sn | 3 | 5 | 15 | 5.58 | 5.48 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 130 | Sn | 32 | 0.6 | 19.2 |  |  | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 130 | Sn | 64 | 0.6 | 38.4 | 5.8 | 5.71 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 140 | Sn | 1 | 5 | 5 |  |  | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 140 | Sn | 1 | 10 | 10 |  |  | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 140 | Sn | 3 | 3 | 9 | 8.13 | 8.01 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 140 | Sn | 3 | 5 | 15 | 7.36 | 7.25 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 140 | Sn | 32 | 0.6 | 19.2 |  |  | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 16 | Narrow | 140 | Sn | 64 | 0.6 | 38.4 | 7.64 | 7.55 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 70 |  | 1 | 5 | 5 | 0.91 | 1.83 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 70 |  | 1 | 10 | 10 | 0.9 | 1.79 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 70 |  | 3 | 3 | 9 | 1 | 1.98 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 70 |  | 3 | 5 | 15 | 0.9 | 1.79 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 70 |  | 32 | 0.6 | 19.2 | 1.05 | 2.1 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 70 |  | 64 | 0.6 | 38.4 | 0.93 | 1.87 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 80 |  | 1 | 5 | 5 | 1.59 | 2.96 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 80 |  | 1 | 10 | 10 | 1.57 | 2.89 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 80 |  | 3 | 3 | 9 | 1.74 | 3.21 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 80 |  | 3 | 5 | 15 | 1.57 | 2.9 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 80 |  | 32 | 0.6 | 19.2 | 1.84 | 3.39 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 80 |  | 64 | 0.6 | 38.4 | 1.63 | 3.01 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 90 |  | 1 | 5 | 5 | 2.46 | 4.35 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 90 |  | 1 | 10 | 10 | 2.42 | 4.25 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 90 |  | 3 | 3 | 9 | 2.68 | 4.72 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 90 |  | 3 | 5 | 15 | 2.43 | 4.26 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 90 |  | 32 | 0.6 | 19.2 | 2.85 | 4.98 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 90 |  | 64 | 0.6 | 38.4 | 2.51 | 4.42 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 100 |  | 1 | 5 | 5 | 3.49 | 5.96 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 100 |  | 1 | 10 | 10 | 3.43 | 5.81 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 100 |  | 3 | 3 | 9 | 3.82 | 6.46 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 100 |  | 3 | 5 | 15 | 3.45 | 5.84 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 100 |  | 32 | 0.6 | 19.2 | 4.03 | 6.82 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 100 |  | 64 | 0.6 | 38.4 | 3.56 | 6.06 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 110 |  | 1 | 5 | 5 | 4.69 | 7.79 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 110 |  | 1 | 10 | 10 | 4.58 | 7.59 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 110 |  | 3 | 3 | 9 | 5.09 | 8.44 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 110 |  | 3 | 5 | 15 | 4.62 | 7.62 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 110 |  | 32 | 0.6 | 19.2 | 5.4 | 8.91 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 110 |  | 64 | 0.6 | 38.4 | 4.77 | 7.9 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 120 |  | 1 | 5 | 5 | 5.98 | 9.8 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 120 |  | 1 | 10 | 10 | 5.87 | 9.54 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 120 |  | 3 | 3 | 9 | 6.52 | 10.6 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 120 |  | 3 | 5 | 15 | 5.9 | 9.57 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 120 |  | 32 | 0.6 | 19.2 | 6.92 | 11.19 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 120 |  | 64 | 0.6 | 38.4 | 6.11 | 9.92 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 130 |  | 1 | 5 | 5 | 7.41 | 12 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 130 |  | 1 | 10 | 10 | 7.27 | 11.68 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 130 |  | 3 | 3 | 9 | 8.08 | 12.97 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 130 |  | 3 | 5 | 15 | 7.33 | 1.71 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 130 |  | 32 | 0.6 | 19.2 | 8.56 | 13.68 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 130 |  | 64 | 0.6 | 38.4 | 7.57 | 12.12 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 140 |  | 1 | 5 | 5 | 8.95 | 14.39 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 140 |  | 1 | 10 | 10 | 8.77 | 13.96 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 140 |  | 3 | 3 | 9 | 9.75 | 15.51 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 140 |  | 3 | 5 | 15 | 8.83 | 14 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 140 |  | 32 | 0.6 | 19.2 | 10.34 | 16.35 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 140 |  | 64 | 0.6 | 38.4 | 9.13 | 14.47 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 100 | Sn | 1 | 5 | 5 |  |  | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 100 | Sn | 1 | 10 | 10 |  |  | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 100 | Sn | 3 | 3 | 9 | 0.73 | 1.1 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 100 | Sn | 3 | 5 | 15 | 0.66 | 1 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 100 | Sn | 32 | 0.6 | 19.2 |  |  | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 100 | Sn | 64 | 0.6 | 38.4 | 0.69 | 1.04 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 110 | Sn | 1 | 5 | 5 |  |  | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 110 | Sn | 1 | 10 | 10 |  |  | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 110 | Sn | 3 | 3 | 9 | 1.16 | 1.74 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 110 | Sn | 3 | 5 | 15 | 1.06 | 1.57 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 110 | Sn | 32 | 0.6 | 19.2 |  |  | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 110 | Sn | 64 | 0.6 | 38.4 | 1.1 | 1.65 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 120 | Sn | 1 | 5 | 5 |  |  | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 120 | Sn | 1 | 10 | 10 |  |  | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 120 | Sn | 3 | 3 | 9 | 1.72 | 2.54 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 120 | Sn | 3 | 5 | 15 | 1.56 | 2.3 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 120 | Sn | 32 | 0.6 | 19.2 |  |  | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 120 | Sn | 64 | 0.6 | 38.4 | 1.63 | 2.41 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 130 | Sn | 1 | 5 | 5 |  |  | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 130 | Sn | 1 | 10 | 10 |  |  | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 130 | Sn | 3 | 3 | 9 | 2.39 | 3.53 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 130 | Sn | 3 | 5 | 15 | 2.17 | 3.19 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 130 | Sn | 32 | 0.6 | 19.2 |  |  | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 130 | Sn | 64 | 0.6 | 38.4 | 2.27 | 3.34 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 140 | Sn | 1 | 5 | 5 |  |  | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 140 | Sn | 1 | 10 | 10 |  |  | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 140 | Sn | 3 | 3 | 9 | 3.18 | 4.7 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 140 | Sn | 3 | 5 | 15 | 2.89 | 4.25 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 140 | Sn | 32 | 0.6 | 19.2 |  |  | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Standard | 140 | Sn | 64 | 0.6 | 38.4 | 3.02 | 4.44 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 70 |  | 1 | 5 | 5 | 0.79 | 1.38 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 70 |  | 1 | 10 | 10 | 0.79 | 1.36 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 70 |  | 3 | 3 | 9 | 0.86 | 1.51 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 70 |  | 3 | 5 | 15 | 0.79 | 1.36 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 70 |  | 32 | 0.6 | 19.2 | 0.91 | 1.59 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 70 |  | 64 | 0.6 | 38.4 | 0.8 | 1.41 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 80 |  | 1 | 5 | 5 | 1.39 | 2.26 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 80 |  | 1 | 10 | 10 | 1.37 | 2.22 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 80 |  | 3 | 3 | 9 | 1.52 | 2.47 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 80 |  | 3 | 5 | 15 | 1.37 | 2.23 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 80 |  | 32 | 0.6 | 19.2 | 1.61 | 2.61 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 80 |  | 64 | 0.6 | 38.4 | 1.42 | 2.31 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 90 |  | 1 | 5 | 5 | 2.17 | 3.37 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 90 |  | 1 | 10 | 10 | 2.13 | 3.3 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 90 |  | 3 | 3 | 9 | 2.36 | 3.67 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 90 |  | 3 | 5 | 15 | 2.13 | 3.31 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 90 |  | 32 | 0.6 | 19.2 | 2.5 | 3.87 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 90 |  | 64 | 0.6 | 38.4 | 2.2 | 3.43 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 100 |  | 1 | 5 | 5 | 3.09 | 4.66 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 100 |  | 1 | 10 | 10 | 3.04 | 4.57 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 100 |  | 3 | 3 | 9 | 3.38 | 5.08 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 100 |  | 3 | 5 | 15 | 3.05 | 4.58 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 100 |  | 32 | 0.6 | 19.2 | 3.57 | 5.36 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 100 |  | 64 | 0.6 | 38.4 | 3.13 | 4.73 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 110 |  | 1 | 5 | 5 | 4.15 | 6.14 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 110 |  | 1 | 10 | 10 | 4.08 | 6.01 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 110 |  | 3 | 3 | 9 | 4.54 | 6.68 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 110 |  | 3 | 5 | 15 | 4.1 | 6.03 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 110 |  | 32 | 0.6 | 19.2 | 4.79 | 7.04 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 110 |  | 64 | 0.6 | 38.4 | 4.22 | 6.22 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 120 |  | 1 | 5 | 5 | 5.35 | 7.78 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 120 |  | 1 | 10 | 10 | 5.25 | 7.6 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 120 |  | 3 | 3 | 9 | 5.83 | 8.45 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 120 |  | 3 | 5 | 15 | 5.26 | 7.62 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 120 |  | 32 | 0.6 | 19.2 | 6.17 | 8.91 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 120 |  | 64 | 0.6 | 38.4 | 5.42 | 7.87 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 130 |  | 1 | 5 | 5 | 6.67 | 9.57 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 130 |  | 1 | 10 | 10 | 6.54 | 9.36 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 130 |  | 3 | 3 | 9 | 7.26 | 10.4 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 130 |  | 3 | 5 | 15 | 6.55 | 9.37 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 130 |  | 32 | 0.6 | 19.2 | 7.68 | 10.95 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 130 |  | 64 | 0.6 | 38.4 | 6.74 | 9.66 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 140 |  | 1 | 5 | 5 | 8.03 | 11.5 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 140 |  | 1 | 10 | 10 | 7.9 | 11.25 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 140 |  | 3 | 3 | 9 | 8.78 | 12.5 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 140 |  | 3 | 5 | 15 | 7.93 | 11.25 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 140 |  | 32 | 0.6 | 19.2 | 9.3 | 13.15 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 140 |  | 64 | 0.6 | 38.4 | 8.18 | 11.61 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 100 | Sn | 1 | 5 | 5 |  |  | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 100 | Sn | 1 | 10 | 10 |  |  | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 100 | Sn | 3 | 3 | 9 | 0.66 | 0.9 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 100 | Sn | 3 | 5 | 15 | 0.59 | 0.81 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 100 | Sn | 32 | 0.6 | 19.2 |  |  | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 100 | Sn | 64 | 0.6 | 38.4 | 0.62 | 0.85 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 110 | Sn | 1 | 5 | 5 |  |  | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 110 | Sn | 1 | 10 | 10 |  |  | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 110 | Sn | 3 | 3 | 9 | 1.05 | 1.42 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 110 | Sn | 3 | 5 | 15 | 0.95 | 1.29 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 110 | Sn | 32 | 0.6 | 19.2 |  |  | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 110 | Sn | 64 | 0.6 | 38.4 | 0.99 | 1.34 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 120 | Sn | 1 | 5 | 5 |  |  | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 120 | Sn | 1 | 10 | 10 |  |  | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 120 | Sn | 3 | 3 | 9 | 1.56 | 2.1 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 120 | Sn | 3 | 5 | 15 | 1.41 | 1.9 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 120 | Sn | 32 | 0.6 | 19.2 |  |  | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 120 | Sn | 64 | 0.6 | 38.4 | 1.47 | 1.98 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 130 | Sn | 1 | 5 | 5 |  |  | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 130 | Sn | 1 | 10 | 10 |  |  | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 130 | Sn | 3 | 3 | 9 | 2.18 | 2.92 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 130 | Sn | 3 | 5 | 15 | 1.98 | 2.64 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 130 | Sn | 32 | 0.6 | 19.2 |  |  | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 130 | Sn | 64 | 0.6 | 38.4 | 2.06 | 2.76 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 140 | Sn | 1 | 5 | 5 |  |  | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 140 | Sn | 1 | 10 | 10 |  |  | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 140 | Sn | 3 | 3 | 9 | 2.92 | 3.91 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 140 | Sn | 3 | 5 | 15 | 2.64 | 3.53 | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 140 | Sn | 32 | 0.6 | 19.2 |  |  | Pro.Pulse A tube | Siemens
Pro.Pulse A tube | 14 | 1 | 32 | Narrow | 140 | Sn | 64 | 0.6 | 38.4 | 2.75 | 6.68 | Pro.Pulse A tube | Siemens
Pro.Pulse B tube | 15 | 1 | 16 | Standard | 70 |  | 32 | 0.6 | 19.2 |  |  | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 16 | Standard | 70 |  | 64 | 0.6 | 38.4 | 4.4 | 4.74 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 16 | Standard | 80 |  | 32 | 0.6 | 19.2 |  |  | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 16 | Standard | 80 |  | 64 | 0.6 | 38.4 | 6.92 | 7.32 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 16 | Standard | 90 |  | 32 | 0.6 | 19.2 |  |  | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 16 | Standard | 90 |  | 64 | 0.6 | 38.4 | 9.92 | 10.38 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 16 | Standard | 100 |  | 32 | 0.6 | 19.2 |  |  | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 16 | Standard | 100 |  | 64 | 0.6 | 38.4 | 13.32 | 13.84 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 16 | Standard | 110 |  | 32 | 0.6 | 19.2 |  |  | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 16 | Standard | 110 |  | 64 | 0.6 | 38.4 | 17.02 | 17.62 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 16 | Standard | 120 |  | 32 | 0.6 | 19.2 |  |  | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 16 | Standard | 120 |  | 64 | 0.6 | 38.4 | 21.08 | 21.72 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 16 | Standard | 130 |  | 32 | 0.6 | 19.2 |  |  | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 16 | Standard | 130 |  | 64 | 0.6 | 38.4 | 25.36 | 26.1 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 16 | Standard | 140 |  | 32 | 0.6 | 19.2 |  |  | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 16 | Standard | 140 |  | 64 | 0.6 | 38.4 | 30 | 30.77 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 16 | Standard | 100 | Sn | 32 | 0.6 | 19.2 | 2.41 | 2.42 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 16 | Standard | 100 | Sn | 64 | 0.6 | 38.4 | 2.08 | 2.11 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 16 | Standard | 110 | Sn | 32 | 0.6 | 19.2 | 3.71 | 3.7 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 16 | Standard | 110 | Sn | 64 | 0.6 | 38.4 | 3.22 | 3.26 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 16 | Standard | 120 | Sn | 32 | 0.6 | 19.2 | 5.35 | 5.41 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 16 | Standard | 120 | Sn | 64 | 0.6 | 38.4 | 4.63 | 4.69 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 16 | Standard | 130 | Sn | 32 | 0.6 | 19.2 | 7.33 | 7.41 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 16 | Standard | 130 | Sn | 64 | 0.6 | 38.4 | 6.33 | 6.43 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 16 | Standard | 140 | Sn | 32 | 0.6 | 19.2 | 9.62 | 9.75 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 16 | Standard | 140 | Sn | 64 | 0.6 | 38.4 | 8.32 | 8.46 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 16 | Narrow | 70 |  | 32 | 0.6 | 19.2 |  |  | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 16 | Narrow | 70 |  | 64 | 0.6 | 38.4 | 3.93 | 3.97 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 16 | Narrow | 80 |  | 32 | 0.6 | 19.2 |  |  | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 16 | Narrow | 80 |  | 64 | 0.6 | 38.4 | 6.25 | 6.23 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 16 | Narrow | 90 |  | 32 | 0.6 | 19.2 |  |  | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 16 | Narrow | 90 |  | 64 | 0.6 | 38.4 | 9.04 | 8.95 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 16 | Narrow | 100 |  | 32 | 0.6 | 19.2 |  |  | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 16 | Narrow | 100 |  | 64 | 0.6 | 38.4 | 12.22 | 12.05 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 16 | Narrow | 110 |  | 32 | 0.6 | 19.2 |  |  | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 16 | Narrow | 110 |  | 64 | 0.6 | 38.4 | 15.7 | 15.46 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 16 | Narrow | 120 |  | 32 | 0.6 | 19.2 |  |  | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 16 | Narrow | 120 |  | 64 | 0.6 | 38.4 | 19.52 | 19.18 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 16 | Narrow | 130 |  | 32 | 0.6 | 19.2 |  |  | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 16 | Narrow | 130 |  | 64 | 0.6 | 38.4 | 23.57 | 23.18 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 16 | Narrow | 140 |  | 32 | 0.6 | 19.2 |  |  | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 16 | Narrow | 140 |  | 64 | 0.6 | 38.4 | 27.92 | 27.46 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 16 | Narrow | 100 | Sn | 32 | 0.6 | 19.2 | 2.26 | 2.2 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 16 | Narrow | 100 | Sn | 64 | 0.6 | 38.4 | 1.96 | 1.91 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 16 | Narrow | 110 | Sn | 32 | 0.6 | 19.2 | 3.51 | 3.43 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 16 | Narrow | 110 | Sn | 64 | 0.6 | 38.4 | 3.04 | 2.97 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 16 | Narrow | 120 | Sn | 32 | 0.6 | 19.2 | 5.08 | 4.97 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 16 | Narrow | 120 | Sn | 64 | 0.6 | 38.4 | 4.39 | 4.31 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 16 | Narrow | 130 | Sn | 32 | 0.6 | 19.2 | 6.96 | 6.83 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 16 | Narrow | 130 | Sn | 64 | 0.6 | 38.4 | 6.01 | 5.92 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 16 | Narrow | 140 | Sn | 32 | 0.6 | 19.2 | 9.17 | 9.02 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 16 | Narrow | 140 | Sn | 64 | 0.6 | 38.4 | 7.92 | 7.81 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 32 | Standard | 70 |  | 32 | 0.6 | 19.2 |  |  | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 32 | Standard | 70 |  | 64 | 0.6 | 38.4 | 0.98 | 1.96 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 32 | Standard | 80 |  | 32 | 0.6 | 19.2 |  |  | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 32 | Standard | 80 |  | 64 | 0.6 | 38.4 | 1.7 | 3.14 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 32 | Standard | 90 |  | 32 | 0.6 | 19.2 |  |  | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 32 | Standard | 90 |  | 64 | 0.6 | 38.4 | 2.96 | 4.59 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 32 | Standard | 100 |  | 32 | 0.6 | 19.2 |  |  | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 32 | Standard | 100 |  | 64 | 0.6 | 38.4 | 3.7 | 6.28 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 32 | Standard | 110 |  | 32 | 0.6 | 19.2 |  |  | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 32 | Standard | 110 |  | 64 | 0.6 | 38.4 | 4.94 | 8.16 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 32 | Standard | 120 |  | 32 | 0.6 | 19.2 |  |  | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 32 | Standard | 120 |  | 64 | 0.6 | 38.4 | 6.31 | 10.24 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 32 | Standard | 130 |  | 32 | 0.6 | 19.2 |  |  | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 32 | Standard | 130 |  | 64 | 0.6 | 38.4 | 7.81 | 12.49 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 32 | Standard | 140 |  | 32 | 0.6 | 19.2 |  |  | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 32 | Standard | 140 |  | 64 | 0.6 | 38.4 | 9.44 | 14.91 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 32 | Standard | 100 | Sn | 32 | 0.6 | 19.2 | 0.83 | 1.24 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 32 | Standard | 100 | Sn | 64 | 0.6 | 38.4 | 0.71 | 1.08 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 32 | Standard | 110 | Sn | 32 | 0.6 | 19.2 | 1.32 | 1.95 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 32 | Standard | 110 | Sn | 64 | 0.6 | 38.4 | 1.14 | 1.7 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 32 | Standard | 120 | Sn | 32 | 0.6 | 19.2 | 1.94 | 2.85 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 32 | Standard | 120 | Sn | 64 | 0.6 | 38.4 | 1.68 | 2.48 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 32 | Standard | 130 | Sn | 32 | 0.6 | 19.2 | 2.7 | 3.94 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 32 | Standard | 130 | Sn | 64 | 0.6 | 38.4 | 2.34 | 3.43 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 32 | Standard | 140 | Sn | 32 | 0.6 | 19.2 | 3.6 | 5.24 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 32 | Standard | 140 | Sn | 64 | 0.6 | 38.4 | 3.12 | 4.56 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 32 | Narrow | 70 |  | 32 | 0.6 | 19.2 |  |  | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 32 | Narrow | 70 |  | 64 | 0.6 | 38.4 | 0.85 | 1.48 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 32 | Narrow | 80 |  | 32 | 0.6 | 19.2 |  |  | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 32 | Narrow | 80 |  | 64 | 0.6 | 38.4 | 1.48 | 2.41 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 32 | Narrow | 90 |  | 32 | 0.6 | 19.2 |  |  | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 32 | Narrow | 90 |  | 64 | 0.6 | 38.4 | 2.3 | 3.57 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 32 | Narrow | 100 |  | 32 | 0.6 | 19.2 |  |  | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 32 | Narrow | 100 |  | 64 | 0.6 | 38.4 | 3.27 | 4.92 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 32 | Narrow | 110 |  | 32 | 0.6 | 19.2 |  |  | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 32 | Narrow | 110 |  | 64 | 0.6 | 38.4 | 4.38 | 6.45 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 32 | Narrow | 120 |  | 32 | 0.6 | 19.2 |  |  | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 32 | Narrow | 120 |  | 64 | 0.6 | 38.4 | 5.62 | 8.14 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 32 | Narrow | 130 |  | 32 | 0.6 | 19.2 |  |  | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 32 | Narrow | 130 |  | 64 | 0.6 | 38.4 | 6.98 | 9.98 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 32 | Narrow | 140 |  | 32 | 0.6 | 19.2 |  |  | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 32 | Narrow | 140 |  | 64 | 0.6 | 38.4 | 8.46 | 11.98 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 32 | Narrow | 100 | Sn | 32 | 0.6 | 19.2 | 0.74 | 1.01 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 32 | Narrow | 100 | Sn | 64 | 0.6 | 38.4 | 0.64 | 0.88 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 32 | Narrow | 110 | Sn | 32 | 0.6 | 19.2 | 1.19 | 1.6 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 32 | Narrow | 110 | Sn | 64 | 0.6 | 38.4 | 1.03 | 1.39 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 32 | Narrow | 120 | Sn | 32 | 0.6 | 19.2 | 1.76 | 2.35 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 32 | Narrow | 120 | Sn | 64 | 0.6 | 38.4 | 1.53 | 2.04 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 32 | Narrow | 130 | Sn | 32 | 0.6 | 19.2 | 2.46 | 3.27 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 32 | Narrow | 130 | Sn | 64 | 0.6 | 38.4 | 2.13 | 2.84 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 32 | Narrow | 140 | Sn | 32 | 0.6 | 19.2 | 3.3 | 4.37 | Pro.Pulse B tube | Siemens
Pro.Pulse B tube | 15 | 1 | 32 | Narrow | 140 | Sn | 64 | 0.6 | 38.4 | 2.85 | 3.79 | Pro.Pulse B tube | Siemens
BodyTom | 16 | 2 | 16 | Standard | 80 |  | 16 | 0.6 | 9.6 | 7.9 | 9.46 | BodyTom | Neurologica
BodyTom | 16 | 2 | 16 | Standard | 100 |  | 16 | 0.6 | 9.6 | 16.23 | 18.1 | BodyTom | Neurologica
BodyTom | 16 | 2 | 16 | Standard | 120 |  | 16 | 0.6 | 9.6 | 24.62 | 27.07 | BodyTom | Neurologica
BodyTom | 16 | 2 | 16 | Standard | 140 |  | 16 | 0.6 | 9.6 | 35 | 37.74 | BodyTom | Neurologica
BodyTom | 16 | 2 | 32 | Standard | 80 |  | 16 | 0.6 | 9.6 | 1.49 | 3.15 | BodyTom | Neurologica
BodyTom | 16 | 2 | 32 | Standard | 100 |  | 16 | 0.6 | 9.6 | 3.08 | 5.85 | BodyTom | Neurologica
BodyTom | 16 | 2 | 32 | Standard | 120 |  | 16 | 0.6 | 9.6 | 5.53 | 9.23 | BodyTom | Neurologica
BodyTom | 16 | 2 | 32 | Standard | 140 |  | 16 | 0.6 | 9.6 | 8.26 | 13.13 | BodyTom | Neurologica
Pro.Specta | 13 | 1 | 16 | Standard | 80 |  | 1 | 5 | 5 | 5.33 | 5.92 | Symbia Pro.specta | Siemens
Pro.Specta | 13 | 1 | 16 | Standard | 80 |  | 1 | 10 | 10 | 5.36 | 5.93 | Symbia Pro.specta | Siemens
Pro.Specta | 13 | 1 | 16 | Standard | 80 |  | 3 | 3 | 9 | 5.96 | 6.59 | Symbia Pro.specta | Siemens
Pro.Specta | 13 | 1 | 16 | Standard | 80 |  | 16 | 0.7 | 11.2 | 6.38 | 7.08 | Symbia Pro.specta | Siemens
Pro.Specta | 13 | 1 | 16 | Standard | 80 | Sn | 16 | 0.7 | 11.2 |  |  | Symbia Pro.specta | Siemens
Pro.Specta | 13 | 1 | 16 | Standard | 110 |  | 1 | 5 | 5 | 13.33 | 14.3 | Symbia Pro.specta | Siemens
Pro.Specta | 13 | 1 | 16 | Standard | 110 |  | 1 | 10 | 10 | 13.39 | 14.32 | Symbia Pro.specta | Siemens
Pro.Specta | 13 | 1 | 16 | Standard | 110 |  | 3 | 3 | 9 | 14.87 | 15.92 | Symbia Pro.specta | Siemens
Pro.Specta | 13 | 1 | 16 | Standard | 110 |  | 16 | 0.7 | 11.2 | 15.98 | 17.1 | Symbia Pro.specta | Siemens
Pro.Specta | 13 | 1 | 16 | Standard | 110 | Sn | 16 | 0.7 | 11.2 | 2.99 | 3.08 | Symbia Pro.specta | Siemens
Pro.Specta | 13 | 1 | 16 | Standard | 130 |  | 1 | 5 | 5 | 20.27 | 21.52 | Symbia Pro.specta | Siemens
Pro.Specta | 13 | 1 | 16 | Standard | 130 |  | 1 | 10 | 10 | 20.27 | 21.52 | Symbia Pro.specta | Siemens
Pro.Specta | 13 | 1 | 16 | Standard | 130 |  | 3 | 3 | 9 | 22.52 | 23.91 | Symbia Pro.specta | Siemens
Pro.Specta | 13 | 1 | 16 | Standard | 130 |  | 16 | 0.7 | 11.2 | 24.24 | 25.7 | Symbia Pro.specta | Siemens
Pro.Specta | 13 | 1 | 16 | Standard | 130 | Sn | 16 | 0.7 | 11.2 | 5.86 | 6.05 | Symbia Pro.specta | Siemens
Pro.Specta | 13 | 1 | 32 | Standard | 80 |  | 1 | 5 | 5 | 1.36 | 2.64 | Symbia Pro.specta | Siemens
Pro.Specta | 13 | 1 | 32 | Standard | 80 |  | 1 | 10 | 10 | 1.36 | 2.64 | Symbia Pro.specta | Siemens
Pro.Specta | 13 | 1 | 32 | Standard | 80 |  | 3 | 3 | 9 | 1.51 | 2.94 | Symbia Pro.specta | Siemens
Pro.Specta | 13 | 1 | 32 | Standard | 80 |  | 16 | 0.7 | 11.2 | 1.62 | 3.15 | Symbia Pro.specta | Siemens
Pro.Specta | 13 | 1 | 32 | Standard | 80 | Sn | 16 | 0.7 | 11.2 |  |  | Symbia Pro.specta | Siemens
Pro.Specta | 13 | 1 | 32 | Standard | 110 |  | 1 | 5 | 5 | 3.98 | 6.82 | Symbia Pro.specta | Siemens
Pro.Specta | 13 | 1 | 32 | Standard | 110 |  | 1 | 10 | 10 | 3.97 | 6.83 | Symbia Pro.specta | Siemens
Pro.Specta | 13 | 1 | 32 | Standard | 110 |  | 3 | 3 | 9 | 4.41 | 7.59 | Symbia Pro.specta | Siemens
Pro.Specta | 13 | 1 | 32 | Standard | 110 |  | 16 | 0.7 | 11.2 | 4.75 | 8.13 | Symbia Pro.specta | Siemens
Pro.Specta | 13 | 1 | 32 | Standard | 110 | Sn | 16 | 0.7 | 11.2 | 1.09 | 1.65 | Symbia Pro.specta | Siemens
Pro.Specta | 13 | 1 | 32 | Standard | 130 |  | 1 | 5 | 5 | 6.4 | 10.57 | Symbia Pro.specta | Siemens
Pro.Specta | 13 | 1 | 32 | Standard | 130 |  | 1 | 10 | 10 | 6.39 | 10.57 | Symbia Pro.specta | Siemens
Pro.Specta | 13 | 1 | 32 | Standard | 130 |  | 3 | 3 | 9 | 7.1 | 11.74 | Symbia Pro.specta | Siemens
Pro.Specta | 13 | 1 | 32 | Standard | 130 |  | 16 | 0.7 | 11.2 | 7.63 | 12.57 | Symbia Pro.specta | Siemens
Pro.Specta | 13 | 1 | 32 | Standard | 130 | Sn | 16 | 0.7 | 11.2 | 2.23 | 3.32 | Symbia Pro.specta | Siemens
Intevo Bold | 12 | 1 | 16 | Standard | 80 |  | 4 | 0.6 | 2.4 | 7.6 | 8.12 | Intevo Bold | Siemens
Intevo Bold | 12 | 1 | 16 | Standard | 80 |  | 12 | 0.6 | 7.2 | 7.7 | 8.2 | Intevo Bold | Siemens
Intevo Bold | 12 | 1 | 16 | Standard | 80 |  | 16 | 0.6 | 9.6 | 6.64 | 7.19 | Intevo Bold | Siemens
Intevo Bold | 12 | 1 | 16 | Standard | 80 |  | 1 | 10 | 10 | 5.55 | 5.9 | Intevo Bold | Siemens
Intevo Bold | 12 | 1 | 16 | Standard | 80 |  | 12 | 1.2 | 14.4 | 6.04 | 6.45 | Intevo Bold | Siemens
Intevo Bold | 12 | 1 | 16 | Standard | 80 |  | 2 | 8 | 16 | 5.43 | 5.81 | Intevo Bold | Siemens
Intevo Bold | 12 | 1 | 16 | Standard | 80 |  | 16 | 1.2 | 19.2 | 6.13 | 6.56 | Intevo Bold | Siemens
Intevo Bold | 12 | 1 | 16 | Standard | 110 |  | 4 | 0.6 | 2.4 | 18 | 18.65 | Intevo Bold | Siemens
Intevo Bold | 12 | 1 | 16 | Standard | 110 |  | 12 | 0.6 | 7.2 | 17.63 | 18.31 | Intevo Bold | Siemens
Intevo Bold | 12 | 1 | 16 | Standard | 110 |  | 16 | 0.6 | 9.6 | 15.57 | 16.18 | Intevo Bold | Siemens
Intevo Bold | 12 | 1 | 16 | Standard | 110 |  | 1 | 10 | 10 | 12.69 | 13.18 | Intevo Bold | Siemens
Intevo Bold | 12 | 1 | 16 | Standard | 110 |  | 12 | 1.2 | 14.4 | 13.88 | 14.43 | Intevo Bold | Siemens
Intevo Bold | 12 | 1 | 16 | Standard | 110 |  | 2 | 8 | 16 | 12.49 | 12.99 | Intevo Bold | Siemens
Intevo Bold | 12 | 1 | 16 | Standard | 110 |  | 16 | 1.2 | 19.2 | 14.14 | 14.7 | Intevo Bold | Siemens
Intevo Bold | 12 | 1 | 16 | Standard | 130 |  | 4 | 0.6 | 2.4 | 25.78 | 26.52 | Intevo Bold | Siemens
Intevo Bold | 12 | 1 | 16 | Standard | 130 |  | 12 | 0.6 | 7.2 | 25.93 | 26.72 | Intevo Bold | Siemens
Intevo Bold | 12 | 1 | 16 | Standard | 130 |  | 16 | 0.6 | 9.6 | 22.91 | 23.62 | Intevo Bold | Siemens
Intevo Bold | 12 | 1 | 16 | Standard | 130 |  | 1 | 10 | 10 | 18.67 | 19.24 | Intevo Bold | Siemens
Intevo Bold | 12 | 1 | 16 | Standard | 130 |  | 12 | 1.2 | 14.4 | 20.42 | 21.05 | Intevo Bold | Siemens
Intevo Bold | 12 | 1 | 16 | Standard | 130 |  | 2 | 8 | 16 | 18.38 | 18.94 | Intevo Bold | Siemens
Intevo Bold | 12 | 1 | 16 | Standard | 130 |  | 16 | 1.2 | 19.2 | 20.81 | 21.45 | Intevo Bold | Siemens
Intevo Bold | 12 | 1 | 32 | Standard | 80 |  | 4 | 0.6 | 2.4 | 1.88 | 3.49 | Intevo Bold | Siemens
Intevo Bold | 12 | 1 | 32 | Standard | 80 |  | 12 | 0.6 | 7.2 | 1.9 | 3.54 | Intevo Bold | Siemens
Intevo Bold | 12 | 1 | 32 | Standard | 80 |  | 16 | 0.6 | 9.6 | 1.68 | 3.12 | Intevo Bold | Siemens
Intevo Bold | 12 | 1 | 32 | Standard | 80 |  | 1 | 10 | 10 | 1.37 | 2.55 | Intevo Bold | Siemens
Intevo Bold | 12 | 1 | 32 | Standard | 80 |  | 12 | 1.2 | 14.4 | 1.5 | 2.78 | Intevo Bold | Siemens
Intevo Bold | 12 | 1 | 32 | Standard | 80 |  | 2 | 8 | 16 | 1.35 | 2.5 | Intevo Bold | Siemens
Intevo Bold | 12 | 1 | 32 | Standard | 80 |  | 16 | 1.2 | 19.2 | 1.52 | 2.9 | Intevo Bold | Siemens
Intevo Bold | 12 | 1 | 32 | Standard | 110 |  | 4 | 0.6 | 2.4 | 5.24 | 8.63 | Intevo Bold | Siemens
Intevo Bold | 12 | 1 | 32 | Standard | 110 |  | 12 | 0.6 | 7.2 | 5.12 | 8.5 | Intevo Bold | Siemens
Intevo Bold | 12 | 1 | 32 | Standard | 110 |  | 16 | 0.6 | 9.6 | 4.52 | 7.51 | Intevo Bold | Siemens
Intevo Bold | 12 | 1 | 32 | Standard | 110 |  | 1 | 10 | 10 | 3.68 | 6.12 | Intevo Bold | Siemens
Intevo Bold | 12 | 1 | 32 | Standard | 110 |  | 12 | 1.2 | 14.4 | 4.05 | 6.69 | Intevo Bold | Siemens
Intevo Bold | 12 | 1 | 32 | Standard | 110 |  | 2 | 8 | 16 | 3.64 | 6.02 | Intevo Bold | Siemens
Intevo Bold | 12 | 1 | 32 | Standard | 110 |  | 16 | 1.2 | 19.2 | 4.12 | 6.82 | Intevo Bold | Siemens
Intevo Bold | 12 | 1 | 32 | Standard | 130 |  | 4 | 0.6 | 2.4 | 7.95 | 12.63 | Intevo Bold | Siemens
Intevo Bold | 12 | 1 | 32 | Standard | 130 |  | 12 | 0.6 | 7.2 | 8.02 | 12.79 | Intevo Bold | Siemens
Intevo Bold | 12 | 1 | 32 | Standard | 130 |  | 16 | 0.6 | 9.6 | 7.09 | 11.29 | Intevo Bold | Siemens
Intevo Bold | 12 | 1 | 32 | Standard | 130 |  | 1 | 10 | 10 | 5.78 | 9.21 | Intevo Bold | Siemens
Intevo Bold | 12 | 1 | 32 | Standard | 130 |  | 12 | 1.2 | 14.4 | 6.31 | 10.04 | Intevo Bold | Siemens
Intevo Bold | 12 | 1 | 32 | Standard | 130 |  | 2 | 8 | 16 | 5.68 | 9.04 | Intevo Bold | Siemens
Intevo Bold | 12 | 1 | 32 | Standard | 130 |  | 16 | 1.2 | 19.2 | 6.41 | 10.21 | Intevo Bold | Siemens
Biograph Vision | 11 | 1 | 16 | Standard | 70 |  | 2 | 1 | 2 | 2.6 | 2.74 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Standard | 70 |  | 6 | 0.6 | 3.6 | 4.22 | 4.43 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Standard | 70 |  | 8 | 0.6 | 4.8 | 3.88 | 4.07 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Standard Z-sharp | 70 |  | 8 | 0.6 | 4.8 | 4.58 | 4.78 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Standard | 70 |  | 1 | 5 | 5 | 2.34 | 2.46 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Standard Z-sharp | 70 |  | 10 | 0.6 | 6 | 3.17 | 3.31 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Standard | 70 |  | 6 | 1.2 | 7.2 | 3.3 | 3.35 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Standard | 70 |  | 1 | 10 | 10 | 2.34 | 2.46 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Standard Z-sharp | 70 |  | 20 | 0.6 | 12 | 3.23 | 3.38 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Standard | 70 |  | 12 | 1.2 | 14.4 | 3 | 3.15 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Standard | 70 |  | 64 | 0.6 | 19.2 | 3.01 | 3.15 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Standard Z-sharp | 70 |  | 64 | 0.6 | 38.4 | 2.87 | 3.01 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Standard | 80 |  | 2 | 1 | 2 | 4.18 | 4.36 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Standard | 80 |  | 6 | 0.6 | 3.6 | 6.76 | 7.06 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Standard | 80 |  | 8 | 0.6 | 4.8 | 6.22 | 6.49 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Standard Z-sharp | 80 |  | 8 | 0.6 | 4.8 | 7.24 | 7.46 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Standard | 80 |  | 1 | 5 | 5 | 3.75 | 3.91 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Standard Z-sharp | 80 |  | 10 | 0.6 | 6 | 5.02 | 5.17 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Standard | 80 |  | 6 | 1.2 | 7.2 | 5.39 | 5.62 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Standard | 80 |  | 1 | 10 | 10 | 3.75 | 3.91 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Standard Z-sharp | 80 |  | 20 | 0.6 | 12 | 5.12 | 5.27 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Standard | 80 |  | 12 | 1.2 | 14.4 | 4.81 | 5.02 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Standard | 80 |  | 64 | 0.6 | 19.2 | 4.77 | 4.91 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Standard Z-sharp | 80 |  | 64 | 0.6 | 38.4 | 4.6 | 4.8 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Standard | 100 |  | 2 | 1 | 2 | 8.44 | 8.65 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Standard | 100 |  | 6 | 0.6 | 3.6 | 13.67 | 14.01 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Standard | 100 |  | 8 | 0.6 | 4.8 | 12.56 | 12.88 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Standard Z-sharp | 100 |  | 8 | 0.6 | 4.8 | 14.69 | 14.93 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Standard | 100 |  | 1 | 5 | 5 | 7.57 | 7.76 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Standard Z-sharp | 100 |  | 10 | 0.6 | 6 | 10.19 | 10.35 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Standard | 100 |  | 6 | 1.2 | 7.2 | 10.89 | 11.16 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Standard | 100 |  | 1 | 10 | 10 | 7.57 | 7.76 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Standard Z-sharp | 100 |  | 20 | 0.6 | 12 | 10.38 | 10.55 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Standard | 100 |  | 12 | 1.2 | 14.4 | 9.71 | 9.96 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Standard | 100 |  | 64 | 0.6 | 19.2 | 8.48 | 8.69 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Standard Z-sharp | 100 |  | 64 | 0.6 | 38.4 | 8.57 | 8.71 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Standard | 120 |  | 2 | 1 | 2 | 13.85 | 14.11 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Standard | 120 |  | 6 | 0.6 | 3.6 | 22.42 | 22.85 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Standard | 120 |  | 8 | 0.6 | 4.8 | 20.6 | 21 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Standard Z-sharp | 120 |  | 8 | 0.6 | 4.8 | 24.33 | 24.63 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Standard | 120 |  | 1 | 5 | 5 | 12.42 | 12.66 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Standard Z-sharp | 120 |  | 10 | 0.6 | 6 | 16.87 | 17.08 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Standard | 120 |  | 6 | 1.2 | 7.2 | 17.86 | 18.2 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Standard | 120 |  | 1 | 10 | 10 | 12.42 | 12.66 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Standard Z-sharp | 120 |  | 20 | 0.6 | 12 | 17.19 | 17.41 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Standard | 120 |  | 12 | 1.2 | 14.4 | 15.93 | 16.24 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Standard | 120 |  | 64 | 0.6 | 19.2 | 13.91 | 14.17 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Standard Z-sharp | 120 |  | 64 | 0.6 | 38.4 | 14.19 | 14.37 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Standard | 140 |  | 2 | 1 | 2 | 20.24 | 20.6 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Standard | 140 |  | 6 | 0.6 | 3.6 | 32.77 | 33.35 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Standard | 140 |  | 8 | 0.6 | 4.8 | 30.12 | 30.65 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Standard Z-sharp | 140 |  | 8 | 0.6 | 4.8 | 36.16 | 36.6 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Standard | 140 |  | 1 | 5 | 5 | 18.16 | 18.48 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Standard Z-sharp | 140 |  | 10 | 0.6 | 6 | 25.07 | 25.4 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Standard | 140 |  | 6 | 1.2 | 7.2 | 26.11 | 26.56 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Standard | 140 |  | 1 | 10 | 10 | 18.16 | 18.48 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Standard Z-sharp | 140 |  | 20 | 0.6 | 12 | 25.55 | 25.89 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Standard | 140 |  | 12 | 1.2 | 14.4 | 23.3 | 23.7 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Standard | 140 |  | 64 | 0.6 | 19.2 | 20.33 | 20.69 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Standard Z-sharp | 140 |  | 64 | 0.6 | 38.4 | 21.09 | 21.37 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Narrow | 70 |  | 2 | 1 | 2 | 2.36 | 2.36 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Narrow | 70 |  | 6 | 0.6 | 3.6 | 3.83 | 3.81 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Narrow | 70 |  | 8 | 0.6 | 4.8 | 3.52 | 3.51 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Narrow Z-sharp | 70 |  | 8 | 0.6 | 4.8 | 4.12 | 4.07 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Narrow | 70 |  | 1 | 5 | 5 | 2.12 | 2.11 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Narrow Z-sharp | 70 |  | 10 | 0.6 | 6 | 2.86 | 2.82 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Narrow | 70 |  | 6 | 1.2 | 7.2 | 3.05 | 3.04 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Narrow | 70 |  | 1 | 10 | 10 | 2.12 | 2.11 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Narrow Z-sharp | 70 |  | 20 | 0.6 | 12 | 2.91 | 2.88 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Narrow | 70 |  | 12 | 1.2 | 14.4 | 2.72 | 2.71 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Narrow | 70 |  | 64 | 0.6 | 19.2 | 2.38 | 2.37 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Narrow Z-sharp | 70 |  | 64 | 0.6 | 38.4 | 2.4 | 2.37 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Narrow | 80 |  | 2 | 1 | 2 | 3.81 | 3.78 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Narrow | 80 |  | 6 | 0.6 | 3.6 | 6.17 | 6.12 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Narrow | 80 |  | 8 | 0.6 | 4.8 | 5.67 | 5.63 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Narrow Z-sharp | 80 |  | 8 | 0.6 | 4.8 | 6.58 | 6.44 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Narrow | 80 |  | 1 | 5 | 5 | 3.42 | 3.39 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Narrow Z-sharp | 80 |  | 10 | 0.6 | 6 | 4.56 | 4.47 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Narrow | 80 |  | 6 | 1.2 | 7.2 | 4.92 | 4.88 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Narrow | 80 |  | 1 | 10 | 10 | 3.42 | 3.39 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Narrow Z-sharp | 80 |  | 20 | 0.6 | 12 | 4.65 | 4.55 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Narrow | 80 |  | 12 | 1.2 | 14.4 | 4.39 | 4.35 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Narrow | 80 |  | 64 | 0.6 | 19.2 | 3.83 | 3.8 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Narrow Z-sharp | 80 |  | 64 | 0.6 | 38.4 | 3.84 | 3.76 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Narrow | 100 |  | 2 | 1 | 2 | 7.81 | 7.69 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Narrow | 100 |  | 6 | 0.6 | 3.6 | 12.65 | 12.45 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Narrow | 100 |  | 8 | 0.6 | 4.8 | 11.63 | 11.44 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Narrow Z-sharp | 100 |  | 8 | 0.6 | 4.8 | 13.53 | 13.14 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Narrow | 100 |  | 1 | 5 | 5 | 7.01 | 6.9 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Narrow Z-sharp | 100 |  | 10 | 0.6 | 6 | 9.38 | 9.11 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Narrow | 100 |  | 6 | 1.2 | 7.2 | 10.08 | 9.92 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Narrow | 100 |  | 1 | 10 | 10 | 7.01 | 6.9 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Narrow Z-sharp | 100 |  | 20 | 0.6 | 12 | 9.56 | 9.29 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Narrow | 100 |  | 12 | 1.2 | 14.4 | 8.99 | 8.85 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Narrow | 100 |  | 64 | 0.6 | 19.2 | 7.85 | 7.72 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Narrow Z-sharp | 100 |  | 64 | 0.6 | 38.4 | 7.89 | 7.67 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Narrow | 120 |  | 2 | 1 | 2 | 12.93 | 12.68 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Narrow | 120 |  | 6 | 0.6 | 3.6 | 20.93 | 20.52 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Narrow | 120 |  | 8 | 0.6 | 4.8 | 19.23 | 18.86 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Narrow Z-sharp | 120 |  | 8 | 0.6 | 4.8 | 22.53 | 21.91 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Narrow | 120 |  | 1 | 5 | 5 | 11.6 | 11.37 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Narrow Z-sharp | 120 |  | 10 | 0.6 | 6 | 15.62 | 15.19 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Narrow | 120 |  | 6 | 1.2 | 7.2 | 16.67 | 16.35 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Narrow | 120 |  | 1 | 10 | 10 | 11.6 | 11.37 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Narrow Z-sharp | 120 |  | 20 | 0.6 | 12 | 15.92 | 15.48 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Narrow | 120 |  | 12 | 1.2 | 14.4 | 14.87 | 14.59 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Narrow | 120 |  | 64 | 0.6 | 19.2 | 12.98 | 12.73 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Narrow Z-sharp | 120 |  | 64 | 0.6 | 38.4 | 13.14 | 12.78 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Narrow | 140 |  | 2 | 1 | 2 | 18.83 | 18.5 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Narrow | 140 |  | 6 | 0.6 | 3.6 | 30.49 | 29.95 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Narrow | 140 |  | 8 | 0.6 | 4.8 | 28.02 | 27.53 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Narrow Z-sharp | 140 |  | 8 | 0.6 | 4.8 | 33.71 | 32.85 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Narrow | 140 |  | 1 | 5 | 5 | 16.89 | 16.6 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Narrow Z-sharp | 140 |  | 10 | 0.6 | 6 | 23.38 | 22.77 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Narrow | 140 |  | 6 | 1.2 | 7.2 | 24.29 | 23.86 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Narrow | 140 |  | 1 | 10 | 10 | 16.89 | 16.6 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Narrow Z-sharp | 140 |  | 20 | 0.6 | 12 | 23.82 | 23.21 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Narrow | 140 |  | 12 | 1.2 | 14.4 | 21.67 | 21.29 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Narrow | 140 |  | 64 | 0.6 | 19.2 | 18.91 | 18.58 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 16 | Narrow Z-sharp | 140 |  | 64 | 0.6 | 38.4 | 19.67 | 19.16 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Standard | 70 |  | 2 | 1 | 2 | 0.61 | 1.13 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Standard | 70 |  | 6 | 0.6 | 3.6 | 0.99 | 1.83 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Standard | 70 |  | 8 | 0.6 | 4.8 | 0.91 | 1.68 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Standard Z-sharp | 70 |  | 8 | 0.6 | 4.8 | 1.07 | 1.97 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Standard | 70 |  | 1 | 5 | 5 | 0.55 | 1.01 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Standard Z-sharp | 70 |  | 10 | 0.6 | 6 | 0.74 | 1.37 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Standard | 70 |  | 6 | 1.2 | 7.2 | 0.79 | 1.46 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Standard | 70 |  | 1 | 10 | 10 | 0.55 | 1.01 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Standard Z-sharp | 70 |  | 20 | 0.6 | 12 | 0.75 | 1.4 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Standard | 70 |  | 12 | 1.2 | 14.4 | 0.7 | 1.3 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Standard | 70 |  | 64 | 0.6 | 19.2 | 0.61 | 1.14 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Standard Z-sharp | 70 |  | 64 | 0.6 | 38.4 | 0.62 | 1.15 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Standard | 80 |  | 2 | 1 | 2 | 1.07 | 1.86 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Standard | 80 |  | 6 | 0.6 | 3.6 | 1.73 | 3.01 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Standard | 80 |  | 8 | 0.6 | 4.8 | 1.59 | 2.77 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Standard Z-sharp | 80 |  | 8 | 0.6 | 4.8 | 1.83 | 3.18 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Standard | 80 |  | 1 | 5 | 5 | 0.96 | 1.67 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Standard Z-sharp | 80 |  | 10 | 0.6 | 6 | 1.27 | 2.21 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Standard | 80 |  | 6 | 1.2 | 7.2 | 1.38 | 2.4 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Standard | 80 |  | 1 | 10 | 10 | 0.96 | 1.67 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Standard Z-sharp | 80 |  | 20 | 0.6 | 12 | 1.29 | 2.25 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Standard | 80 |  | 12 | 1.2 | 14.4 | 1.23 | 2.14 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Standard | 80 |  | 64 | 0.6 | 19.2 | 1.08 | 1.87 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Standard Z-sharp | 80 |  | 64 | 0.6 | 38.4 | 1.07 | 1.86 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Standard | 100 |  | 2 | 1 | 2 | 2.43 | 3.9 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Standard | 100 |  | 6 | 0.6 | 3.6 | 3.94 | 6.32 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Standard | 100 |  | 8 | 0.6 | 4.8 | 3.62 | 5.81 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Standard Z-sharp | 100 |  | 8 | 0.6 | 4.8 | 4.19 | 6.75 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Standard | 100 |  | 1 | 5 | 5 | 2.18 | 3.5 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Standard Z-sharp | 100 |  | 10 | 0.6 | 6 | 2.91 | 4.68 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Standard | 100 |  | 6 | 1.2 | 7.2 | 3.14 | 5.03 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Standard | 100 |  | 1 | 10 | 10 | 2.18 | 3.5 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Standard Z-sharp | 100 |  | 20 | 0.6 | 12 | 2.96 | 4.77 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Standard | 100 |  | 12 | 1.2 | 14.4 | 2.8 | 4.49 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Standard | 100 |  | 64 | 0.6 | 19.2 | 2.44 | 3.92 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Standard Z-sharp | 100 |  | 64 | 0.6 | 38.4 | 2.45 | 3.94 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Standard | 120 |  | 2 | 1 | 2 | 4.28 | 6.61 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Standard | 120 |  | 6 | 0.6 | 3.6 | 6.93 | 10.7 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Standard | 120 |  | 8 | 0.6 | 4.8 | 6.37 | 9.84 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Standard Z-sharp | 120 |  | 8 | 0.6 | 4.8 | 7.5 | 11.55 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Standard | 120 |  | 1 | 5 | 5 | 3.84 | 5.93 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Standard Z-sharp | 120 |  | 10 | 0.6 | 6 | 5 | 8.01 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Standard | 120 |  | 6 | 1.2 | 7.2 | 5.52 | 8.53 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Standard | 120 |  | 1 | 10 | 10 | 3.84 | 5.93 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Standard Z-sharp | 120 |  | 20 | 0.6 | 12 | 5.3 | 8.16 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Standard | 120 |  | 12 | 1.2 | 14.4 | 4.93 | 7.61 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Standard | 120 |  | 64 | 0.6 | 19.2 | 4.3 | 6.64 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Standard Z-sharp | 120 |  | 64 | 0.6 | 38.4 | 4.37 | 6.74 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Standard | 140 |  | 2 | 1 | 2 | 6.57 | 9.94 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Standard | 140 |  | 6 | 0.6 | 3.6 | 10.64 | 16.09 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Standard | 140 |  | 8 | 0.6 | 4.8 | 9.78 | 14.79 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Standard Z-sharp | 140 |  | 8 | 0.6 | 4.8 | 11.59 | 17.62 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Standard | 140 |  | 1 | 5 | 5 | 5.89 | 8.91 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Standard Z-sharp | 140 |  | 10 | 0.6 | 6 | 8.03 | 12.22 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Standard | 140 |  | 6 | 1.2 | 7.2 | 8.47 | 12.81 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Standard | 140 |  | 1 | 10 | 10 | 5.89 | 8.91 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Standard Z-sharp | 140 |  | 20 | 0.6 | 12 | 8.19 | 12.45 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Standard | 140 |  | 12 | 1.2 | 14.4 | 7.56 | 11.43 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Standard | 140 |  | 64 | 0.6 | 19.2 | 6.6 | 9.98 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Standard Z-sharp | 140 |  | 64 | 0.6 | 38.4 | 6.76 | 10.28 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Narrow | 70 |  | 2 | 1 | 2 | 0.54 | 0.9 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Narrow | 70 |  | 6 | 0.6 | 3.6 | 0.87 | 1.45 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Narrow | 70 |  | 8 | 0.6 | 4.8 | 0.8 | 1.33 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Narrow Z-sharp | 70 |  | 8 | 0.6 | 4.8 | 0.93 | 1.54 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Narrow | 70 |  | 1 | 5 | 5 | 0.48 | 0.8 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Narrow Z-sharp | 70 |  | 10 | 0.6 | 6 | 0.65 | 1.07 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Narrow | 70 |  | 6 | 1.2 | 7.2 | 0.69 | 1.16 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Narrow | 70 |  | 1 | 10 | 10 | 0.48 | 0.8 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Narrow Z-sharp | 70 |  | 20 | 0.6 | 12 | 0.66 | 1.09 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Narrow | 70 |  | 12 | 1.2 | 14.4 | 0.62 | 1.03 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Narrow | 70 |  | 64 | 0.6 | 19.2 | 0.54 | 0.9 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Narrow Z-sharp | 70 |  | 64 | 0.6 | 38.4 | 0.55 | 0.9 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Narrow | 80 |  | 2 | 1 | 2 | 0.94 | 1.46 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Narrow | 80 |  | 6 | 0.6 | 3.6 | 1.52 | 2.36 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Narrow | 80 |  | 8 | 0.6 | 4.8 | 1.39 | 2.17 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Narrow Z-sharp | 80 |  | 8 | 0.6 | 4.8 | 1.62 | 2.48 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Narrow | 80 |  | 1 | 5 | 5 | 0.84 | 1.31 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Narrow Z-sharp | 80 |  | 10 | 0.6 | 6 | 1.12 | 1.72 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Narrow | 80 |  | 6 | 1.2 | 7.2 | 1.21 | 1.88 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Narrow | 80 |  | 1 | 10 | 10 | 0.84 | 1.31 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Narrow Z-sharp | 80 |  | 20 | 0.6 | 12 | 1.15 | 1.76 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Narrow | 80 |  | 12 | 1.2 | 14.4 | 1.08 | 1.68 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Narrow | 80 |  | 64 | 0.6 | 19.2 | 0.94 | 1.47 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Narrow Z-sharp | 80 |  | 64 | 0.6 | 38.4 | 0.95 | 1.45 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Narrow | 100 |  | 2 | 1 | 2 | 2.17 | 3.12 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Narrow | 100 |  | 6 | 0.6 | 3.6 | 3.51 | 5.06 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Narrow | 100 |  | 8 | 0.6 | 4.8 | 3.22 | 4.65 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Narrow Z-sharp | 100 |  | 8 | 0.6 | 4.8 | 3.75 | 5.34 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Narrow | 100 |  | 1 | 5 | 5 | 1.94 | 2.8 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Narrow Z-sharp | 100 |  | 10 | 0.6 | 6 | 2.6 | 3.7 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Narrow | 100 |  | 6 | 1.2 | 7.2 | 2.79 | 4.03 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Narrow | 100 |  | 1 | 10 | 10 | 1.94 | 2.8 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Narrow Z-sharp | 100 |  | 20 | 0.6 | 12 | 2.65 | 3.78 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Narrow | 100 |  | 12 | 1.2 | 14.4 | 2.49 | 3.59 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Narrow | 100 |  | 64 | 0.6 | 19.2 | 2.17 | 3.14 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Narrow Z-sharp | 100 |  | 64 | 0.6 | 38.4 | 2.19 | 3.12 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Narrow | 120 |  | 2 | 1 | 2 | 3.84 | 5.36 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Narrow | 120 |  | 6 | 0.6 | 3.6 | 6.22 | 8.68 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Narrow | 120 |  | 8 | 0.6 | 4.8 | 5.72 | 7.98 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Narrow Z-sharp | 120 |  | 8 | 0.6 | 4.8 | 6.67 | 9.21 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Narrow | 120 |  | 1 | 5 | 5 | 3.45 | 4.81 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Narrow Z-sharp | 120 |  | 10 | 0.6 | 6 | 4.62 | 6.39 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Narrow | 120 |  | 6 | 1.2 | 7.2 | 4.95 | 6.91 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Narrow | 120 |  | 1 | 10 | 10 | 3.45 | 4.81 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Narrow Z-sharp | 120 |  | 20 | 0.6 | 12 | 4.71 | 6.51 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Narrow | 120 |  | 12 | 1.2 | 14.4 | 4.42 | 6.17 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Narrow | 120 |  | 64 | 0.6 | 19.2 | 3.86 | 5.39 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Narrow Z-sharp | 120 |  | 64 | 0.6 | 38.4 | 3.89 | 5.38 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Narrow | 140 |  | 2 | 1 | 2 | 5.89 | 8.03 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Narrow | 140 |  | 6 | 0.6 | 3.6 | 9.53 | 13 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Narrow | 140 |  | 8 | 0.6 | 4.8 | 8.76 | 11.95 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Narrow Z-sharp | 140 |  | 8 | 0.6 | 4.8 | 10.45 | 14.2 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Narrow | 140 |  | 1 | 5 | 5 | 5.28 | 7.2 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Narrow Z-sharp | 140 |  | 10 | 0.6 | 6 | 7.25 | 9.84 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Narrow | 140 |  | 6 | 1.2 | 7.2 | 7.59 | 10.36 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Narrow | 140 |  | 1 | 10 | 10 | 5.28 | 7.2 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Narrow Z-sharp | 140 |  | 20 | 0.6 | 12 | 7.38 | 10.03 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Narrow | 140 |  | 12 | 1.2 | 14.4 | 6.77 | 9.24 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Narrow | 140 |  | 64 | 0.6 | 19.2 | 5.91 | 8.07 | Biograph Vision | Siemens
Biograph Vision | 11 | 1 | 32 | Narrow Z-sharp | 140 |  | 64 | 0.6 | 38.4 | 6.1 | 8.28 | Biograph Vision | Siemens
X.cite | 10 | 1 | 16 | Standard | 70 |  | 6 | 0.6 | 3.6 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 70 |  | 1 | 5 | 5 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 70 |  | 1 | 10 | 10 | 3.16 | 3.4 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 70 |  | 3 | 3 | 9 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 70 |  | 3 | 5 | 15 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard Z-sharp | 70 |  | 32 | 0.6 | 19.2 | 4.47 | 4.8 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 70 |  | 64 | 0.6 | 38.4 | 3.93 | 4.24 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard Z-sharp | 70 |  | 128 | 0.6 | 76.8 | 3.91 | 4.21 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 80 |  | 6 | 0.6 | 3.6 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 80 |  | 1 | 5 | 5 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 80 |  | 1 | 10 | 10 | 4.87 | 5.14 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 80 |  | 3 | 3 | 9 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 80 |  | 3 | 5 | 15 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard Z-sharp | 80 |  | 32 | 0.6 | 19.2 | 6.88 | 7.25 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 80 |  | 64 | 0.6 | 38.4 | 6.05 | 6.42 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard Z-sharp | 80 |  | 128 | 0.6 | 76.8 | 6.02 | 6.37 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 90 |  | 6 | 0.6 | 3.6 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 90 |  | 1 | 5 | 5 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 90 |  | 1 | 10 | 10 | 6.88 | 7.19 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 90 |  | 3 | 3 | 9 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 90 |  | 3 | 5 | 15 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard Z-sharp | 90 |  | 32 | 0.6 | 19.2 | 9.73 | 10.15 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 90 |  | 64 | 0.6 | 38.4 | 8.56 | 8.97 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard Z-sharp | 90 |  | 128 | 0.6 | 76.8 | 8.52 | 8.9 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 100 |  | 6 | 0.6 | 3.6 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 100 |  | 1 | 5 | 5 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 100 |  | 1 | 10 | 10 | 9.15 | 9.49 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 100 |  | 3 | 3 | 9 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 100 |  | 3 | 5 | 15 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard Z-sharp | 100 |  | 32 | 0.6 | 19.2 | 12.95 | 13.39 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 100 |  | 64 | 0.6 | 38.4 | 11.38 | 11.84 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard Z-sharp | 100 |  | 128 | 0.6 | 76.8 | 11.33 | 11.75 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 110 |  | 6 | 0.6 | 3.6 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 110 |  | 1 | 5 | 5 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 110 |  | 1 | 10 | 10 | 11.63 | 12.01 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 110 |  | 3 | 3 | 9 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 110 |  | 3 | 5 | 15 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard Z-sharp | 110 |  | 32 | 0.6 | 19.2 | 16.45 | 16.94 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 110 |  | 64 | 0.6 | 38.4 | 14.46 | 14.98 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard Z-sharp | 110 |  | 128 | 0.6 | 76.8 | 14.4 | 14.86 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 120 |  | 6 | 0.6 | 3.6 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 120 |  | 1 | 5 | 5 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 120 |  | 1 | 10 | 10 | 14.33 | 14.71 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 120 |  | 3 | 3 | 9 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 120 |  | 3 | 5 | 15 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard Z-sharp | 120 |  | 32 | 0.6 | 19.2 | 20.26 | 20.75 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 120 |  | 64 | 0.6 | 38.4 | 17.81 | 18.36 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard Z-sharp | 120 |  | 128 | 0.6 | 76.8 | 17.73 | 18.21 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 130 |  | 6 | 0.6 | 3.6 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 130 |  | 1 | 5 | 5 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 130 |  | 1 | 10 | 10 | 17.15 | 17.6 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 130 |  | 3 | 3 | 9 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 130 |  | 3 | 5 | 15 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard Z-sharp | 130 |  | 32 | 0.6 | 19.2 | 24.27 | 24.83 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 130 |  | 64 | 0.6 | 38.4 | 21.33 | 21.96 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard Z-sharp | 130 |  | 128 | 0.6 | 76.8 | 21.24 | 21.79 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 140 |  | 6 | 0.6 | 3.6 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 140 |  | 1 | 5 | 5 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 140 |  | 1 | 10 | 10 | 20.15 | 20.65 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 140 |  | 3 | 3 | 9 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 140 |  | 3 | 5 | 15 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard Z-sharp | 140 |  | 32 | 0.6 | 19.2 | 28.51 | 29.14 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 140 |  | 64 | 0.6 | 38.4 | 25.06 | 25.77 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard Z-sharp | 140 |  | 128 | 0.6 | 76.8 | 24.95 | 25.57 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 150 |  | 6 | 0.6 | 3.6 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 150 |  | 1 | 5 | 5 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 150 |  | 1 | 10 | 10 | 23.36 | 23.89 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 150 |  | 3 | 3 | 9 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 150 |  | 3 | 5 | 15 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard Z-sharp | 150 |  | 32 | 0.6 | 19.2 | 33.04 | 33.7 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 150 |  | 64 | 0.6 | 38.4 | 29.05 | 29.81 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard Z-sharp | 150 |  | 128 | 0.6 | 76.8 | 28.92 | 29.57 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 100 | SnLD | 6 | 0.6 | 3.6 | 2.7 | 2.69 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 100 |  | 1 | 5 | 5 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 100 |  | 1 | 10 | 10 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 100 |  | 3 | 3 | 9 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 100 |  | 3 | 5 | 15 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard Z-sharp | 100 |  | 32 | 0.6 | 19.2 | 2.08 | 2.08 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 100 |  | 64 | 0.6 | 38.4 | 1.83 | 1.84 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard Z-sharp | 100 |  | 128 | 0.6 | 76.8 | 1.82 | 1.83 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 110 |  | 6 | 0.6 | 3.6 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 110 |  | 1 | 5 | 5 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 110 |  | 1 | 10 | 10 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 110 |  | 3 | 3 | 9 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 110 |  | 3 | 5 | 15 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard Z-sharp | 110 |  | 32 | 0.6 | 19.2 | 3.19 | 3.18 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 110 |  | 64 | 0.6 | 38.4 | 2.8 | 2.821 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard Z-sharp | 110 |  | 128 | 0.6 | 76.8 | 2.79 | 2.8 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 120 |  | 6 | 0.6 | 3.6 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 120 |  | 1 | 5 | 5 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 120 |  | 1 | 10 | 10 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 120 |  | 3 | 3 | 9 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 120 |  | 3 | 5 | 15 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard Z-sharp | 120 |  | 32 | 0.6 | 19.2 | 4.55 | 4.55 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 120 |  | 64 | 0.6 | 38.4 | 4 | 4.03 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard Z-sharp | 120 |  | 128 | 0.6 | 76.8 | 3.97 | 4 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 130 |  | 6 | 0.6 | 3.6 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 130 |  | 1 | 5 | 5 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 130 |  | 1 | 10 | 10 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 130 |  | 3 | 3 | 9 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 130 |  | 3 | 5 | 15 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard Z-sharp | 130 |  | 32 | 0.6 | 19.2 | 6.17 | 6.19 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 130 |  | 64 | 0.6 | 38.4 | 5.43 | 5.49 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard Z-sharp | 130 |  | 128 | 0.6 | 76.8 | 5.39 | 5.44 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 140 |  | 6 | 0.6 | 3.6 | 10.39 | 10.47 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 140 |  | 1 | 5 | 5 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 140 |  | 1 | 10 | 10 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 140 |  | 3 | 3 | 9 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 140 |  | 3 | 5 | 15 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard Z-sharp | 140 |  | 32 | 0.6 | 19.2 | 8.05 | 8.1 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 140 |  | 64 | 0.6 | 38.4 | 7.09 | 7.18 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard Z-sharp | 140 |  | 128 | 0.6 | 76.8 | 7.04 | 7.12 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 150 |  | 6 | 0.6 | 3.6 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 150 |  | 1 | 5 | 5 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 150 |  | 1 | 10 | 10 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 150 |  | 3 | 3 | 9 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 150 |  | 3 | 5 | 15 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard Z-sharp | 150 |  | 32 | 0.6 | 19.2 | 10.19 | 10.28 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 150 |  | 64 | 0.6 | 38.4 | 8.97 | 9.11 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard Z-sharp | 150 |  | 128 | 0.6 | 76.8 | 8.91 | 9.04 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 150 | SnDE | 6 | 0.6 | 3.6 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 150 |  | 1 | 5 | 5 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 150 |  | 1 | 10 | 10 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 150 |  | 3 | 3 | 9 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 150 |  | 3 | 5 | 15 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard Z-sharp | 150 |  | 32 | 0.6 | 19.2 | 6.01 | 6.1 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 150 |  | 64 | 0.6 | 38.4 | 5.3 | 5.41 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard Z-sharp | 150 |  | 128 | 0.6 | 76.8 | 5.27 | 6.36 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 120 | AuSn | 6 | 0.6 | 3.6 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 120 |  | 1 | 5 | 5 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 120 |  | 1 | 10 | 10 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 120 |  | 3 | 3 | 9 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 120 |  | 3 | 5 | 15 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard Z-sharp | 120 |  | 32 | 0.6 | 19.2 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 120 |  | 64 | 0.6 | 38.4 | 4.09 | 4.11 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard Z-sharp | 120 |  | 128 | 0.6 | 76.8 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 140 | AuSn | 6 | 0.6 | 3.6 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 140 |  | 1 | 5 | 5 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 140 |  | 1 | 10 | 10 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 140 |  | 3 | 3 | 9 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 140 |  | 3 | 5 | 15 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard Z-sharp | 140 |  | 32 | 0.6 | 19.2 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard | 140 |  | 64 | 0.6 | 38.4 | 6.73 | 6.8 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Standard Z-sharp | 140 |  | 128 | 0.6 | 76.8 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 70 |  | 6 | 0.6 | 3.6 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 70 |  | 1 | 5 | 5 | 3.1 | 3.12 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 70 |  | 1 | 10 | 10 | 2.82 | 2.84 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 70 |  | 3 | 3 | 9 | 3.93 | 4 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 70 |  | 3 | 5 | 15 | 3.75 | 3.81 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow Z-sharp | 70 |  | 32 | 0.6 | 19.2 | 3.98 | 4.01 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 70 |  | 64 | 0.6 | 38.4 | 3.5 | 3.55 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow Z-sharp | 70 |  | 128 | 0.6 | 76.8 | 3.48 | 3.52 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 80 |  | 6 | 0.6 | 3.6 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 80 |  | 1 | 5 | 5 | 4.83 | 4.8 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 80 |  | 1 | 10 | 10 | 4.38 | 4.37 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 80 |  | 3 | 3 | 9 | 6.13 | 6.15 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 80 |  | 3 | 5 | 15 | 5.85 | 5.86 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow Z-sharp | 80 |  | 32 | 0.6 | 19.2 | 6.19 | 6.16 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 80 |  | 64 | 0.6 | 38.4 | 5.45 | 5.46 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow Z-sharp | 80 |  | 128 | 0.6 | 76.8 | 5.42 | 5.41 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 90 |  | 6 | 0.6 | 3.6 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 90 |  | 1 | 5 | 5 | 6.89 | 6.8 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 90 |  | 1 | 10 | 10 | 6.26 | 6.19 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 90 |  | 3 | 3 | 9 | 8.75 | 8.72 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 90 |  | 3 | 5 | 15 | 8.34 | 8.3 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow Z-sharp | 90 |  | 32 | 0.6 | 19.2 | 8.84 | 8.73 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 90 |  | 64 | 0.6 | 38.4 | 7.78 | 7.74 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow Z-sharp | 90 |  | 128 | 0.6 | 76.8 | 7.74 | 7.67 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 100 |  | 6 | 0.6 | 3.6 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 100 |  | 1 | 5 | 5 | 9.22 | 9.06 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 100 |  | 1 | 10 | 10 | 8.36 | 8.24 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 100 |  | 3 | 3 | 9 | 11.69 | 11.61 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 100 |  | 3 | 5 | 15 | 11.15 | 11.06 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow Z-sharp | 100 |  | 32 | 0.6 | 19.2 | 11.81 | 11.63 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 100 |  | 64 | 0.6 | 38.4 | 10.4 | 10.3 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow Z-sharp | 100 |  | 128 | 0.6 | 76.8 | 10.35 | 10.21 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 110 |  | 6 | 0.6 | 3.6 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 110 |  | 1 | 5 | 5 | 11.77 | 11.53 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 110 |  | 1 | 10 | 10 | 10.69 | 10.5 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 110 |  | 3 | 3 | 9 | 14.94 | 14.79 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 110 |  | 3 | 5 | 15 | 14.23 | 14.09 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow Z-sharp | 110 |  | 32 | 0.6 | 19.2 | 15.1 | 14.81 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 110 |  | 64 | 0.6 | 38.4 | 13.29 | 13.13 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow Z-sharp | 110 |  | 128 | 0.6 | 76.8 | 13.23 | 13.01 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 120 |  | 6 | 0.6 | 3.6 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 120 |  | 1 | 5 | 5 | 14.55 | 14.23 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 120 |  | 1 | 10 | 10 | 13.2 | 12.95 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 120 |  | 3 | 3 | 9 | 18.44 | 18.24 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 120 |  | 3 | 5 | 15 | 17.59 | 17.38 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow Z-sharp | 120 |  | 32 | 0.6 | 19.2 | 18.64 | 18.26 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 120 |  | 64 | 0.6 | 38.4 | 16.41 | 16.19 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow Z-sharp | 120 |  | 128 | 0.6 | 76.8 | 16.33 | 16.04 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 130 |  | 6 | 0.6 | 3.6 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 130 |  | 1 | 5 | 5 | 17.51 | 17.11 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 130 |  | 1 | 10 | 10 | 15.87 | 15.57 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 130 |  | 3 | 3 | 9 | 22.17 | 21.93 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 130 |  | 3 | 5 | 15 | 21.17 | 20.9 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow Z-sharp | 130 |  | 32 | 0.6 | 19.2 | 22.41 | 21.96 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 130 |  | 64 | 0.6 | 38.4 | 19.73 | 19.46 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow Z-sharp | 130 |  | 128 | 0.6 | 76.8 | 19.64 | 19.29 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 140 |  | 6 | 0.6 | 3.6 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 140 |  | 1 | 5 | 5 | 20.65 | 20.18 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 140 |  | 1 | 10 | 10 | 18.71 | 18.36 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 140 |  | 3 | 3 | 9 | 26.14 | 25.86 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 140 |  | 3 | 5 | 15 | 24.97 | 24.65 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow Z-sharp | 140 |  | 32 | 0.6 | 19.2 | 26.42 | 25.89 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 140 |  | 64 | 0.6 | 38.4 | 23.26 | 22.95 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow Z-sharp | 140 |  | 128 | 0.6 | 76.8 | 23.15 | 22.74 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 150 |  | 6 | 0.6 | 3.6 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 150 |  | 1 | 5 | 5 | 23.98 | 23.43 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 150 |  | 1 | 10 | 10 | 21.71 | 21.31 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 150 |  | 3 | 3 | 9 | 30.33 | 30.02 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 150 |  | 3 | 5 | 15 | 28.92 | 28.6 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow Z-sharp | 150 |  | 32 | 0.6 | 19.2 | 30.66 | 30.06 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 150 |  | 64 | 0.6 | 38.4 | 26.99 | 26.64 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow Z-sharp | 150 |  | 128 | 0.6 | 76.8 | 26.86 | 26.4 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 100 | SnLD | 6 | 0.6 | 3.6 | 2.53 | 2.45 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 100 |  | 1 | 5 | 5 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 100 |  | 1 | 10 | 10 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 100 |  | 3 | 3 | 9 | 1094 | 1.89 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 100 |  | 3 | 5 | 15 | 1.84 | 1.8 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow Z-sharp | 100 |  | 32 | 0.6 | 19.2 | 1.96 | 1.9 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 100 |  | 64 | 0.6 | 38.4 | 1.72 | 1.68 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow Z-sharp | 100 |  | 128 | 0.6 | 76.8 | 1.71 | 1.66 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 110 |  | 6 | 0.6 | 3.6 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 110 |  | 1 | 5 | 5 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 110 |  | 1 | 10 | 10 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 110 |  | 3 | 3 | 9 | 2.97 | 2.91 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 110 |  | 3 | 5 | 15 | 2.82 | 2.77 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow Z-sharp | 110 |  | 32 | 0.6 | 19.2 | 3 | 2.92 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 110 |  | 64 | 0.6 | 38.4 | 2.64 | 2.58 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow Z-sharp | 110 |  | 128 | 0.6 | 76.8 | 2.62 | 2.56 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 120 |  | 6 | 0.6 | 3.6 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 120 |  | 1 | 5 | 5 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 120 |  | 1 | 10 | 10 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 120 |  | 3 | 3 | 9 | 4.25 | 4.17 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 120 |  | 3 | 5 | 15 | 4.04 | 3.97 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow Z-sharp | 120 |  | 32 | 0.6 | 19.2 | 4.3 | 4.19 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 120 |  | 64 | 0.6 | 38.4 | 3.78 | 3.7 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow Z-sharp | 120 |  | 128 | 0.6 | 76.8 | 3.76 | 3.67 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 130 |  | 6 | 0.6 | 3.6 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 130 |  | 1 | 5 | 5 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 130 |  | 1 | 10 | 10 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 130 |  | 3 | 3 | 9 | 5.77 | 5.69 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 130 |  | 3 | 5 | 15 | 5.49 | 5.42 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow Z-sharp | 130 |  | 32 | 0.6 | 19.2 | 5.84 | 5.71 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 130 |  | 64 | 0.6 | 38.4 | 4.14 | 5.05 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow Z-sharp | 130 |  | 128 | 0.6 | 76.8 | 5.11 | 5.01 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 140 |  | 6 | 0.6 | 3.6 | 9.87 | 9.68 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 140 |  | 1 | 5 | 5 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 140 |  | 1 | 10 | 10 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 140 |  | 3 | 3 | 9 | 7.55 | 7.48 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 140 |  | 3 | 5 | 15 | 7.18 | 7.12 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow Z-sharp | 140 |  | 32 | 0.6 | 19.2 | 7.64 | 7.5 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 140 |  | 64 | 0.6 | 38.4 | 6.73 | 6.63 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow Z-sharp | 140 |  | 128 | 0.6 | 76.8 | 6.68 | 6.58 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 150 |  | 6 | 0.6 | 3.6 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 150 |  | 1 | 5 | 5 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 150 |  | 1 | 10 | 10 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 150 |  | 3 | 3 | 9 | 9.58 | 9.51 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 150 |  | 3 | 5 | 15 | 9.11 | 9.05 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow Z-sharp | 150 |  | 32 | 0.6 | 19.2 | 9.69 | 9.55 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 150 |  | 64 | 0.6 | 38.4 | 8.53 | 8.43 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow Z-sharp | 150 |  | 128 | 0.6 | 76.8 | 8.47 | 8.37 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 150 | SnDE | 6 | 0.6 | 3.6 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 150 |  | 1 | 5 | 5 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 150 |  | 1 | 10 | 10 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 150 |  | 3 | 3 | 9 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 150 |  | 3 | 5 | 15 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow Z-sharp | 150 |  | 32 | 0.6 | 19.2 | 5.73 | 5.67 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 150 |  | 64 | 0.6 | 38.4 | 5.06 | 5.02 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow Z-sharp | 150 |  | 128 | 0.6 | 76.8 | 5.02 | 4.98 | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 120 | AuSn | 6 | 0.6 | 3.6 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 120 |  | 1 | 5 | 5 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 120 |  | 1 | 10 | 10 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 120 |  | 3 | 3 | 9 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 120 |  | 3 | 5 | 15 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow Z-sharp | 120 |  | 32 | 0.6 | 19.2 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 120 |  | 64 | 0.6 | 38.4 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow Z-sharp | 120 |  | 128 | 0.6 | 76.8 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 140 | AuSn | 6 | 0.6 | 3.6 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 140 |  | 1 | 5 | 5 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 140 |  | 1 | 10 | 10 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 140 |  | 3 | 3 | 9 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 140 |  | 3 | 5 | 15 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow Z-sharp | 140 |  | 32 | 0.6 | 19.2 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow | 140 |  | 64 | 0.6 | 38.4 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 16 | Narrow Z-sharp | 140 |  | 128 | 0.6 | 76.8 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 70 |  | 6 | 0.6 | 3.6 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 70 |  | 1 | 5 | 5 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 70 |  | 1 | 10 | 10 | 0.69 | 1.37 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 70 |  | 3 | 3 | 9 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 70 |  | 3 | 5 | 15 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard Z-sharp | 70 |  | 32 | 0.6 | 19.2 | 0.97 | 1.94 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 70 |  | 64 | 0.6 | 38.4 | 0.86 | 1.72 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard Z-sharp | 70 |  | 128 | 0.6 | 76.8 | 0.85 | 1.7 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 80 |  | 6 | 0.6 | 3.6 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 80 |  | 1 | 5 | 5 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 80 |  | 1 | 10 | 10 | 1.18 | 2.16 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 80 |  | 3 | 3 | 9 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 80 |  | 3 | 5 | 15 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard Z-sharp | 80 |  | 32 | 0.6 | 19.2 | 1.66 | 3.05 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 80 |  | 64 | 0.6 | 38.4 | 1.46 | 2.7 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard Z-sharp | 80 |  | 128 | 0.6 | 76.8 | 1.45 | 2.68 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 90 |  | 6 | 0.6 | 3.6 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 90 |  | 1 | 5 | 5 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 90 |  | 1 | 10 | 10 | 1.79 | 3.11 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 90 |  | 3 | 3 | 9 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 90 |  | 3 | 5 | 15 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard Z-sharp | 90 |  | 32 | 0.6 | 19.2 | 2.52 | 4.39 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 90 |  | 64 | 0.6 | 38.4 | 2.21 | 3.89 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard Z-sharp | 90 |  | 128 | 0.6 | 76.8 | 2.2 | 3.86 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 100 |  | 6 | 0.6 | 3.6 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 100 |  | 1 | 5 | 5 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 100 |  | 1 | 10 | 10 | 2.51 | 4.21 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 100 |  | 3 | 3 | 9 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 100 |  | 3 | 5 | 15 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard Z-sharp | 100 |  | 32 | 0.6 | 19.2 | 3.54 | 5.95 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 100 |  | 64 | 0.6 | 38.4 | 3.1 | 5.27 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard Z-sharp | 100 |  | 128 | 0.6 | 76.8 | 3.09 | 5.23 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 110 |  | 6 | 0.6 | 3.6 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 110 |  | 1 | 5 | 5 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 110 |  | 1 | 10 | 10 | 3.33 | 5.44 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 110 |  | 3 | 3 | 9 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 110 |  | 3 | 5 | 15 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard Z-sharp | 110 |  | 32 | 0.6 | 19.2 | 4.68 | 7.69 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 110 |  | 64 | 0.6 | 38.4 | 4.12 | 6.8 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard Z-sharp | 110 |  | 128 | 0.6 | 76.8 | 4.09 | 6.75 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 120 |  | 6 | 0.6 | 3.6 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 120 |  | 1 | 5 | 5 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 120 |  | 1 | 10 | 10 | 4.23 | 6.8 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 120 |  | 3 | 3 | 9 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 120 |  | 3 | 5 | 15 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard Z-sharp | 120 |  | 32 | 0.6 | 19.2 | 5.97 | 9.61 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 120 |  | 64 | 0.6 | 38.4 | 5.24 | 8.49 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard Z-sharp | 120 |  | 128 | 0.6 | 76.8 | 5.2 | 8.42 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 130 |  | 6 | 0.6 | 3.6 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 130 |  | 1 | 5 | 5 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 130 |  | 1 | 10 | 10 | 5.22 | 8.26 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 130 |  | 3 | 3 | 9 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 130 |  | 3 | 5 | 15 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard Z-sharp | 130 |  | 32 | 0.6 | 19.2 | 7.37 | 11.66 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 130 |  | 64 | 0.6 | 38.4 | 6.46 | 10.31 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard Z-sharp | 130 |  | 128 | 0.6 | 76.8 | 6.42 | 10.23 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 140 |  | 6 | 0.6 | 3.6 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 140 |  | 1 | 5 | 5 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 140 |  | 1 | 10 | 10 | 6.29 | 9.82 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 140 |  | 3 | 3 | 9 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 140 |  | 3 | 5 | 15 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard Z-sharp | 140 |  | 32 | 0.6 | 19.2 | 8.87 | 13.88 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 140 |  | 64 | 0.6 | 38.4 | 7.8 | 12.27 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard Z-sharp | 140 |  | 128 | 0.6 | 76.8 | 7.73 | 12.15 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 150 |  | 6 | 0.6 | 3.6 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 150 |  | 1 | 5 | 5 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 150 |  | 1 | 10 | 10 | 7.43 | 11.51 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 150 |  | 3 | 3 | 9 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 150 |  | 3 | 5 | 15 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard Z-sharp | 150 |  | 32 | 0.6 | 19.2 | 10.49 | 16.23 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 150 |  | 64 | 0.6 | 38.4 | 9.23 | 14.37 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard Z-sharp | 150 |  | 128 | 0.6 | 76.8 | 9.15 | 14.23 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 100 | SnLD | 6 | 0.6 | 3.6 | 0.92 | 1.35 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 100 |  | 1 | 5 | 5 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 100 |  | 1 | 10 | 10 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 100 |  | 3 | 3 | 9 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 100 |  | 3 | 5 | 15 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard Z-sharp | 100 |  | 32 | 0.6 | 19.2 | 0.71 | 1.05 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 100 |  | 64 | 0.6 | 38.4 | 0.63 | 0.93 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard Z-sharp | 100 |  | 128 | 0.6 | 76.8 | 0.62 | 0.92 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 110 |  | 6 | 0.6 | 3.6 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 110 |  | 1 | 5 | 5 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 110 |  | 1 | 10 | 10 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 110 |  | 3 | 3 | 9 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 110 |  | 3 | 5 | 15 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard Z-sharp | 110 |  | 32 | 0.6 | 19.2 | 1.12 | 1.63 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 110 |  | 64 | 0.6 | 38.4 | 0.99 | 1.44 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard Z-sharp | 110 |  | 128 | 0.6 | 76.8 | 0.98 | 1.43 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 120 |  | 6 | 0.6 | 3.6 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 120 |  | 1 | 5 | 5 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 120 |  | 1 | 10 | 10 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 120 |  | 3 | 3 | 9 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 120 |  | 3 | 5 | 15 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard Z-sharp | 120 |  | 32 | 0.6 | 19.2 | 1.63 | 2.36 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 120 |  | 64 | 0.6 | 38.4 | 1.44 | 2.09 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard Z-sharp | 120 |  | 128 | 0.6 | 76.8 | 1.43 | 2.07 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 130 |  | 6 | 0.6 | 3.6 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 130 |  | 1 | 5 | 5 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 130 |  | 1 | 10 | 10 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 130 |  | 3 | 3 | 9 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 130 |  | 3 | 5 | 15 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard Z-sharp | 130 |  | 32 | 0.6 | 19.2 | 2.26 | 3.25 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 130 |  | 64 | 0.6 | 38.4 | 1.99 | 2.88 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard Z-sharp | 130 |  | 128 | 0.6 | 76.8 | 1.98 | 2.85 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 140 |  | 6 | 0.6 | 3.6 | 3.86 | 5.54 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 140 |  | 1 | 5 | 5 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 140 |  | 1 | 10 | 10 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 140 |  | 3 | 3 | 9 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 140 |  | 3 | 5 | 15 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard Z-sharp | 140 |  | 32 | 0.6 | 19.2 | 3 | 4.3 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 140 |  | 64 | 0.6 | 38.4 | 2.63 | 3.81 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard Z-sharp | 140 |  | 128 | 0.6 | 76.8 | 2.62 | 3.77 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 150 |  | 6 | 0.6 | 3.6 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 150 |  | 1 | 5 | 5 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 150 |  | 1 | 10 | 10 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 150 |  | 3 | 3 | 9 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 150 |  | 3 | 5 | 15 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard Z-sharp | 150 |  | 32 | 0.6 | 19.2 | 3.84 | 5.5 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 150 |  | 64 | 0.6 | 38.4 | 3.38 | 4.87 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard Z-sharp | 150 |  | 128 | 0.6 | 76.8 | 3.36 | 4.83 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 150 | SnDE | 6 | 0.6 | 3.6 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 150 |  | 1 | 5 | 5 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 150 |  | 1 | 10 | 10 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 150 |  | 3 | 3 | 9 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 150 |  | 3 | 5 | 15 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard Z-sharp | 150 |  | 32 | 0.6 | 19.2 | 2.34 | 3.32 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 150 |  | 64 | 0.6 | 38.4 | 2.06 | 2.95 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard Z-sharp | 150 |  | 128 | 0.6 | 76.8 | 2.06 | 2.93 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 120 | AuSn | 6 | 0.6 | 3.6 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 120 |  | 1 | 5 | 5 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 120 |  | 1 | 10 | 10 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 120 |  | 3 | 3 | 9 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 120 |  | 3 | 5 | 15 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard Z-sharp | 120 |  | 32 | 0.6 | 19.2 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 120 |  | 64 | 0.6 | 38.4 | 1.38 | 2.06 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard Z-sharp | 120 |  | 128 | 0.6 | 76.8 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 140 | AuSn | 6 | 0.6 | 3.6 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 140 |  | 1 | 5 | 5 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 140 |  | 1 | 10 | 10 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 140 |  | 3 | 3 | 9 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 140 |  | 3 | 5 | 15 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard Z-sharp | 140 |  | 32 | 0.6 | 19.2 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard | 140 |  | 64 | 0.6 | 38.4 | 2.38 | 3.5 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Standard Z-sharp | 140 |  | 128 | 0.6 | 76.8 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 70 |  | 6 | 0.6 | 3.6 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 70 |  | 1 | 5 | 5 | 0.65 | 1.15 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 70 |  | 1 | 10 | 10 | 0.6 | 1.05 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 70 |  | 3 | 3 | 9 | 0.83 | 1.45 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 70 |  | 3 | 5 | 15 | 0.79 | 1.4 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow Z-sharp | 70 |  | 32 | 0.6 | 19.2 | 0.84 | 1.48 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 70 |  | 64 | 0.6 | 38.4 | 0.74 | 1.31 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow Z-sharp | 70 |  | 128 | 0.6 | 76.8 | 0.73 | 1.3 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 80 |  | 6 | 0.6 | 3.6 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 80 |  | 1 | 5 | 5 | 1.12 | 1.83 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 80 |  | 1 | 10 | 10 | 1.03 | 1.67 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 80 |  | 3 | 3 | 9 | 1.42 | 2.31 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 80 |  | 3 | 5 | 15 | 1.36 | 2.22 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow Z-sharp | 80 |  | 32 | 0.6 | 19.2 | 1.44 | 2.35 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 80 |  | 64 | 0.6 | 38.4 | 1.27 | 2.09 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow Z-sharp | 80 |  | 128 | 0.6 | 76.8 | 1.26 | 2.07 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 90 |  | 6 | 0.6 | 3.6 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 90 |  | 1 | 5 | 5 | 1.72 | 2.66 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 90 |  | 1 | 10 | 10 | 1.57 | 2.44 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 90 |  | 3 | 3 | 9 | 2.18 | 3.37 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 90 |  | 3 | 5 | 15 | 2.08 | 3.24 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow Z-sharp | 90 |  | 32 | 0.6 | 19.2 | 2.21 | 3.43 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 90 |  | 64 | 0.6 | 38.4 | 1.94 | 3.04 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow Z-sharp | 90 |  | 128 | 0.6 | 76.8 | 1.92 | 3.01 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 100 |  | 6 | 0.6 | 3.6 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 100 |  | 1 | 5 | 5 | 2.42 | 3.63 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 100 |  | 1 | 10 | 10 | 2.21 | 3.32 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 100 |  | 3 | 3 | 9 | 3.08 | 4.59 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 100 |  | 3 | 5 | 15 | 2.93 | 4.42 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow Z-sharp | 100 |  | 32 | 0.6 | 19.2 | 3.12 | 4.68 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 100 |  | 64 | 0.6 | 38.4 | 2.74 | 4.15 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow Z-sharp | 100 |  | 128 | 0.6 | 76.8 | 2.71 | 4.1 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 110 |  | 6 | 0.6 | 3.6 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 110 |  | 1 | 5 | 5 | 3.21 | 4.71 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 110 |  | 1 | 10 | 10 | 2.94 | 4.31 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 110 |  | 3 | 3 | 9 | 4.09 | 5.96 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 110 |  | 3 | 5 | 15 | 3.89 | 5.74 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow Z-sharp | 110 |  | 32 | 0.6 | 19.2 | 4.14 | 6.08 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 110 |  | 64 | 0.6 | 38.4 | 3.65 | 5.39 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow Z-sharp | 110 |  | 128 | 0.6 | 76.8 | 3.61 | 5.33 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 120 |  | 6 | 0.6 | 3.6 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 120 |  | 1 | 5 | 5 | 4.11 | 5.92 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 120 |  | 1 | 10 | 10 | 3.76 | 5.41 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 120 |  | 3 | 3 | 9 | 5.22 | 7.49 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 120 |  | 3 | 5 | 15 | 4.97 | 7.2 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow Z-sharp | 120 |  | 32 | 0.6 | 19.2 | 5.29 | 7.63 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 120 |  | 64 | 0.6 | 38.4 | 4.65 | 6.76 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow Z-sharp | 120 |  | 128 | 0.6 | 76.8 | 4.61 | 6.69 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 130 |  | 6 | 0.6 | 3.6 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 130 |  | 1 | 5 | 5 | 5.08 | 7.22 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 130 |  | 1 | 10 | 10 | 4.65 | 6.61 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 130 |  | 3 | 3 | 9 | 6.47 | 9.16 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 130 |  | 3 | 5 | 15 | 6.14 | 8.78 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow Z-sharp | 130 |  | 32 | 0.6 | 19.2 | 6.55 | 9.31 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 130 |  | 64 | 0.6 | 38.4 | 5.76 | 8.25 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow Z-sharp | 130 |  | 128 | 0.6 | 76.8 | 5.71 | 8.17 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 140 |  | 6 | 0.6 | 3.6 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 140 |  | 1 | 5 | 5 | 6.14 | 8.63 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 140 |  | 1 | 10 | 10 | 5.62 | 7.91 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 140 |  | 3 | 3 | 9 | 7.81 | 10.94 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 140 |  | 3 | 5 | 15 | 7.43 | 10.5 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow Z-sharp | 140 |  | 32 | 0.6 | 19.2 | 7.92 | 11.13 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 140 |  | 64 | 0.6 | 38.4 | 6.96 | 9.86 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow Z-sharp | 140 |  | 128 | 0.6 | 76.8 | 6.9 | 9.76 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 150 |  | 6 | 0.6 | 3.6 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 150 |  | 1 | 5 | 5 | 7.29 | 10.15 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 150 |  | 1 | 10 | 10 | 6.66 | 9.29 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 150 |  | 3 | 3 | 9 | 9.26 | 12.85 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 150 |  | 3 | 5 | 15 | 8.8 | 12.32 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow Z-sharp | 150 |  | 32 | 0.6 | 19.2 | 9.38 | 13.08 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 150 |  | 64 | 0.6 | 38.4 | 8.26 | 11.6 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow Z-sharp | 150 |  | 128 | 0.6 | 76.8 | 8.18 | 11.47 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 100 | SnLD | 6 | 0.6 | 3.6 | 0.83 | 1.1 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 100 |  | 1 | 5 | 5 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 100 |  | 1 | 10 | 10 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 100 |  | 3 | 3 | 9 | 0.63 | 0.84 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 100 |  | 3 | 5 | 15 | 0.6 | 0.81 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow Z-sharp | 100 |  | 32 | 0.6 | 19.2 | 0.64 | 0.85 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 100 |  | 64 | 0.6 | 38.4 | 0.56 | 0.76 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow Z-sharp | 100 |  | 128 | 0.6 | 76.8 | 0.56 | 0.75 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 110 |  | 6 | 0.6 | 3.6 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 110 |  | 1 | 5 | 5 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 110 |  | 1 | 10 | 10 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 110 |  | 3 | 3 | 9 | 1 | 1.31 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 110 |  | 3 | 5 | 15 | 0.95 | 1.26 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow Z-sharp | 110 |  | 32 | 0.6 | 19.2 | 1.01 | 1.33 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 110 |  | 64 | 0.6 | 38.4 | 0.89 | 1.18 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow Z-sharp | 110 |  | 128 | 0.6 | 76.8 | 0.88 | 1.17 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 120 |  | 6 | 0.6 | 3.6 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 120 |  | 1 | 5 | 5 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 120 |  | 1 | 10 | 10 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 120 |  | 3 | 3 | 9 | 1.46 | 1.91 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 120 |  | 3 | 5 | 15 | 1.39 | 1.84 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow Z-sharp | 120 |  | 32 | 0.6 | 19.2 | 1.48 | 1.94 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 120 |  | 64 | 0.6 | 38.4 | 1.3 | 1.75 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow Z-sharp | 120 |  | 128 | 0.6 | 76.8 | 1.29 | 1.71 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 130 |  | 6 | 0.6 | 3.6 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 130 |  | 1 | 5 | 5 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 130 |  | 1 | 10 | 10 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 130 |  | 3 | 3 | 9 | 2.02 | 2.63 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 130 |  | 3 | 5 | 15 | 1.93 | 2.53 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow Z-sharp | 130 |  | 32 | 0.6 | 19.2 | 2.05 | 2.68 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 130 |  | 64 | 0.6 | 38.4 | 1.8 | 2.38 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow Z-sharp | 130 |  | 128 | 0.6 | 76.8 | 1.79 | 2.36 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 140 |  | 6 | 0.6 | 3.6 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 140 |  | 1 | 5 | 5 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 140 |  | 1 | 10 | 10 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 140 |  | 3 | 3 | 9 | 2.69 | 3.5 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 140 |  | 3 | 5 | 15 | 2.56 | 3.36 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow Z-sharp | 140 |  | 32 | 0.6 | 19.2 | 2.73 | 3.56 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 140 |  | 64 | 0.6 | 38.4 | 2.4 | 3.16 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow Z-sharp | 140 |  | 128 | 0.6 | 76.8 | 2.38 | 3.13 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 150 |  | 6 | 0.6 | 3.6 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 150 |  | 1 | 5 | 5 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 150 |  | 1 | 10 | 10 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 150 |  | 3 | 3 | 9 | 3.46 | 4.49 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 150 |  | 3 | 5 | 15 | 3.3 | 4.31 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow Z-sharp | 150 |  | 32 | 0.6 | 19.2 | 3.51 | 4.57 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 150 |  | 64 | 0.6 | 38.4 | 3.08 | 4.06 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow Z-sharp | 150 |  | 128 | 0.6 | 76.8 | 3.06 | 4.02 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 150 | SnDE | 6 | 0.6 | 3.6 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 150 |  | 1 | 5 | 5 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 150 |  | 1 | 10 | 10 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 150 |  | 3 | 3 | 9 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 150 |  | 3 | 5 | 15 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow Z-sharp | 150 |  | 32 | 0.6 | 19.2 | 2.14 | 2.78 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 150 |  | 64 | 0.6 | 38.4 | 1.89 | 2.47 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow Z-sharp | 150 |  | 128 | 0.6 | 76.8 | 1.88 | 2.45 | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 120 | AuSn | 6 | 0.6 | 3.6 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 120 |  | 1 | 5 | 5 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 120 |  | 1 | 10 | 10 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 120 |  | 3 | 3 | 9 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 120 |  | 3 | 5 | 15 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow Z-sharp | 120 |  | 32 | 0.6 | 19.2 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 120 |  | 64 | 0.6 | 38.4 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow Z-sharp | 120 |  | 128 | 0.6 | 76.8 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 140 | AuSn | 6 | 0.6 | 3.6 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 140 |  | 1 | 5 | 5 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 140 |  | 1 | 10 | 10 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 140 |  | 3 | 3 | 9 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 140 |  | 3 | 5 | 15 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow Z-sharp | 140 |  | 32 | 0.6 | 19.2 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow | 140 |  | 64 | 0.6 | 38.4 |  |  | X.cite | Siemens
X.cite | 10 | 1 | 32 | Narrow Z-sharp | 140 |  | 128 | 0.6 | 76.8 |  |  | X.cite | Siemens
go.Top | 7 | 1 | 16 | Standard | 70 | Sn | 6 | 0.6 | 3.6 |  |  | go.Top | Siemens
go.Top | 7 | 1 | 16 | Narrow | 70 | Sn | 6 | 0.6 | 3.6 |  |  | go.Top | Siemens
go.Top | 7 | 1 | 16 | Standard | 70 |  | 1 | 5 | 5 | 3.7 | 4 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Narrow | 70 |  | 1 | 5 | 5 | 3.3 | 3.3 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Standard | 70 |  | 1 | 10 | 10 | 3.7 | 3.9 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Narrow | 70 |  | 1 | 10 | 10 | 3.3 | 3.3 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Standard | 70 |  | 3 | 5 | 15 | 3.7 | 4 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Standard | 70 | Sn | 3 | 5 | 15 |  |  | go.Top | Siemens
go.Top | 7 | 1 | 16 | Narrow | 70 |  | 3 | 5 | 15 | 3.3 | 3.3 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Narrow | 70 | Sn | 3 | 5 | 5 |  |  | go.Top | Siemens
go.Top | 7 | 1 | 16 | Standard | 70 |  | 64 | 0.6 | 38.4 | 3.8 | 4.1 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Standard | 70 | Sn | 64 | 0.6 | 38.4 |  |  | go.Top | Siemens
go.Top | 7 | 1 | 16 | Narrow | 70 |  | 64 | 0.6 | 38.4 | 3.4 | 3.4 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Narrow | 70 | Sn | 64 | 0.6 | 38.4 |  |  | go.Top | Siemens
go.Top | 7 | 1 | 16 | Standard | 70 | AuSn | 64 | 0.6 | 38.4 |  |  | go.Top | Siemens
go.Top | 7 | 1 | 16 | Standard | 80 | Sn | 6 | 0.6 | 3.6 |  |  | go.Top | Siemens
go.Top | 7 | 1 | 16 | Narrow | 80 | Sn | 6 | 0.6 | 3.6 |  |  | go.Top | Siemens
go.Top | 7 | 1 | 16 | Standard | 80 |  | 1 | 5 | 5 | 5.7 | 6 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Narrow | 80 |  | 1 | 5 | 5 | 5.1 | 5.1 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Standard | 80 |  | 1 | 10 | 10 | 5.7 | 6 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Narrow | 80 |  | 1 | 10 | 10 | 5.1 | 5.1 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Standard | 80 |  | 3 | 5 | 15 | 5.7 | 6 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Standard | 80 | Sn | 3 | 5 | 15 |  |  | go.Top | Siemens
go.Top | 7 | 1 | 16 | Narrow | 80 |  | 3 | 5 | 15 | 5.1 | 5.1 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Narrow | 80 | Sn | 3 | 5 | 5 |  |  | go.Top | Siemens
go.Top | 7 | 1 | 16 | Standard | 80 |  | 64 | 0.6 | 38.4 | 5.9 | 6.2 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Standard | 80 | Sn | 64 | 0.6 | 38.4 |  |  | go.Top | Siemens
go.Top | 7 | 1 | 16 | Narrow | 80 |  | 64 | 0.6 | 38.4 | 5.3 | 5.3 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Narrow | 80 | Sn | 64 | 0.6 | 38.4 |  |  | go.Top | Siemens
go.Top | 7 | 1 | 16 | Standard | 80 | AuSn | 64 | 0.6 | 38.4 |  |  | go.Top | Siemens
go.Top | 7 | 1 | 16 | Standard | 90 | Sn | 6 | 0.6 | 3.6 |  |  | go.Top | Siemens
go.Top | 7 | 1 | 16 | Narrow | 90 | Sn | 6 | 0.6 | 3.6 |  |  | go.Top | Siemens
go.Top | 7 | 1 | 16 | Standard | 90 |  | 1 | 5 | 5 | 8.2 | 8.5 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Narrow | 90 |  | 1 | 5 | 5 | 7.4 | 7.3 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Standard | 90 |  | 1 | 10 | 10 | 8.1 | 8.4 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Narrow | 90 |  | 1 | 10 | 10 | 7.4 | 7.2 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Standard | 90 |  | 3 | 5 | 15 | 8.2 | 8.5 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Standard | 90 | Sn | 3 | 5 | 15 |  |  | go.Top | Siemens
go.Top | 7 | 1 | 16 | Narrow | 90 |  | 3 | 5 | 15 | 7.4 | 7.3 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Narrow | 90 | Sn | 3 | 5 | 5 |  |  | go.Top | Siemens
go.Top | 7 | 1 | 16 | Standard | 90 |  | 64 | 0.6 | 38.4 | 8.4 | 8.8 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Standard | 90 | Sn | 64 | 0.6 | 38.4 |  |  | go.Top | Siemens
go.Top | 7 | 1 | 16 | Narrow | 90 |  | 64 | 0.6 | 38.4 | 7.7 | 7.6 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Narrow | 90 | Sn | 64 | 0.6 | 38.4 |  |  | go.Top | Siemens
go.Top | 7 | 1 | 16 | Standard | 90 | AuSn | 64 | 0.6 | 38.4 |  |  | go.Top | Siemens
go.Top | 7 | 1 | 16 | Standard | 100 | Sn | 6 | 0.6 | 3.6 | 1.3 | 1.3 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Narrow | 100 | Sn | 6 | 0.6 | 3.6 | 1.2 | 1.2 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Standard | 100 |  | 1 | 5 | 5 | 11.1 | 11.5 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Narrow | 100 |  | 1 | 5 | 5 | 10.1 | 9.9 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Standard | 100 |  | 1 | 10 | 10 | 11 | 11.3 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Narrow | 100 |  | 1 | 10 | 10 | 10.1 | 9.8 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Standard | 100 |  | 3 | 5 | 15 | 11 | 11.4 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Standard | 100 | Sn | 3 | 5 | 15 | 0.9 | 0.9 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Narrow | 100 |  | 3 | 5 | 15 | 10.1 | 9.9 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Narrow | 100 | Sn | 3 | 5 | 5 | 0.9 | 0.8 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Standard | 100 |  | 64 | 0.6 | 38.4 | 11.4 | 11.8 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Standard | 100 | Sn | 64 | 0.6 | 38.4 | 0.9 | 0.9 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Narrow | 100 |  | 64 | 0.6 | 38.4 | 10.4 | 10.3 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Narrow | 100 | Sn | 64 | 0.6 | 38.4 | 0.9 | 0.9 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Standard | 100 | AuSn | 64 | 0.6 | 38.4 |  |  | go.Top | Siemens
go.Top | 7 | 1 | 16 | Standard | 110 | Sn | 6 | 0.6 | 3.6 |  |  | go.Top | Siemens
go.Top | 7 | 1 | 16 | Narrow | 110 | Sn | 6 | 0.6 | 3.6 |  |  | go.Top | Siemens
go.Top | 7 | 1 | 16 | Standard | 110 |  | 1 | 5 | 5 | 14.2 | 14.7 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Narrow | 110 |  | 1 | 5 | 5 | 13.1 | 12.8 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Standard | 110 |  | 1 | 10 | 10 | 14.1 | 14.5 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Narrow | 110 |  | 1 | 10 | 10 | 13 | 12.7 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Standard | 110 |  | 3 | 5 | 15 | 14.2 | 14.6 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Standard | 110 | Sn | 3 | 5 | 15 | 1.5 | 1.5 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Narrow | 110 |  | 3 | 5 | 15 | 13 | 12.8 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Narrow | 110 | Sn | 3 | 5 | 5 | 1.4 | 1.4 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Standard | 110 |  | 64 | 0.6 | 38.4 | 14.6 | 15.1 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Standard | 110 | Sn | 64 | 0.6 | 38.4 | 1.6 | 1.6 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Narrow | 110 |  | 64 | 0.6 | 38.4 | 13.5 | 13.2 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Narrow | 110 | Sn | 64 | 0.6 | 38.4 | 1.5 | 1.5 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Standard | 110 | AuSn | 64 | 0.6 | 38.4 |  |  | go.Top | Siemens
go.Top | 7 | 1 | 16 | Standard | 120 | Sn | 6 | 0.6 | 3.6 |  |  | go.Top | Siemens
go.Top | 7 | 1 | 16 | Narrow | 120 | Sn | 6 | 0.6 | 3.6 |  |  | go.Top | Siemens
go.Top | 7 | 1 | 16 | Standard | 120 |  | 1 | 5 | 5 | 17.8 | 18.2 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Narrow | 120 |  | 1 | 5 | 5 | 16.3 | 16 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Standard | 120 |  | 1 | 10 | 10 | 17.6 | 18 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Narrow | 120 |  | 1 | 10 | 10 | 16.2 | 15.6 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Standard | 120 |  | 3 | 5 | 15 | 17.7 | 18.1 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Standard | 120 | Sn | 3 | 5 | 15 | 2.3 | 2.3 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Narrow | 120 |  | 3 | 5 | 15 | 16.3 | 16 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Narrow | 120 | Sn | 3 | 5 | 5 | 2.2 | 2.2 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Standard | 120 |  | 64 | 0.6 | 38.4 | 18.3 | 18.8 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Standard | 120 | Sn | 64 | 0.6 | 38.4 | 2.4 | 2.4 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Narrow | 120 |  | 64 | 0.6 | 38.4 | 16.9 | 16.5 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Narrow | 120 | Sn | 64 | 0.6 | 38.4 | 2.3 | 2.2 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Standard | 120 | AuSn | 64 | 0.6 | 38.4 | 5.4 | 5.5 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Standard | 130 | Sn | 6 | 0.6 | 3.6 |  |  | go.Top | Siemens
go.Top | 7 | 1 | 16 | Narrow | 130 | Sn | 6 | 0.6 | 3.6 |  |  | go.Top | Siemens
go.Top | 7 | 1 | 16 | Standard | 130 |  | 1 | 5 | 5 | 21.3 | 21.9 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Narrow | 130 |  | 1 | 5 | 5 | 19.6 | 19.3 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Standard | 130 |  | 1 | 10 | 10 | 21 | 21.5 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Narrow | 130 |  | 1 | 10 | 10 | 19.5 | 19.1 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Standard | 130 |  | 3 | 5 | 15 | 21.2 | 21.7 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Standard | 130 | Sn | 3 | 5 | 15 | 3.3 | 3.4 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Narrow | 130 |  | 3 | 5 | 15 | 19.6 | 19.2 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Narrow | 130 | Sn | 3 | 5 | 5 | 3.2 | 3.1 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Standard | 130 |  | 64 | 0.6 | 38.4 | 21.9 | 22.4 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Standard | 130 | Sn | 64 | 0.6 | 38.4 | 3.5 | 3.5 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Narrow | 130 |  | 64 | 0.6 | 38.4 | 20.2 | 19.9 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Narrow | 130 | Sn | 64 | 0.6 | 38.4 | 3.3 | 3.2 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Standard | 130 | AuSn | 64 | 0.6 | 38.4 |  |  | go.Top | Siemens
go.Top | 7 | 1 | 16 | Standard | 140 | Sn | 6 | 0.6 | 3.6 | 6.4 | 6.5 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Narrow | 140 | Sn | 6 | 0.6 | 3.6 | 6.1 | 6 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Standard | 140 |  | 1 | 5 | 5 | 25 | 25.6 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Narrow | 140 |  | 1 | 5 | 5 | 23.2 | 22.8 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Standard | 140 |  | 1 | 10 | 10 | 24.7 | 25.3 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Narrow | 140 |  | 1 | 10 | 10 | 23 | 22.5 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Standard | 140 |  | 3 | 5 | 15 | 24.9 | 25.4 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Standard | 140 | Sn | 3 | 5 | 15 | 4.6 | 4.6 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Narrow | 140 |  | 3 | 5 | 15 | 23.2 | 22.6 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Narrow | 140 | Sn | 3 | 5 | 5 | 4.4 | 4.3 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Standard | 140 |  | 64 | 0.6 | 38.4 | 25.8 | 26.3 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Standard | 140 | Sn | 64 | 0.6 | 38.4 | 4.7 | 4.8 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Narrow | 140 |  | 64 | 0.6 | 38.4 | 23.9 | 23.4 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Narrow | 140 | Sn | 64 | 0.6 | 38.4 | 4.5 | 4.4 | go.Top | Siemens
go.Top | 7 | 1 | 16 | Standard | 140 | AuSn | 64 | 0.6 | 38.4 | 8.7 | 8.7 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Standard | 70 | Sn | 6 | 0.6 | 3.6 |  |  | go.Top | Siemens
go.Top | 7 | 1 | 32 | Narrow | 70 | Sn | 6 | 0.6 | 3.6 |  |  | go.Top | Siemens
go.Top | 7 | 1 | 32 | Standard | 70 |  | 1 | 5 | 5 | 0.9 | 1.7 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Narrow | 70 |  | 1 | 5 | 5 | 0.7 | 1.3 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Standard | 70 |  | 1 | 10 | 10 | 0.8 | 1.7 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Narrow | 70 |  | 1 | 10 | 10 | 3.3 | 3.3 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Standard | 70 |  | 3 | 5 | 15 | 0.8 | 1.7 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Standard | 70 | Sn | 3 | 5 | 15 |  |  | go.Top | Siemens
go.Top | 7 | 1 | 32 | Narrow | 70 |  | 3 | 5 | 15 | 0.7 | 1.3 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Narrow | 70 | Sn | 3 | 5 | 5 |  |  | go.Top | Siemens
go.Top | 7 | 1 | 32 | Standard | 70 |  | 64 | 0.6 | 38.4 | 0.9 | 1.7 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Standard | 70 | Sn | 64 | 0.6 | 38.4 |  |  | go.Top | Siemens
go.Top | 7 | 1 | 32 | Narrow | 70 |  | 64 | 0.6 | 38.4 | 0.8 | 1.3 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Narrow | 70 | Sn | 64 | 0.6 | 38.4 |  |  | go.Top | Siemens
go.Top | 7 | 1 | 32 | Standard | 70 | AuSn | 64 | 0.6 | 38.4 |  |  | go.Top | Siemens
go.Top | 7 | 1 | 32 | Standard | 80 | Sn | 6 | 0.6 | 3.6 |  |  | go.Top | Siemens
go.Top | 7 | 1 | 32 | Narrow | 80 | Sn | 6 | 0.6 | 3.6 |  |  | go.Top | Siemens
go.Top | 7 | 1 | 32 | Standard | 80 |  | 1 | 5 | 5 | 1.4 | 2.7 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Narrow | 80 |  | 1 | 5 | 5 | 1.2 | 2 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Standard | 80 |  | 1 | 10 | 10 | 1.4 | 2.6 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Narrow | 80 |  | 1 | 10 | 10 | 5.1 | 5.1 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Standard | 80 |  | 3 | 5 | 15 | 1.4 | 2.6 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Standard | 80 | Sn | 3 | 5 | 15 |  |  | go.Top | Siemens
go.Top | 7 | 1 | 32 | Narrow | 80 |  | 3 | 5 | 15 | 1.2 | 2 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Narrow | 80 | Sn | 3 | 5 | 5 |  |  | go.Top | Siemens
go.Top | 7 | 1 | 32 | Standard | 80 |  | 64 | 0.6 | 38.4 | 1.5 | 2.7 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Standard | 80 | Sn | 64 | 0.6 | 38.4 |  |  | go.Top | Siemens
go.Top | 7 | 1 | 32 | Narrow | 80 |  | 64 | 0.6 | 38.4 | 1.3 | 2.1 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Narrow | 80 | Sn | 64 | 0.6 | 38.4 |  |  | go.Top | Siemens
go.Top | 7 | 1 | 32 | Standard | 80 | AuSn | 64 | 0.6 | 38.4 |  |  | go.Top | Siemens
go.Top | 7 | 1 | 32 | Standard | 90 | Sn | 6 | 0.6 | 3.6 |  |  | go.Top | Siemens
go.Top | 7 | 1 | 32 | Narrow | 90 | Sn | 6 | 0.6 | 3.6 |  |  | go.Top | Siemens
go.Top | 7 | 1 | 32 | Standard | 90 |  | 1 | 5 | 5 | 2.2 | 3.9 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Narrow | 90 |  | 1 | 5 | 5 | 1.9 | 3 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Standard | 90 |  | 1 | 10 | 10 | 2.2 | 3.8 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Narrow | 90 |  | 1 | 10 | 10 | 7.4 | 7.2 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Standard | 90 |  | 3 | 5 | 15 | 2.2 | 3.8 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Standard | 90 | Sn | 3 | 5 | 15 |  |  | go.Top | Siemens
go.Top | 7 | 1 | 32 | Narrow | 90 |  | 3 | 5 | 15 | 1.9 | 3 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Narrow | 90 | Sn | 3 | 5 | 5 |  |  | go.Top | Siemens
go.Top | 7 | 1 | 32 | Standard | 90 |  | 64 | 0.6 | 38.4 | 2.3 | 4 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Standard | 90 | Sn | 64 | 0.6 | 38.4 |  |  | go.Top | Siemens
go.Top | 7 | 1 | 32 | Narrow | 90 |  | 64 | 0.6 | 38.4 | 2 | 3.1 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Narrow | 90 | Sn | 64 | 0.6 | 38.4 |  |  | go.Top | Siemens
go.Top | 7 | 1 | 32 | Standard | 90 | AuSn | 64 | 0.6 | 38.4 |  |  | go.Top | Siemens
go.Top | 7 | 1 | 32 | Standard | 100 | Sn | 6 | 0.6 | 3.6 | 0.5 | 0.7 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Narrow | 100 | Sn | 6 | 0.6 | 3.6 | 0.4 | 0.6 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Standard | 100 |  | 1 | 5 | 5 | 3.2 | 5.4 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Narrow | 100 |  | 1 | 5 | 5 | 2.8 | 4.2 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Standard | 100 |  | 1 | 10 | 10 | 3.1 | 5.2 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Narrow | 100 |  | 1 | 10 | 10 | 10.1 | 9.8 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Standard | 100 |  | 3 | 5 | 15 | 3.1 | 5.3 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Standard | 100 | Sn | 3 | 5 | 15 | 0.3 | 0.5 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Narrow | 100 |  | 3 | 5 | 15 | 2.8 | 4.1 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Narrow | 100 | Sn | 3 | 5 | 5 | 0.3 | 0.4 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Standard | 100 |  | 64 | 0.6 | 38.4 | 3.2 | 5.4 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Standard | 100 | Sn | 64 | 0.6 | 38.4 | 0.3 | 0.5 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Narrow | 100 |  | 64 | 0.6 | 38.4 | 2.9 | 4.3 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Narrow | 100 | Sn | 64 | 0.6 | 38.4 | 0.3 | 0.4 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Standard | 100 | AuSn | 64 | 0.6 | 38.4 |  |  | go.Top | Siemens
go.Top | 7 | 1 | 32 | Standard | 110 | Sn | 6 | 0.6 | 3.6 |  |  | go.Top | Siemens
go.Top | 7 | 1 | 32 | Narrow | 110 | Sn | 6 | 0.6 | 3.6 |  |  | go.Top | Siemens
go.Top | 7 | 1 | 32 | Standard | 110 |  | 1 | 5 | 5 | 4.2 | 7 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Narrow | 110 |  | 1 | 5 | 5 | 3.7 | 5.5 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Standard | 110 |  | 1 | 10 | 10 | 4.2 | 6.9 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Narrow | 110 |  | 1 | 10 | 10 | 13 | 12.7 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Standard | 110 |  | 3 | 5 | 15 | 4.2 | 6.9 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Standard | 110 | Sn | 3 | 5 | 15 | 0.6 | 0.8 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Narrow | 110 |  | 3 | 5 | 15 | 3.7 | 5.4 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Narrow | 110 | Sn | 3 | 5 | 5 | 0.5 | 0.7 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Standard | 110 |  | 64 | 0.6 | 38.4 | 4.3 | 7.1 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Standard | 110 | Sn | 64 | 0.6 | 38.4 | 0.6 | 0.9 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Narrow | 110 |  | 64 | 0.6 | 38.4 | 3.8 | 5.6 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Narrow | 110 | Sn | 64 | 0.6 | 38.4 | 0.5 | 0.7 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Standard | 110 | AuSn | 64 | 0.6 | 38.4 |  |  | go.Top | Siemens
go.Top | 7 | 1 | 32 | Standard | 120 | Sn | 6 | 0.6 | 3.6 |  |  | go.Top | Siemens
go.Top | 7 | 1 | 32 | Narrow | 120 | Sn | 6 | 0.6 | 3.6 |  |  | go.Top | Siemens
go.Top | 7 | 1 | 32 | Standard | 120 |  | 1 | 5 | 5 | 5.4 | 8.9 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Narrow | 120 |  | 1 | 5 | 5 | 4.8 | 7 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Standard | 120 |  | 1 | 10 | 10 | 5.3 | 8.6 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Narrow | 120 |  | 1 | 10 | 10 | 16.2 | 15.6 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Standard | 120 |  | 3 | 5 | 15 | 5.4 | 8.7 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Standard | 120 | Sn | 3 | 5 | 15 | 0.9 | 1.3 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Narrow | 120 |  | 3 | 5 | 15 | 4.8 | 6.9 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Narrow | 120 | Sn | 3 | 5 | 5 | 0.8 | 1.1 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Standard | 120 |  | 64 | 0.6 | 38.4 | 5.6 | 9 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Standard | 120 | Sn | 64 | 0.6 | 38.4 | 0.9 | 1.3 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Narrow | 120 |  | 64 | 0.6 | 38.4 | 5 | 7.1 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Narrow | 120 | Sn | 64 | 0.6 | 38.4 | 0.8 | 1.1 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Standard | 120 | AuSn | 64 | 0.6 | 38.4 | 1.8 | 2.8 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Standard | 130 | Sn | 6 | 0.6 | 3.6 |  |  | go.Top | Siemens
go.Top | 7 | 1 | 32 | Narrow | 130 | Sn | 6 | 0.6 | 3.6 |  |  | go.Top | Siemens
go.Top | 7 | 1 | 32 | Standard | 130 |  | 1 | 5 | 5 | 6.7 | 10.8 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Narrow | 130 |  | 1 | 5 | 5 | 5.9 | 8.5 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Standard | 130 |  | 1 | 10 | 10 | 6.6 | 10.5 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Narrow | 130 |  | 1 | 10 | 10 | 19.5 | 19.1 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Standard | 130 |  | 3 | 5 | 15 | 6.6 | 10.5 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Standard | 130 | Sn | 3 | 5 | 15 | 1.3 | 1.9 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Narrow | 130 |  | 3 | 5 | 15 | 5.9 | 8.4 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Narrow | 130 | Sn | 3 | 5 | 5 | 1.2 | 1.5 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Standard | 130 |  | 64 | 0.6 | 38.4 | 6.8 | 10.9 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Standard | 130 | Sn | 64 | 0.6 | 38.4 | 1.3 | 1.9 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Narrow | 130 |  | 64 | 0.6 | 38.4 | 6.1 | 8.7 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Narrow | 130 | Sn | 64 | 0.6 | 38.4 | 1.2 | 1.6 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Standard | 130 | AuSn | 64 | 0.6 | 38.4 |  |  | go.Top | Siemens
go.Top | 7 | 1 | 32 | Standard | 140 | Sn | 6 | 0.6 | 3.6 | 2.5 | 3.7 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Narrow | 140 | Sn | 6 | 0.6 | 3.6 | 2.3 | 3 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Standard | 140 |  | 1 | 5 | 5 | 8.1 | 12.9 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Narrow | 140 |  | 1 | 5 | 5 | 7.2 | 10.2 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Standard | 140 |  | 1 | 10 | 10 | 7.9 | 12.5 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Narrow | 140 |  | 1 | 10 | 10 | 23 | 22.5 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Standard | 140 |  | 3 | 5 | 15 | 8 | 12.5 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Standard | 140 | Sn | 3 | 5 | 15 | 1.8 | 2.6 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Narrow | 140 |  | 3 | 5 | 15 | 7.1 | 10 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Narrow | 140 | Sn | 3 | 5 | 5 | 1.6 | 2.1 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Standard | 140 |  | 64 | 0.6 | 38.4 | 8.3 | 13 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Standard | 140 | Sn | 64 | 0.6 | 38.4 | 1.9 | 2.7 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Narrow | 140 |  | 64 | 0.6 | 38.4 | 7.4 | 10.4 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Narrow | 140 | Sn | 64 | 0.6 | 38.4 | 1.7 | 2.2 | go.Top | Siemens
go.Top | 7 | 1 | 32 | Standard | 140 | AuSn | 64 | 0.6 | 38.4 | 3.1 | 4.6 | go.Top | Siemens
Definition Edge | 6 | 1 | 16 | Standard | 70 |  | 2 | 1 | 2 | 2.6 | 2.74 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Standard | 70 |  | 6 | 0.6 | 3.6 | 4.22 | 4.43 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Standard | 70 |  | 8 | 0.6 | 4.8 | 3.88 | 4.07 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Standard Z-sharp | 70 |  | 8 | 0.6 | 4.8 | 4.58 | 4.78 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Standard | 70 |  | 1 | 5 | 5 | 2.34 | 2.46 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Standard | 70 |  | 10 | 0.6 | 6 | 3.17 | 3.31 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Standard | 70 |  | 6 | 1.2 | 7.2 | 3.36 | 3.53 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Standard | 70 |  | 1 | 10 | 10 | 2.34 | 2.46 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Standard | 70 |  | 20 | 0.6 | 12 | 3.23 | 3.38 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Standard | 70 |  | 12 | 1.2 | 14.4 | 3 | 3.15 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Standard | 70 |  | 32 | 1.2 | 38.4 | 2.62 | 2.75 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Standard | 70 |  | 64 | 0.6 | 38.4 | 2.67 | 2.79 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Standard | 80 |  | 2 | 1 | 2 | 4.18 | 4.36 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Standard | 80 |  | 6 | 0.6 | 3.6 | 6.76 | 7.06 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Standard | 80 |  | 8 | 0.6 | 4.8 | 6.22 | 6.49 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Standard Z-sharp | 80 |  | 8 | 0.6 | 4.8 | 7.24 | 7.46 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Standard | 80 |  | 1 | 5 | 5 | 3.75 | 3.91 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Standard | 80 |  | 10 | 0.6 | 6 | 5.02 | 5.17 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Standard | 80 |  | 6 | 1.2 | 7.2 | 5.39 | 5.62 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Standard | 80 |  | 1 | 10 | 10 | 3.75 | 3.91 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Standard | 80 |  | 20 | 0.6 | 12 | 5.12 | 5.27 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Standard | 80 |  | 12 | 1.2 | 14.4 | 4.81 | 5.02 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Standard | 80 |  | 32 | 1.2 | 38.4 | 4.2 | 4.38 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Standard | 80 |  | 64 | 0.6 | 38.4 | 4.23 | 4.35 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Standard | 100 |  | 2 | 1 | 2 | 8.44 | 8.65 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Standard | 100 |  | 6 | 0.6 | 3.6 | 13.67 | 14.01 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Standard | 100 |  | 8 | 0.6 | 4.8 | 12.56 | 12.88 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Standard Z-sharp | 100 |  | 8 | 0.6 | 4.8 | 14.69 | 14.93 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Standard | 100 |  | 1 | 5 | 5 | 7.57 | 7.76 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Standard | 100 |  | 10 | 0.6 | 6 | 10.19 | 10.35 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Standard | 100 |  | 6 | 1.2 | 7.2 | 10.89 | 11.16 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Standard | 100 |  | 1 | 10 | 10 | 7.57 | 7.76 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Standard | 100 |  | 20 | 0.6 | 12 | 10.38 | 10.55 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Standard | 100 |  | 12 | 1.2 | 14.4 | 9.71 | 9.96 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Standard | 100 |  | 32 | 1.2 | 38.4 | 8.48 | 8.69 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Standard | 100 |  | 64 | 0.6 | 38.4 | 8.57 | 8.71 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Standard | 120 |  | 2 | 1 | 2 | 13.85 | 14.11 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Standard | 120 |  | 6 | 0.6 | 3.6 | 22.42 | 22.85 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Standard | 120 |  | 8 | 0.6 | 4.8 | 20.6 | 21 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Standard Z-sharp | 120 |  | 8 | 0.6 | 4.8 | 24.33 | 24.63 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Standard | 120 |  | 1 | 5 | 5 | 12.42 | 12.66 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Standard | 120 |  | 10 | 0.6 | 6 | 16.87 | 17.08 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Standard | 120 |  | 6 | 1.2 | 7.2 | 17.86 | 18.2 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Standard | 120 |  | 1 | 10 | 10 | 12.42 | 12.66 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Standard | 120 |  | 20 | 0.6 | 12 | 17.19 | 17.41 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Standard | 120 |  | 12 | 1.2 | 14.4 | 15.93 | 16.24 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Standard | 120 |  | 32 | 1.2 | 38.4 | 13.91 | 14.17 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Standard | 120 |  | 64 | 0.6 | 38.4 | 14.19 | 14.37 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Standard | 140 |  | 2 | 1 | 2 | 20.24 | 20.6 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Standard | 140 |  | 6 | 0.6 | 3.6 | 32.77 | 33.35 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Standard | 140 |  | 8 | 0.6 | 4.8 | 30.12 | 30.65 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Standard Z-sharp | 140 |  | 8 | 0.6 | 4.8 | 36.16 | 36.64 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Standard | 140 |  | 1 | 5 | 5 | 18.16 | 18.48 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Standard | 140 |  | 10 | 0.6 | 6 | 25.07 | 25.4 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Standard | 140 |  | 6 | 1.2 | 7.2 | 26.11 | 26.56 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Standard | 140 |  | 1 | 10 | 10 | 18.16 | 18.48 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Standard | 140 |  | 20 | 0.6 | 12 | 25.55 | 25.89 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Standard | 140 |  | 12 | 1.2 | 14.4 | 23.3 | 23.7 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Standard | 140 |  | 32 | 1.2 | 38.4 | 20.33 | 20.69 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Standard | 140 |  | 64 | 0.6 | 38.4 | 21.09 | 21.37 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Narrow | 70 |  | 2 | 1 | 2 | 0.61 | 1.13 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Narrow | 70 |  | 6 | 0.6 | 3.6 | 0.99 | 1.83 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Narrow | 70 |  | 8 | 0.6 | 4.8 | 0.91 | 1.68 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Narrow Z-sharp | 70 |  | 8 | 0.6 | 4.8 | 1.07 | 1.97 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Narrow | 70 |  | 1 | 5 | 5 | 0.55 | 1.01 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Narrow | 70 |  | 10 | 0.6 | 6 | 0.74 | 1.37 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Narrow | 70 |  | 6 | 1.2 | 7.2 | 0.79 | 1.46 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Narrow | 70 |  | 1 | 10 | 10 | 0.55 | 1.01 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Narrow | 70 |  | 20 | 0.6 | 12 | 0.75 | 1.4 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Narrow | 70 |  | 12 | 1.2 | 14.4 | 0.7 | 1.3 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Narrow | 70 |  | 32 | 1.2 | 38.4 | 0.61 | 1.14 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Narrow | 70 |  | 64 | 0.6 | 38.4 | 0.67 | 1.15 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Narrow | 80 |  | 2 | 1 | 2 | 1.07 | 1.86 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Narrow | 80 |  | 6 | 0.6 | 3.6 | 1.73 | 3.01 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Narrow | 80 |  | 8 | 0.6 | 4.8 | 1.59 | 2.77 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Narrow Z-sharp | 80 |  | 8 | 0.6 | 4.8 | 1.83 | 3.18 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Narrow | 80 |  | 1 | 5 | 5 | 0.96 | 1.67 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Narrow | 80 |  | 10 | 0.6 | 6 | 1.27 | 2.21 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Narrow | 80 |  | 6 | 1.2 | 7.2 | 1.38 | 2.4 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Narrow | 80 |  | 1 | 10 | 10 | 0.96 | 1.67 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Narrow | 80 |  | 20 | 0.6 | 12 | 1.29 | 2.25 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Narrow | 80 |  | 12 | 1.2 | 14.4 | 1.23 | 2.14 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Narrow | 80 |  | 32 | 1.2 | 38.4 | 1.08 | 1.87 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Narrow | 80 |  | 64 | 0.6 | 38.4 | 1.07 | 1.86 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Narrow | 100 |  | 2 | 1 | 2 | 2.43 | 3.9 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Narrow | 100 |  | 6 | 0.6 | 3.6 | 3.94 | 6.32 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Narrow | 100 |  | 8 | 0.6 | 4.8 | 3.62 | 5.81 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Narrow Z-sharp | 100 |  | 8 | 0.6 | 4.8 | 4.19 | 6.75 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Narrow | 100 |  | 1 | 5 | 5 | 2.18 | 3.5 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Narrow | 100 |  | 10 | 0.6 | 6 | 2.91 | 4.68 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Narrow | 100 |  | 6 | 1.2 | 7.2 | 3.14 | 5.03 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Narrow | 100 |  | 1 | 10 | 10 | 2.18 | 3.5 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Narrow | 100 |  | 20 | 0.6 | 12 | 2.96 | 4.77 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Narrow | 100 |  | 12 | 1.2 | 14.4 | 2.8 | 4.49 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Narrow | 100 |  | 32 | 1.2 | 38.4 | 2.44 | 3.92 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Narrow | 100 |  | 64 | 0.6 | 38.4 | 2.45 | 3.94 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Narrow | 120 |  | 2 | 1 | 2 | 4.28 | 6.61 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Narrow | 120 |  | 6 | 0.6 | 3.6 | 6.93 | 10.7 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Narrow | 120 |  | 8 | 0.6 | 4.8 | 6.37 | 9.84 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Narrow Z-sharp | 120 |  | 8 | 0.6 | 4.8 | 7.5 | 11.55 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Narrow | 120 |  | 1 | 5 | 5 | 3.84 | 5.93 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Narrow | 120 |  | 10 | 0.6 | 6 | 5.2 | 8.01 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Narrow | 120 |  | 6 | 1.2 | 7.2 | 5.52 | 8.53 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Narrow | 120 |  | 1 | 10 | 10 | 3.84 | 5.93 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Narrow | 120 |  | 20 | 0.6 | 12 | 5.3 | 8.16 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Narrow | 120 |  | 12 | 1.2 | 14.4 | 4.93 | 7.61 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Narrow | 120 |  | 32 | 1.2 | 38.4 | 4.3 | 6.64 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Narrow | 120 |  | 64 | 0.6 | 38.4 | 4.37 | 6.74 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Narrow | 140 |  | 2 | 1 | 2 | 6.57 | 9.94 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Narrow | 140 |  | 6 | 0.6 | 3.6 | 10.64 | 16.09 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Narrow | 140 |  | 8 | 0.6 | 4.8 | 9.78 | 14.79 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Narrow Z-sharp | 140 |  | 8 | 0.6 | 4.8 | 11.59 | 17.62 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Narrow | 140 |  | 1 | 5 | 5 | 5.89 | 8.91 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Narrow | 140 |  | 10 | 0.6 | 6 | 8.03 | 12.22 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Narrow | 140 |  | 6 | 1.2 | 7.2 | 8.47 | 12.81 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Narrow | 140 |  | 1 | 10 | 10 | 5.89 | 8.91 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Narrow | 140 |  | 20 | 0.6 | 12 | 8.19 | 12.45 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Narrow | 140 |  | 12 | 1.2 | 14.4 | 7.56 | 11.43 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Narrow | 140 |  | 32 | 1.2 | 38.4 | 6.6 | 9.98 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 16 | Narrow | 140 |  | 64 | 0.6 | 38.4 | 6.76 | 10.28 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Standard | 70 |  | 2 | 1 | 2 | 2.36 | 2.36 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Standard | 70 |  | 6 | 0.6 | 3.6 | 3.83 | 3.81 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Standard | 70 |  | 8 | 0.6 | 4.8 | 3.52 | 3.51 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Standard Z-sharp | 70 |  | 8 | 0.6 | 4.8 | 4.12 | 4.07 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Standard | 70 |  | 1 | 5 | 5 | 2.12 | 2.11 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Standard | 70 |  | 10 | 0.6 | 6 | 2.86 | 2.82 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Standard | 70 |  | 6 | 1.2 | 7.2 | 3.05 | 3.04 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Standard | 70 |  | 1 | 10 | 10 | 2.12 | 2.11 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Standard | 70 |  | 20 | 0.6 | 12 | 2.91 | 2.88 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Standard | 70 |  | 12 | 1.2 | 14.4 | 2.72 | 2.71 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Standard | 70 |  | 32 | 1.2 | 38.4 | 2.38 | 2.37 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Standard | 70 |  | 64 | 0.6 | 38.4 | 2.4 | 2.37 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Standard | 80 |  | 2 | 1 | 2 | 3.81 | 3.78 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Standard | 80 |  | 6 | 0.6 | 3.6 | 6.17 | 6.12 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Standard | 80 |  | 8 | 0.6 | 4.8 | 5.67 | 5.63 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Standard Z-sharp | 80 |  | 8 | 0.6 | 4.8 | 6.58 | 6.44 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Standard | 80 |  | 1 | 5 | 5 | 3.42 | 3.39 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Standard | 80 |  | 10 | 0.6 | 6 | 4.56 | 4.47 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Standard | 80 |  | 6 | 1.2 | 7.2 | 4.92 | 4.88 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Standard | 80 |  | 1 | 10 | 10 | 3.42 | 3.39 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Standard | 80 |  | 20 | 0.6 | 12 | 4.65 | 4.55 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Standard | 80 |  | 12 | 1.2 | 14.4 | 4.39 | 4.39 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Standard | 80 |  | 32 | 1.2 | 38.4 | 3.83 | 3.8 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Standard | 80 |  | 64 | 0.6 | 38.4 | 3.84 | 3.76 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Standard | 100 |  | 2 | 1 | 2 | 7.81 | 7.69 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Standard | 100 |  | 6 | 0.6 | 3.6 | 12.65 | 12.45 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Standard | 100 |  | 8 | 0.6 | 4.8 | 11.63 | 11.44 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Standard Z-sharp | 100 |  | 8 | 0.6 | 4.8 | 13.53 | 13.14 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Standard | 100 |  | 1 | 5 | 5 | 7.01 | 6.9 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Standard | 100 |  | 10 | 0.6 | 6 | 9.38 | 9.11 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Standard | 100 |  | 6 | 1.2 | 7.2 | 10.08 | 9.92 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Standard | 100 |  | 1 | 10 | 10 | 7.01 | 6.9 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Standard | 100 |  | 20 | 0.6 | 12 | 9.56 | 9.29 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Standard | 100 |  | 12 | 1.2 | 14.4 | 8.99 | 8.85 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Standard | 100 |  | 32 | 1.2 | 38.4 | 7.85 | 7.72 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Standard | 100 |  | 64 | 0.6 | 38.4 | 7.89 | 7.67 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Standard | 120 |  | 2 | 1 | 2 | 12.93 | 12.68 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Standard | 120 |  | 6 | 0.6 | 3.6 | 20.93 | 20.52 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Standard | 120 |  | 8 | 0.6 | 4.8 | 19.23 | 18.86 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Standard Z-sharp | 120 |  | 8 | 0.6 | 4.8 | 22.53 | 21.91 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Standard | 120 |  | 1 | 5 | 5 | 11.6 | 11.37 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Standard | 120 |  | 10 | 0.6 | 6 | 15.62 | 16.19 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Standard | 120 |  | 6 | 1.2 | 7.2 | 16.67 | 16.35 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Standard | 120 |  | 1 | 10 | 10 | 11.6 | 11.37 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Standard | 120 |  | 20 | 0.6 | 12 | 15.92 | 15.48 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Standard | 120 |  | 12 | 1.2 | 14.4 | 14.87 | 14.59 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Standard | 120 |  | 32 | 1.2 | 38.4 | 12.98 | 12.73 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Standard | 120 |  | 64 | 0.6 | 38.4 | 13.14 | 12.78 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Standard | 140 |  | 2 | 1 | 2 | 18.83 | 18.5 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Standard | 140 |  | 6 | 0.6 | 3.6 | 30.49 | 29.95 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Standard | 140 |  | 8 | 0.6 | 4.8 | 28.02 | 27.53 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Standard Z-sharp | 140 |  | 8 | 0.6 | 4.8 | 33.71 | 32.85 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Standard | 140 |  | 1 | 5 | 5 | 16.89 | 16.6 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Standard | 140 |  | 10 | 0.6 | 6 | 23.38 | 22.77 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Standard | 140 |  | 6 | 1.2 | 7.2 | 24.29 | 23.86 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Standard | 140 |  | 1 | 10 | 10 | 16.89 | 16.6 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Standard | 140 |  | 20 | 0.6 | 12 | 23.82 | 23.21 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Standard | 140 |  | 12 | 1.2 | 14.4 | 21.67 | 21.29 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Standard | 140 |  | 32 | 1.2 | 38.4 | 18.91 | 18.58 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Standard | 140 |  | 64 | 0.6 | 38.4 | 19.67 | 19.16 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Narrow | 70 |  | 2 | 1 | 2 | 0.54 | 0.9 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Narrow | 70 |  | 6 | 0.6 | 3.6 | 0.87 | 1.45 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Narrow | 70 |  | 8 | 0.6 | 4.8 | 0.8 | 1.33 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Narrow Z-sharp | 70 |  | 8 | 0.6 | 4.8 | 0.93 | 1.54 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Narrow | 70 |  | 1 | 5 | 5 | 0.48 | 0.8 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Narrow | 70 |  | 10 | 0.6 | 6 | 0.65 | 1.07 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Narrow | 70 |  | 6 | 1.2 | 7.2 | 0.69 | 1.16 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Narrow | 70 |  | 1 | 10 | 10 | 0.48 | 0.8 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Narrow | 70 |  | 20 | 0.6 | 12 | 0.66 | 1.09 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Narrow | 70 |  | 12 | 1.2 | 14.4 | 0.62 | 1.03 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Narrow | 70 |  | 32 | 1.2 | 38.4 | 0.54 | 0.9 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Narrow | 70 |  | 64 | 0.6 | 38.4 | 0.55 | 0.9 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Narrow | 80 |  | 2 | 1 | 2 | 0.94 | 1.46 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Narrow | 80 |  | 6 | 0.6 | 3.6 | 1.52 | 2.36 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Narrow | 80 |  | 8 | 0.6 | 4.8 | 1.39 | 2.17 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Narrow Z-sharp | 80 |  | 8 | 0.6 | 4.8 | 1.62 | 2.48 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Narrow | 80 |  | 1 | 5 | 5 | 0.84 | 1.31 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Narrow | 80 |  | 10 | 0.6 | 6 | 1.12 | 1.72 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Narrow | 80 |  | 6 | 1.2 | 7.2 | 1.21 | 1.88 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Narrow | 80 |  | 1 | 10 | 10 | 0.84 | 1.31 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Narrow | 80 |  | 20 | 0.6 | 12 | 1.15 | 1.76 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Narrow | 80 |  | 12 | 1.2 | 14.4 | 1.08 | 1.68 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Narrow | 80 |  | 32 | 1.2 | 38.4 | 0.94 | 1.47 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Narrow | 80 |  | 64 | 0.6 | 38.4 | 0.95 | 1.45 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Narrow | 100 |  | 2 | 1 | 2 | 2.17 | 3.12 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Narrow | 100 |  | 6 | 0.6 | 3.6 | 3.51 | 5.06 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Narrow | 100 |  | 8 | 0.6 | 4.8 | 3.22 | 4.65 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Narrow Z-sharp | 100 |  | 8 | 0.6 | 4.8 | 3.75 | 5.34 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Narrow | 100 |  | 1 | 5 | 5 | 1.94 | 2.8 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Narrow | 100 |  | 10 | 0.6 | 6 | 2.6 | 3.7 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Narrow | 100 |  | 6 | 1.2 | 7.2 | 2.79 | 4.03 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Narrow | 100 |  | 1 | 10 | 10 | 1.94 | 2.8 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Narrow | 100 |  | 20 | 0.6 | 12 | 2.65 | 3.78 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Narrow | 100 |  | 12 | 1.2 | 14.4 | 2.49 | 3.59 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Narrow | 100 |  | 32 | 1.2 | 38.4 | 2.17 | 3.14 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Narrow | 100 |  | 64 | 0.6 | 38.4 | 2.19 | 3.12 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Narrow | 120 |  | 2 | 1 | 2 | 3.84 | 5.36 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Narrow | 120 |  | 6 | 0.6 | 3.6 | 6.22 | 8.68 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Narrow | 120 |  | 8 | 0.6 | 4.8 | 5.72 | 7.98 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Narrow Z-sharp | 120 |  | 8 | 0.6 | 4.8 | 6.67 | 9.21 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Narrow | 120 |  | 1 | 5 | 5 | 3.45 | 4.81 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Narrow | 120 |  | 10 | 0.6 | 6 | 4.62 | 6.39 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Narrow | 120 |  | 6 | 1.2 | 7.2 | 4.95 | 6.91 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Narrow | 120 |  | 1 | 10 | 10 | 3.45 | 4.81 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Narrow | 120 |  | 20 | 0.6 | 12 | 4.71 | 6.51 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Narrow | 120 |  | 12 | 1.2 | 14.4 | 4.42 | 6.17 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Narrow | 120 |  | 32 | 1.2 | 38.4 | 3.86 | 5.39 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Narrow | 120 |  | 64 | 0.6 | 38.4 | 3.89 | 5.38 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Narrow | 140 |  | 2 | 1 | 2 | 5.89 | 8.03 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Narrow | 140 |  | 6 | 0.6 | 3.6 | 9.53 | 13 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Narrow | 140 |  | 8 | 0.6 | 4.8 | 8.76 | 11.95 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Narrow Z-sharp | 140 |  | 8 | 0.6 | 4.8 | 10.45 | 14.2 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Narrow | 140 |  | 1 | 5 | 5 | 5.28 | 7.2 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Narrow | 140 |  | 10 | 0.6 | 6 | 7.25 | 9.84 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Narrow | 140 |  | 6 | 1.2 | 7.2 | 7.59 | 10.36 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Narrow | 140 |  | 1 | 10 | 10 | 5.28 | 7.2 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Narrow | 140 |  | 20 | 0.6 | 12 | 7.38 | 10.03 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Narrow | 140 |  | 12 | 1.2 | 14.4 | 6.77 | 9.24 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Narrow | 140 |  | 32 | 1.2 | 38.4 | 5.91 | 8.07 | Definition Edge | Siemens
Definition Edge | 6 | 1 | 32 | Narrow | 140 |  | 64 | 0.6 | 38.4 | 6.1 | 8.28 | Definition Edge | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 70 |  | 2 | 1 | 2 | 2.57 | 2.69 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 70 |  | 6 | 0.6 | 3.6 | 4.17 | 4.35 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard Z-sharp | 70 |  | 8 | 0.6 | 4.8 | 4.39 | 4.56 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 70 |  | 1 | 5 | 5 | 2.31 | 2.41 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard Z-sharp | 70 |  | 10 | 0.6 | 6 | 3.04 | 3.16 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 70 |  | 1 | 10 | 10 | 2.31 | 2.41 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 70 |  | 12 | 1.2 | 14.4 | 2.96 | 3.09 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard Z-sharp | 70 |  | 20 | 0.6 | 12 | 3.1 | 3.22 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard Z-sharp | 70 |  | 32 | 0.6 | 19.2 | 2.86 | 2.98 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 70 |  | 12 | 1.2 | 14.4 | 2.59 | 2.7 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 70 |  | 64 | 0.6 | 38.4 | 2.52 | 2.62 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 80 |  | 2 | 1 | 2 | 4.12 | 4.24 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 80 |  | 6 | 0.6 | 3.6 | 6.66 | 6.87 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard Z-sharp | 80 |  | 8 | 0.6 | 4.8 | 7.03 | 7.21 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 80 |  | 1 | 5 | 5 | 3.69 | 3.8 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard Z-sharp | 80 |  | 10 | 0.6 | 6 | 4.87 | 5 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 80 |  | 1 | 10 | 10 | 3.69 | 3.8 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 80 |  | 12 | 1.2 | 14.4 | 4.74 | 4.88 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard Z-sharp | 80 |  | 20 | 0.6 | 12 | 4.97 | 5.1 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard Z-sharp | 80 |  | 32 | 0.6 | 19.2 | 4.58 | 4.71 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 80 |  | 12 | 1.2 | 14.4 | 4.13 | 4.26 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 80 |  | 64 | 0.6 | 38.4 | 4.04 | 4.15 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 90 |  | 2 | 1 | 2 | 6.06 | 6.2 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 90 |  | 6 | 0.6 | 3.6 | 9.81 | 10 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard Z-sharp | 90 |  | 8 | 0.6 | 4.8 | 10.4 | 10.6 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 90 |  | 1 | 5 | 5 | 5.44 | 5.56 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard Z-sharp | 90 |  | 10 | 0.6 | 6 | 7.19 | 7.32 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 90 |  | 1 | 10 | 10 | 5.44 | 5.56 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 90 |  | 12 | 1.2 | 14.4 | 6.97 | 7.13 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard Z-sharp | 90 |  | 20 | 0.6 | 12 | 7.33 | 7.46 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard Z-sharp | 90 |  | 32 | 0.6 | 19.2 | 6.75 | 6.89 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 90 |  | 12 | 1.2 | 14.4 | 6.09 | 6.22 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 90 |  | 64 | 0.6 | 38.4 | 5.97 | 6.07 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 100 |  | 2 | 1 | 2 | 8.35 | 8.49 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 100 |  | 6 | 0.6 | 3.6 | 13.5 | 13.8 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard Z-sharp | 100 |  | 8 | 0.6 | 4.8 | 14.3 | 14.5 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 100 |  | 1 | 5 | 5 | 7.49 | 7.62 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard Z-sharp | 100 |  | 10 | 0.6 | 6 | 9.91 | 10 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 100 |  | 1 | 10 | 10 | 7.49 | 7.62 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 100 |  | 12 | 1.2 | 14.4 | 9.6 | 9.77 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard Z-sharp | 100 |  | 20 | 0.6 | 12 | 10.1 | 10.2 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard Z-sharp | 100 |  | 32 | 0.6 | 19.2 | 9.3 | 9.44 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 100 |  | 12 | 1.2 | 14.4 | 8.38 | 8.53 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 100 |  | 64 | 0.6 | 38.4 | 8.22 | 8.32 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 110 |  | 2 | 1 | 2 | 11.2 | 11.4 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 110 |  | 6 | 0.6 | 3.6 | 18.1 | 18.4 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard Z-sharp | 110 |  | 8 | 0.6 | 4.8 | 19.2 | 19.4 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 110 |  | 1 | 5 | 5 | 10.1 | 10.2 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard Z-sharp | 110 |  | 10 | 0.6 | 6 | 13.3 | 13.4 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 110 |  | 1 | 10 | 10 | 10.1 | 10.2 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 110 |  | 12 | 1.2 | 14.4 | 12.9 | 13.1 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard Z-sharp | 110 |  | 20 | 0.6 | 12 | 13.5 | 13.7 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard Z-sharp | 110 |  | 32 | 0.6 | 19.2 | 12.5 | 12.6 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 110 |  | 12 | 1.2 | 14.4 | 11.3 | 11.4 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 110 |  | 64 | 0.6 | 38.4 | 11 | 11.1 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 120 |  | 2 | 1 | 2 | 14 | 14.2 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 120 |  | 6 | 0.6 | 3.6 | 22.7 | 23 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard Z-sharp | 120 |  | 8 | 0.6 | 4.8 | 23.9 | 24.1 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 120 |  | 1 | 5 | 5 | 12.6 | 12.7 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard Z-sharp | 120 |  | 10 | 0.6 | 6 | 16.6 | 16.7 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 120 |  | 1 | 10 | 10 | 12.6 | 12.7 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 120 |  | 12 | 1.2 | 14.4 | 16.1 | 16.3 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard Z-sharp | 120 |  | 20 | 0.6 | 12 | 16.9 | 17.1 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard Z-sharp | 120 |  | 32 | 0.6 | 19.2 | 15.6 | 15.8 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 120 |  | 12 | 1.2 | 14.4 | 14.1 | 14.2 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 120 |  | 64 | 0.6 | 38.4 | 13.8 | 13.9 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 130 |  | 2 | 1 | 2 | 16.9 | 17.2 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 130 |  | 6 | 0.6 | 3.6 | 27.4 | 27.8 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard Z-sharp | 130 |  | 8 | 0.6 | 4.8 | 29 | 29.3 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 130 |  | 1 | 5 | 5 | 15.2 | 15.4 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard Z-sharp | 130 |  | 10 | 0.6 | 6 | 20.1 | 20.3 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 130 |  | 1 | 10 | 10 | 15.2 | 15.4 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 130 |  | 12 | 1.2 | 14.4 | 19.5 | 19.8 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard Z-sharp | 130 |  | 20 | 0.6 | 12 | 20.5 | 20.7 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard Z-sharp | 130 |  | 32 | 0.6 | 19.2 | 18.9 | 19.1 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 130 |  | 12 | 1.2 | 14.4 | 17 | 17.3 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 130 |  | 64 | 0.6 | 38.4 | 16.7 | 16.8 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 140 |  | 2 | 1 | 2 | 20.3 | 20.5 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 140 |  | 6 | 0.6 | 3.6 | 32.9 | 33.3 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard Z-sharp | 140 |  | 8 | 0.6 | 4.8 | 34.8 | 35.1 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 140 |  | 1 | 5 | 5 | 18.2 | 18.4 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard Z-sharp | 140 |  | 10 | 0.6 | 6 | 24.1 | 24.3 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 140 |  | 1 | 10 | 10 | 18.2 | 18.4 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 140 |  | 12 | 1.2 | 14.4 | 23.4 | 23.6 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard Z-sharp | 140 |  | 20 | 0.6 | 12 | 24.6 | 24.8 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard Z-sharp | 140 |  | 32 | 0.6 | 19.2 | 22.6 | 22.9 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 140 |  | 12 | 1.2 | 14.4 | 20.4 | 20.9 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 140 |  | 64 | 0.6 | 38.4 | 20 | 20.2 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 100 | Sn | 2 | 1 | 2 | 1.54 | 1.54 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 100 | Sn | 6 | 0.6 | 3.6 | 2.5 | 2.49 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard Z-sharp | 100 | Sn | 8 | 0.6 | 4.8 | 2.65 | 2.63 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 100 | Sn | 1 | 5 | 5 | 1.38 | 1.38 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard Z-sharp | 100 | Sn | 10 | 0.6 | 6 | 1.84 | 1.82 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 100 | Sn | 1 | 10 | 10 | 1.38 | 1.38 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 100 | Sn | 12 | 1.2 | 14.4 | 1.77 | 1.77 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard Z-sharp | 100 | Sn | 20 | 0.6 | 12 | 1.87 | 1.86 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard Z-sharp | 100 | Sn | 32 | 0.6 | 19.2 | 1.72 | 1.71 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 100 | Sn | 12 | 1.2 | 14.4 | 1.55 | 1.54 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 100 | Sn | 64 | 0.6 | 38.4 | 1.52 | 1.51 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 140 | Sn | 2 | 1 | 2 | 6.49 | 6.55 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 140 | Sn | 6 | 0.6 | 3.6 | 10.5 | 10.6 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard Z-sharp | 140 | Sn | 8 | 0.6 | 4.8 | 11.2 | 11.2 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 140 | Sn | 1 | 5 | 5 | 5.82 | 5.87 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard Z-sharp | 140 | Sn | 10 | 0.6 | 6 | 7.74 | 7.79 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 140 | Sn | 1 | 10 | 10 | 5.82 | 5.87 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 140 | Sn | 12 | 1.2 | 14.4 | 7.47 | 7.54 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard Z-sharp | 140 | Sn | 20 | 0.6 | 12 | 7.88 | 7.94 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard Z-sharp | 140 | Sn | 32 | 0.6 | 19.2 | 7.25 | 7.3 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 140 | Sn | 12 | 1.2 | 14.4 | 6.52 | 6.58 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Standard | 140 | Sn | 64 | 0.6 | 38.4 | 6.42 | 6.46 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 70 |  | 2 | 1 | 2 | 2.32 | 2.29 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 70 |  | 6 | 0.6 | 3.6 | 3.75 | 3.71 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow Z-sharp | 70 |  | 8 | 0.6 | 4.8 | 3.95 | 3.88 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 70 |  | 1 | 5 | 5 | 2.08 | 20.6 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow Z-sharp | 70 |  | 10 | 0.6 | 6 | 2.74 | 2.69 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 70 |  | 1 | 10 | 10 | 2.08 | 20.6 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 70 |  | 12 | 1.2 | 14.4 | 2.67 | 2.64 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow Z-sharp | 70 |  | 20 | 0.6 | 12 | 2.79 | 2.74 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow Z-sharp | 70 |  | 32 | 0.6 | 19.2 | 2.58 | 2.54 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 70 |  | 12 | 1.2 | 14.4 | 2.33 | 2.3 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 70 |  | 64 | 0.6 | 38.4 | 2.27 | 2.23 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 80 |  | 2 | 1 | 2 | 3.74 | 3.67 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 80 |  | 6 | 0.6 | 3.6 | 6.05 | 5.94 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow Z-sharp | 80 |  | 8 | 0.6 | 4.8 | 6.38 | 6.23 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 80 |  | 1 | 5 | 5 | 3.35 | 3.29 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow Z-sharp | 80 |  | 10 | 0.6 | 6 | 4.43 | 4.32 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 80 |  | 1 | 10 | 10 | 3.35 | 3.29 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 80 |  | 12 | 1.2 | 14.4 | 4.3 | 4.22 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow Z-sharp | 80 |  | 20 | 0.6 | 12 | 4.51 | 4.4 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow Z-sharp | 80 |  | 32 | 0.6 | 19.2 | 4.16 | 4.07 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 80 |  | 12 | 1.2 | 14.4 | 3.76 | 3.68 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 80 |  | 64 | 0.6 | 38.4 | 3.67 | 3.58 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 90 |  | 2 | 1 | 2 | 5.54 | 5.42 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 90 |  | 6 | 0.6 | 3.6 | 8.97 | 8.77 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow Z-sharp | 90 |  | 8 | 0.6 | 4.8 | 9.47 | 9.2 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 90 |  | 1 | 5 | 5 | 4.97 | 4.86 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow Z-sharp | 90 |  | 10 | 0.6 | 6 | 6.57 | 6.38 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 90 |  | 1 | 10 | 10 | 4.97 | 4.86 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 90 |  | 12 | 1.2 | 14.4 | 6.38 | 6.24 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow Z-sharp | 90 |  | 20 | 0.6 | 12 | 6.69 | 6.5 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow Z-sharp | 90 |  | 32 | 0.6 | 19.2 | 6.17 | 6.01 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 90 |  | 12 | 1.2 | 14.4 | 5.56 | 5.44 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 90 |  | 64 | 0.6 | 38.4 | 5.45 | 5.29 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 100 |  | 2 | 1 | 2 | 7.66 | 7.48 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 100 |  | 6 | 0.6 | 3.6 | 12.4 | 12.1 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow Z-sharp | 100 |  | 8 | 0.6 | 4.8 | 13.1 | 12.7 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 100 |  | 1 | 5 | 5 | 6.87 | 6.71 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow Z-sharp | 100 |  | 10 | 0.6 | 6 | 9.08 | 8.81 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 100 |  | 1 | 10 | 10 | 6.87 | 6.71 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 100 |  | 12 | 1.2 | 14.4 | 8.81 | 8.6 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow Z-sharp | 100 |  | 20 | 0.6 | 12 | 9.25 | 8.98 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow Z-sharp | 100 |  | 32 | 0.6 | 19.2 | 8.53 | 8.3 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 100 |  | 12 | 1.2 | 14.4 | 7.69 | 7.51 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 100 |  | 64 | 0.6 | 38.4 | 7.53 | 7.31 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 110 |  | 2 | 1 | 2 | 10.3 | 10.1 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 110 |  | 6 | 0.6 | 3.6 | 16.7 | 16.3 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow Z-sharp | 110 |  | 8 | 0.6 | 4.8 | 17.6 | 17.1 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 110 |  | 1 | 5 | 5 | 9.26 | 9.03 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow Z-sharp | 110 |  | 10 | 0.6 | 6 | 12.2 | 11.9 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 110 |  | 1 | 10 | 10 | 9.26 | 9.03 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 110 |  | 12 | 1.2 | 14.4 | 11.9 | 11.6 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow Z-sharp | 110 |  | 20 | 0.6 | 12 | 12.5 | 12.1 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow Z-sharp | 110 |  | 32 | 0.6 | 19.2 | 11.5 | 11.2 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 110 |  | 12 | 1.2 | 14.4 | 10.4 | 10.1 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 110 |  | 64 | 0.6 | 38.4 | 10.1 | 9.83 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 120 |  | 2 | 1 | 2 | 12.9 | 12.6 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 120 |  | 6 | 0.6 | 3.6 | 20.9 | 20.4 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow Z-sharp | 120 |  | 8 | 0.6 | 4.8 | 22.1 | 21.4 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 120 |  | 1 | 5 | 5 | 11.6 | 11.3 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow Z-sharp | 120 |  | 10 | 0.6 | 6 | 15.3 | 14.9 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 120 |  | 1 | 10 | 10 | 11.6 | 11.3 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 120 |  | 12 | 1.2 | 14.4 | 14.9 | 14.5 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow Z-sharp | 120 |  | 20 | 0.6 | 12 | 15.6 | 15.1 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow Z-sharp | 120 |  | 32 | 0.6 | 19.2 | 14.4 | 14 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 120 |  | 12 | 1.2 | 14.4 | 13 | 12.7 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 120 |  | 64 | 0.6 | 38.4 | 12.7 | 12.3 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 130 |  | 2 | 1 | 2 | 15.7 | 15.4 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 130 |  | 6 | 0.6 | 3.6 | 25.4 | 24.9 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow Z-sharp | 130 |  | 8 | 0.6 | 4.8 | 26.8 | 26.1 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 130 |  | 1 | 5 | 5 | 14.1 | 13.8 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow Z-sharp | 130 |  | 10 | 0.6 | 6 | 18.6 | 18.1 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 130 |  | 1 | 10 | 10 | 14.1 | 13.8 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 130 |  | 12 | 1.2 | 14.4 | 18.1 | 17.7 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow Z-sharp | 130 |  | 20 | 0.6 | 12 | 19 | 18.5 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow Z-sharp | 130 |  | 32 | 0.6 | 19.2 | 17.5 | 17.1 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 130 |  | 12 | 1.2 | 14.4 | 15.8 | 15.4 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 130 |  | 64 | 0.6 | 38.4 | 15.4 | 15 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 140 |  | 2 | 1 | 2 | 18.8 | 18.4 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 140 |  | 6 | 0.6 | 3.6 | 30.5 | 29.8 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow Z-sharp | 140 |  | 8 | 0.6 | 4.8 | 32.3 | 31.4 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 140 |  | 1 | 5 | 5 | 16.9 | 16.5 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow Z-sharp | 140 |  | 10 | 0.6 | 6 | 22.4 | 21.8 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 140 |  | 1 | 10 | 10 | 16.9 | 16.5 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 140 |  | 12 | 1.2 | 14.4 | 21.7 | 21.2 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow Z-sharp | 140 |  | 20 | 0.6 | 12 | 22.8 | 22.2 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow Z-sharp | 140 |  | 32 | 0.6 | 19.2 | 21 | 20.5 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 140 |  | 12 | 1.2 | 14.4 | 18.9 | 18.5 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 140 |  | 64 | 0.6 | 38.4 | 18.5 | 18 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 100 | Sn | 2 | 1 | 2 | 1.45 | 1.4 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 100 | Sn | 6 | 0.6 | 3.6 | 2.35 | 2.27 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow Z-sharp | 100 | Sn | 8 | 0.6 | 4.8 | 2.49 | 2.4 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 100 | Sn | 1 | 5 | 5 | 1.3 | 1.26 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow Z-sharp | 100 | Sn | 10 | 0.6 | 6 | 1.73 | 1.66 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 100 | Sn | 1 | 10 | 10 | 1.3 | 1.26 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 100 | Sn | 12 | 1.2 | 14.4 | 1.67 | 1.61 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow Z-sharp | 100 | Sn | 20 | 0.6 | 12 | 1.76 | 1.69 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow Z-sharp | 100 | Sn | 32 | 0.6 | 19.2 | 1.62 | 1.56 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 100 | Sn | 12 | 1.2 | 14.4 | 1.46 | 1.41 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 100 | Sn | 64 | 0.6 | 38.4 | 1.43 | 1.38 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 140 | Sn | 2 | 1 | 2 | 6.16 | 6.06 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 140 | Sn | 6 | 0.6 | 3.6 | 9.97 | 9.8 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow Z-sharp | 140 | Sn | 8 | 0.6 | 4.8 | 10.6 | 10.4 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 140 | Sn | 1 | 5 | 5 | 5.52 | 5.43 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow Z-sharp | 140 | Sn | 10 | 0.6 | 6 | 7.35 | 7.2 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 140 | Sn | 1 | 10 | 10 | 5.52 | 5.43 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 140 | Sn | 12 | 1.2 | 14.4 | 7.09 | 6.97 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow Z-sharp | 140 | Sn | 20 | 0.6 | 12 | 7.49 | 7.34 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow Z-sharp | 140 | Sn | 32 | 0.6 | 19.2 | 6.88 | 6.75 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 140 | Sn | 12 | 1.2 | 14.4 | 6.19 | 6.08 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 16 | Narrow | 140 | Sn | 64 | 0.6 | 38.4 | 6.09 | 5.97 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 70 |  | 2 | 1 | 2 | 0.5 | 0.82 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 70 |  | 6 | 0.6 | 3.6 | 0.81 | 1.33 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow Z-sharp | 70 |  | 8 | 0.6 | 4.8 | 0.86 | 1.39 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 70 |  | 1 | 5 | 5 | 0.45 | 0.74 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow Z-sharp | 70 |  | 10 | 0.6 | 6 | 0.6 | 0.97 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 70 |  | 1 | 10 | 10 | 0.45 | 0.74 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 70 |  | 12 | 1.2 | 14.4 | 0.58 | 0.94 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow Z-sharp | 70 |  | 20 | 0.6 | 12 | 0.61 | 0.98 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow Z-sharp | 70 |  | 32 | 0.6 | 19.2 | 0.56 | 0.91 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 70 |  | 12 | 1.2 | 14.4 | 0.5 | 0.82 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 70 |  | 64 | 0.6 | 38.4 | 0.49 | 0.8 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 80 |  | 2 | 1 | 2 | 0.91 | 1.41 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 80 |  | 6 | 0.6 | 3.6 | 1.48 | 2.28 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow Z-sharp | 80 |  | 8 | 0.6 | 4.8 | 1.56 | 2.38 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 80 |  | 1 | 5 | 5 | 0.82 | 1.26 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow Z-sharp | 80 |  | 10 | 0.6 | 6 | 1.08 | 1.65 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 80 |  | 1 | 10 | 10 | 0.82 | 1.26 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 80 |  | 12 | 1.2 | 14.4 | 1.05 | 1.62 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow Z-sharp | 80 |  | 20 | 0.6 | 12 | 1.11 | 1.68 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow Z-sharp | 80 |  | 32 | 0.6 | 19.2 | 1.02 | 1.56 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 80 |  | 12 | 1.2 | 14.4 | 0.92 | 1.41 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 80 |  | 64 | 0.6 | 38.4 | 0.9 | 1.37 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 90 |  | 2 | 1 | 2 | 1.46 | 2.17 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 90 |  | 6 | 0.6 | 3.6 | 2.36 | 3.51 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow Z-sharp | 90 |  | 8 | 0.6 | 4.8 | 2.49 | 3.67 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 90 |  | 1 | 5 | 5 | 1.31 | 1.95 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow Z-sharp | 90 |  | 10 | 0.6 | 6 | 1.73 | 2.54 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 90 |  | 1 | 10 | 10 | 1.31 | 1.95 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 90 |  | 12 | 1.2 | 14.4 | 1.68 | 2.5 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow Z-sharp | 90 |  | 20 | 0.6 | 12 | 1.76 | 2.59 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow Z-sharp | 90 |  | 32 | 0.6 | 19.2 | 1.62 | 2.4 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 90 |  | 12 | 1.2 | 14.4 | 1.47 | 2.18 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 90 |  | 64 | 0.6 | 38.4 | 1.43 | 2.11 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 100 |  | 2 | 1 | 2 | 2.1 | 3.04 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 100 |  | 6 | 0.6 | 3.6 | 3.4 | 4.93 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow Z-sharp | 100 |  | 8 | 0.6 | 4.8 | 3.57 | 5.15 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 100 |  | 1 | 5 | 5 | 1.89 | 2.73 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow Z-sharp | 100 |  | 10 | 0.6 | 6 | 2.49 | 3.57 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 100 |  | 1 | 10 | 10 | 1.89 | 2.73 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 100 |  | 12 | 1.2 | 14.4 | 2.42 | 3.5 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow Z-sharp | 100 |  | 20 | 0.6 | 12 | 2.54 | 3.64 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow Z-sharp | 100 |  | 32 | 0.6 | 19.2 | 2.34 | 3.37 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 100 |  | 12 | 1.2 | 14.4 | 2.11 | 3.06 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 100 |  | 64 | 0.6 | 38.4 | 2.07 | 2.96 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 110 |  | 2 | 1 | 2 | 2.91 | 4.14 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 110 |  | 6 | 0.6 | 3.6 | 4.72 | 6.7 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow Z-sharp | 110 |  | 8 | 0.6 | 4.8 | 4.98 | 7.01 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 110 |  | 1 | 5 | 5 | 2.61 | 3.71 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow Z-sharp | 110 |  | 10 | 0.6 | 6 | 3.45 | 7.86 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 110 |  | 1 | 10 | 10 | 2.61 | 3.71 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 110 |  | 12 | 1.2 | 14.4 | 3.35 | 4.76 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow Z-sharp | 110 |  | 20 | 0.6 | 12 | 3.52 | 4.95 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow Z-sharp | 110 |  | 32 | 0.6 | 19.2 | 3.24 | 4.55 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 110 |  | 12 | 1.2 | 14.4 | 2.93 | 4.1 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 110 |  | 64 | 0.6 | 38.4 | 2.86 | 4.03 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 120 |  | 2 | 1 | 2 | 3.76 | 5.27 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 120 |  | 6 | 0.6 | 3.6 | 6.09 | 8.53 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow Z-sharp | 120 |  | 8 | 0.6 | 4.8 | 6.42 | 8.92 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 120 |  | 1 | 5 | 5 | 3.37 | 4.73 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow Z-sharp | 120 |  | 10 | 0.6 | 6 | 4.45 | 6.18 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 120 |  | 1 | 10 | 10 | 3.37 | 4.73 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 120 |  | 12 | 1.2 | 14.4 | 4.33 | 6.06 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow Z-sharp | 120 |  | 20 | 0.6 | 12 | 4.54 | 6.3 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow Z-sharp | 120 |  | 32 | 0.6 | 19.2 | 4.18 | 5.84 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 120 |  | 12 | 1.2 | 14.4 | 3.78 | 5.29 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 120 |  | 64 | 0.6 | 38.4 | 3.67 | 5.13 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 130 |  | 2 | 1 | 2 | 4.73 | 6.56 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 130 |  | 6 | 0.6 | 3.6 | 7.65 | 10.6 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow Z-sharp | 130 |  | 8 | 0.6 | 4.8 | 8.08 | 11.1 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 130 |  | 1 | 5 | 5 | 4.24 | 5.88 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow Z-sharp | 130 |  | 10 | 0.6 | 6 | 5.6 | 7.71 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 130 |  | 1 | 10 | 10 | 4.24 | 5.88 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 130 |  | 12 | 1.2 | 14.4 | 5.44 | 7.55 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow Z-sharp | 130 |  | 20 | 0.6 | 12 | 5.71 | 7.86 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow Z-sharp | 130 |  | 32 | 0.6 | 19.2 | 5.26 | 7.27 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 130 |  | 12 | 1.2 | 14.4 | 4.75 | 6.59 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 130 |  | 64 | 0.6 | 38.4 | 4.65 | 6.39 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 140 |  | 2 | 1 | 2 | 5.81 | 8 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 140 |  | 6 | 0.6 | 3.6 | 9.41 | 13 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow Z-sharp | 140 |  | 8 | 0.6 | 4.8 | 9.96 | 13.6 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 140 |  | 1 | 5 | 5 | 5.21 | 7.18 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow Z-sharp | 140 |  | 10 | 0.6 | 6 | 6.9 | 9.4 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 140 |  | 1 | 10 | 10 | 5.21 | 7.18 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 140 |  | 12 | 1.2 | 14.4 | 6.69 | 9.21 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow Z-sharp | 140 |  | 20 | 0.6 | 12 | 7.04 | 9.58 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow Z-sharp | 140 |  | 32 | 0.6 | 19.2 | 6.48 | 8.87 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 140 |  | 12 | 1.2 | 14.4 | 5.84 | 8.04 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 140 |  | 64 | 0.6 | 38.4 | 5.72 | 7.8 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 100 | Sn | 2 | 1 | 2 | 0.48 | 0.63 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 100 | Sn | 6 | 0.6 | 3.6 | 0.78 | 1.02 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow Z-sharp | 100 | Sn | 8 | 0.6 | 4.8 | 0.82 | 1.08 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 100 | Sn | 1 | 5 | 5 | 0.43 | 0.57 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow Z-sharp | 100 | Sn | 10 | 0.6 | 6 | 0.57 | 0.75 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 100 | Sn | 1 | 10 | 10 | 0.43 | 0.57 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 100 | Sn | 12 | 1.2 | 14.4 | 0.55 | 0.73 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow Z-sharp | 100 | Sn | 20 | 0.6 | 12 | 0.58 | 0.76 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow Z-sharp | 100 | Sn | 32 | 0.6 | 19.2 | 0.54 | 0.7 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 100 | Sn | 12 | 1.2 | 14.4 | 0.48 | 0.64 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 100 | Sn | 64 | 0.6 | 38.4 | 0.47 | 0.62 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 140 | Sn | 2 | 1 | 2 | 2.23 | 2.9 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 140 | Sn | 6 | 0.6 | 3.6 | 3.61 | 4.7 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow Z-sharp | 140 | Sn | 8 | 0.6 | 4.8 | 3.83 | 4.94 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 140 | Sn | 1 | 5 | 5 | 2 | 2.6 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow Z-sharp | 140 | Sn | 10 | 0.6 | 6 | 2.66 | 3.43 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 140 | Sn | 1 | 10 | 10 | 2 | 2.6 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 140 | Sn | 12 | 1.2 | 14.4 | 2.57 | 3.34 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow Z-sharp | 140 | Sn | 20 | 0.6 | 12 | 2.71 | 3.49 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow Z-sharp | 140 | Sn | 32 | 0.6 | 19.2 | 2.49 | 3.22 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 140 | Sn | 12 | 1.2 | 14.4 | 2.24 | 2.92 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Narrow | 140 | Sn | 64 | 0.6 | 38.4 | 2.2 | 2.84 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 70 |  | 2 | 1 | 2 | 0.58 | 1.05 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 70 |  | 6 | 0.6 | 3.6 | 0.94 | 1.7 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard Z-sharp | 70 |  | 8 | 0.6 | 4.8 | 0.99 | 1.79 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 70 |  | 1 | 5 | 5 | 0.52 | 0.94 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard Z-sharp | 70 |  | 10 | 0.6 | 6 | 0.69 | 1.24 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 70 |  | 1 | 10 | 10 | 0.52 | 0.94 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 70 |  | 12 | 1.2 | 14.4 | 0.66 | 1.21 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard Z-sharp | 70 |  | 20 | 0.6 | 12 | 0.7 | 1.27 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard Z-sharp | 70 |  | 32 | 0.6 | 19.2 | 0.64 | 1.17 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 70 |  | 12 | 1.2 | 14.4 | 0.58 | 1.05 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 70 |  | 64 | 0.6 | 38.4 | 0.57 | 1.03 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 80 |  | 2 | 1 | 2 | 1.04 | 1.79 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 80 |  | 6 | 0.6 | 3.6 | 1.69 | 2.9 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard Z-sharp | 80 |  | 8 | 0.6 | 4.8 | 1.79 | 3.04 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 80 |  | 1 | 5 | 5 | 0.94 | 1.6 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard Z-sharp | 80 |  | 10 | 0.6 | 6 | 1.24 | 2.11 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 80 |  | 1 | 10 | 10 | 0.94 | 1.6 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 80 |  | 12 | 1.2 | 14.4 | 1.2 | 2.06 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard Z-sharp | 80 |  | 20 | 0.6 | 12 | 1.26 | 2.15 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard Z-sharp | 80 |  | 32 | 0.6 | 19.2 | 1.16 | 1.99 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 80 |  | 12 | 1.2 | 14.4 | 1.05 | 1.8 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 80 |  | 64 | 0.6 | 38.4 | 1.03 | 1.75 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 90 |  | 2 | 1 | 2 | 1.66 | 2.73 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 90 |  | 6 | 0.6 | 3.6 | 2.68 | 4.42 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard Z-sharp | 90 |  | 8 | 0.6 | 4.8 | 2.84 | 4.64 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 90 |  | 1 | 5 | 5 | 1.49 | 2.45 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard Z-sharp | 90 |  | 10 | 0.6 | 6 | 1.97 | 3.22 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 90 |  | 1 | 10 | 10 | 1.49 | 2.45 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 90 |  | 12 | 1.2 | 14.4 | 1.91 | 3.14 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard Z-sharp | 90 |  | 20 | 0.6 | 12 | 2 | 3.28 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard Z-sharp | 90 |  | 32 | 0.6 | 19.2 | 1.85 | 3.03 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 90 |  | 12 | 1.2 | 14.4 | 1.66 | 2.74 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 90 |  | 64 | 0.6 | 38.4 | 1.63 | 2.67 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 100 |  | 2 | 1 | 2 | 2.38 | 3.81 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 100 |  | 6 | 0.6 | 3.6 | 3.85 | 6.17 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard Z-sharp | 100 |  | 8 | 0.6 | 4.8 | 4.07 | 6.48 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 100 |  | 1 | 5 | 5 | 2.13 | 3.42 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard Z-sharp | 100 |  | 10 | 0.6 | 6 | 2.82 | 4.5 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 100 |  | 1 | 10 | 10 | 2.13 | 3.42 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 100 |  | 12 | 1.2 | 14.4 | 2.73 | 4.39 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard Z-sharp | 100 |  | 20 | 0.6 | 12 | 2.88 | 4.58 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard Z-sharp | 100 |  | 32 | 0.6 | 19.2 | 2.65 | 4.23 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 100 |  | 12 | 1.2 | 14.4 | 2.39 | 3.83 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 100 |  | 64 | 0.6 | 38.4 | 2.34 | 3.73 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 110 |  | 2 | 1 | 2 | 3.28 | 5.16 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 110 |  | 6 | 0.6 | 3.6 | 5.31 | 8.36 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard Z-sharp | 110 |  | 8 | 0.6 | 4.8 | 5.62 | 8.77 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 110 |  | 1 | 5 | 5 | 2.94 | 4.63 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard Z-sharp | 110 |  | 10 | 0.6 | 6 | 3.9 | 6.08 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 110 |  | 1 | 10 | 10 | 2.94 | 4.63 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 110 |  | 12 | 1.2 | 14.4 | 3.77 | 5.94 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard Z-sharp | 110 |  | 20 | 0.6 | 12 | 3.97 | 6.2 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard Z-sharp | 110 |  | 32 | 0.6 | 19.2 | 3.66 | 5.73 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 110 |  | 12 | 1.2 | 14.4 | 3.29 | 5.18 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 110 |  | 64 | 0.6 | 38.4 | 3.23 | 5.04 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 120 |  | 2 | 1 | 2 | 4.22 | 6.53 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 120 |  | 6 | 0.6 | 3.6 | 6.84 | 10.6 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard Z-sharp | 120 |  | 8 | 0.6 | 4.8 | 7.23 | 11.1 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 120 |  | 1 | 5 | 5 | 3.79 | 5.86 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard Z-sharp | 120 |  | 10 | 0.6 | 6 | 5.01 | 7.7 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 120 |  | 1 | 10 | 10 | 3.79 | 5.86 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 120 |  | 12 | 1.2 | 14.4 | 4.86 | 7.52 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard Z-sharp | 120 |  | 20 | 0.6 | 12 | 5.11 | 7.85 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard Z-sharp | 120 |  | 32 | 0.6 | 19.2 | 4.7 | 7.25 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 120 |  | 12 | 1.2 | 14.4 | 4.24 | 6.56 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 120 |  | 64 | 0.6 | 38.4 | 4.16 | 6.39 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 130 |  | 2 | 1 | 2 | 5.29 | 8.1 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 130 |  | 6 | 0.6 | 3.6 | 8.57 | 13.1 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard Z-sharp | 130 |  | 8 | 0.6 | 4.8 | 9.07 | 13.8 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 130 |  | 1 | 5 | 5 | 4.75 | 7.26 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard Z-sharp | 130 |  | 10 | 0.6 | 6 | 6.29 | 9.55 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 130 |  | 1 | 10 | 10 | 4.75 | 7.26 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 130 |  | 12 | 1.2 | 14.4 | 6.09 | 9.32 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard Z-sharp | 130 |  | 20 | 0.6 | 12 | 6.41 | 9.74 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard Z-sharp | 130 |  | 32 | 0.6 | 19.2 | 5.9 | 8.99 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 130 |  | 12 | 1.2 | 14.4 | 5.31 | 8.13 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 130 |  | 64 | 0.6 | 38.4 | 5.22 | 7.92 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 140 |  | 2 | 1 | 2 | 6.49 | 9.84 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 140 |  | 6 | 0.6 | 3.6 | 10.5 | 15.9 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard Z-sharp | 140 |  | 8 | 0.6 | 4.8 | 11.1 | 16.7 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 140 |  | 1 | 5 | 5 | 5.82 | 8.83 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard Z-sharp | 140 |  | 10 | 0.6 | 6 | 7.72 | 11.6 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 140 |  | 1 | 10 | 10 | 5.82 | 8.83 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 140 |  | 12 | 1.2 | 14.4 | 7.47 | 11.3 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard Z-sharp | 140 |  | 20 | 0.6 | 12 | 7.86 | 11.8 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard Z-sharp | 140 |  | 32 | 0.6 | 19.2 | 7.24 | 10.9 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 140 |  | 12 | 1.2 | 14.4 | 6.52 | 9.89 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 140 |  | 64 | 0.6 | 38.4 | 6.4 | 9.63 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 100 | Sn | 2 | 1 | 2 | 0.54 | 0.77 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 100 | Sn | 6 | 0.6 | 3.6 | 0.87 | 1.25 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard Z-sharp | 100 | Sn | 8 | 0.6 | 4.8 | 0.92 | 1.32 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 100 | Sn | 1 | 5 | 5 | 0.48 | 0.69 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard Z-sharp | 100 | Sn | 10 | 0.6 | 6 | 0.64 | 0.91 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 100 | Sn | 1 | 10 | 10 | 0.48 | 0.69 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 100 | Sn | 12 | 1.2 | 14.4 | 0.62 | 0.89 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard Z-sharp | 100 | Sn | 20 | 0.6 | 12 | 0.65 | 0.93 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard Z-sharp | 100 | Sn | 32 | 0.6 | 19.2 | 0.6 | 0.86 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 100 | Sn | 12 | 1.2 | 14.4 | 0.54 | 0.78 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 100 | Sn | 64 | 0.6 | 38.4 | 0.53 | 0.76 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 140 | Sn | 2 | 1 | 2 | 2.45 | 3.47 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 140 | Sn | 6 | 0.6 | 3.6 | 3.97 | 5.62 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard Z-sharp | 140 | Sn | 8 | 0.6 | 4.8 | 4.21 | 5.94 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 140 | Sn | 1 | 5 | 5 | 2.2 | 3.12 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard Z-sharp | 140 | Sn | 10 | 0.6 | 6 | 2.92 | 4.12 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 140 | Sn | 1 | 10 | 10 | 2.2 | 3.12 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 140 | Sn | 12 | 1.2 | 14.4 | 2.82 | 4 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard Z-sharp | 140 | Sn | 20 | 0.6 | 12 | 2.97 | 4.2 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard Z-sharp | 140 | Sn | 32 | 0.6 | 19.2 | 2.73 | 3.87 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 140 | Sn | 12 | 1.2 | 14.4 | 2.46 | 3.49 | Somatom Drive | Siemens
Somatom Drive | 5 | 1 | 32 | Standard | 140 | Sn | 64 | 0.6 | 38.4 | 2.42 | 3.42 | Somatom Drive | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 70 |  | 2 | 1 | 2 | 2.53 | 2.72 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 70 |  | 6 | 0.6 | 3.6 | 4.21 | 4.53 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 70 |  | 1 | 5 | 5 | 2.44 | 2.63 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 70 |  | 1 | 10 | 10 | 2.44 | 2.63 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 70 |  | 12 | 1.2 | 14.4 | 3.48 | 3.75 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 70 |  | 32 | 0.6 | 19.2 | 2.75 | 2.96 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 70 |  | 32 | 0.6 | 19.2 | 2.96 | 3.18 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 70 |  | 32 | 0.6 | 19.2 | 3.36 | 3.62 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 70 |  | 48 | 1.2 | 57.6 | 3.05 | 3.28 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 70 |  | 64 | 1.2 | 76.8 | 2.9 | 3.12 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 70 |  | 96 | 0.6 | 57.6 | 2.72 | 2.93 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 70 |  | 24 | 1.2 | 28.8 | 3 | 3.23 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 70 |  | 48 | 1.2 | 57.6 | 2.73 | 2.93 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 80 |  | 2 | 1 | 2 | 4.04 | 4.24 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 80 |  | 6 | 0.6 | 3.6 | 6.73 | 7.07 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 80 |  | 1 | 5 | 5 | 3.9 | 4.1 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 80 |  | 1 | 10 | 10 | 3.9 | 4.1 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 80 |  | 12 | 1.2 | 14.4 | 5.56 | 5.85 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 80 |  | 32 | 0.6 | 19.2 | 4.4 | 4.63 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 80 |  | 32 | 0.6 | 19.2 | 4.72 | 4.97 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 80 |  | 32 | 0.6 | 19.2 | 5.37 | 5.65 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 80 |  | 48 | 1.2 | 57.6 | 4.87 | 5.12 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 80 |  | 64 | 1.2 | 76.8 | 4.62 | 4.86 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 80 |  | 96 | 0.6 | 57.6 | 4.35 | 4.58 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 80 |  | 24 | 1.2 | 28.8 | 4.79 | 5.03 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 80 |  | 48 | 1.2 | 57.6 | 4.35 | 4.58 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 90 |  | 2 | 1 | 2 | 5.93 | 6.08 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 90 |  | 6 | 0.6 | 3.6 | 9.88 | 10.13 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 90 |  | 1 | 5 | 5 | 5.73 | 5.87 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 90 |  | 1 | 10 | 10 | 5.73 | 5.87 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 90 |  | 12 | 1.2 | 14.4 | 8.17 | 8.37 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 90 |  | 32 | 0.6 | 19.2 | 6.46 | 6.62 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 90 |  | 32 | 0.6 | 19.2 | 6.94 | 7.11 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 90 |  | 32 | 0.6 | 19.2 | 7.89 | 8.08 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 90 |  | 48 | 1.2 | 57.6 | 7.16 | 7.33 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 90 |  | 64 | 1.2 | 76.8 | 6.79 | 6.96 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 90 |  | 96 | 0.6 | 57.6 | 6.39 | 6.55 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 90 |  | 24 | 1.2 | 28.8 | 7.03 | 7.2 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 90 |  | 48 | 1.2 | 57.6 | 6.4 | 6.55 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 100 |  | 2 | 1 | 2 | 8.08 | 8.23 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 100 |  | 6 | 0.6 | 3.6 | 13.47 | 13.72 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 100 |  | 1 | 5 | 5 | 7.81 | 7.95 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 100 |  | 1 | 10 | 10 | 7.81 | 7.95 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 100 |  | 12 | 1.2 | 14.4 | 11.14 | 11.34 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 100 |  | 32 | 0.6 | 19.2 | 8.81 | 8.97 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 100 |  | 32 | 0.6 | 19.2 | 9.45 | 9.63 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 100 |  | 32 | 0.6 | 19.2 | 10.75 | 10.95 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 100 |  | 48 | 1.2 | 57.6 | 9.76 | 9.94 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 100 |  | 64 | 1.2 | 76.8 | 9.26 | 9.43 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 100 |  | 96 | 0.6 | 57.6 | 8.71 | 8.87 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 100 |  | 24 | 1.2 | 28.8 | 9.58 | 9.76 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 100 |  | 48 | 1.2 | 57.6 | 8.72 | 8.88 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 110 |  | 2 | 1 | 2 | 10.45 | 10.61 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 110 |  | 6 | 0.6 | 3.6 | 17.42 | 17.68 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 110 |  | 1 | 5 | 5 | 10.1 | 10.25 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 110 |  | 1 | 10 | 10 | 10.1 | 10.25 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 110 |  | 12 | 1.2 | 14.4 | 14.41 | 14.62 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 110 |  | 32 | 0.6 | 19.2 | 11.39 | 11.56 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 110 |  | 32 | 0.6 | 19.2 | 12.23 | 12.41 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 110 |  | 32 | 0.6 | 19.2 | 13.91 | 14.11 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 110 |  | 48 | 1.2 | 57.6 | 12.62 | 12.81 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 110 |  | 64 | 1.2 | 76.8 | 11.98 | 12.16 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 110 |  | 96 | 0.6 | 57.6 | 11.27 | 11.44 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 110 |  | 24 | 1.2 | 28.8 | 12.4 | 12.58 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 110 |  | 48 | 1.2 | 57.6 | 11.28 | 11.45 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 120 |  | 2 | 1 | 2 | 13.06 | 13.22 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 120 |  | 6 | 0.6 | 3.6 | 21.76 | 22.04 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 120 |  | 1 | 5 | 5 | 12.62 | 12.77 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 120 |  | 1 | 10 | 10 | 12.62 | 12.77 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 120 |  | 12 | 1.2 | 14.4 | 18 | 18.22 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 120 |  | 32 | 0.6 | 19.2 | 14.23 | 14.41 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 120 |  | 32 | 0.6 | 19.2 | 15.28 | 15.47 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 120 |  | 32 | 0.6 | 19.2 | 17.37 | 17.59 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 120 |  | 48 | 1.2 | 57.6 | 15.76 | 15.96 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 120 |  | 64 | 1.2 | 76.8 | 14.96 | 15.15 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 120 |  | 96 | 0.6 | 57.6 | 14.07 | 14.25 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 120 |  | 24 | 1.2 | 28.8 | 15.49 | 15.68 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 120 |  | 48 | 1.2 | 57.6 | 14.09 | 14.26 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 130 |  | 2 | 1 | 2 | 15.73 | 15.92 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 130 |  | 6 | 0.6 | 3.6 | 26.21 | 26.53 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 130 |  | 1 | 5 | 5 | 15.2 | 15.38 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 130 |  | 1 | 10 | 10 | 15.2 | 15.38 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 130 |  | 12 | 1.2 | 14.4 | 21.68 | 21.93 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 130 |  | 32 | 0.6 | 19.2 | 17.14 | 17.34 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 130 |  | 32 | 0.6 | 19.2 | 18.4 | 18.62 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 130 |  | 32 | 0.6 | 19.2 | 20.92 | 21.17 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 130 |  | 48 | 1.2 | 57.6 | 18.99 | 19.21 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 130 |  | 64 | 1.2 | 76.8 | 18.02 | 18.24 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 130 |  | 96 | 0.6 | 57.6 | 16.95 | 17.15 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 130 |  | 24 | 1.2 | 28.8 | 18.65 | 18.87 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 130 |  | 48 | 1.2 | 57.6 | 16.97 | 17.17 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 140 |  | 2 | 1 | 2 | 18.55 | 18.75 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 140 |  | 6 | 0.6 | 3.6 | 30.92 | 31.25 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 140 |  | 1 | 5 | 5 | 17.93 | 18.12 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 140 |  | 1 | 10 | 10 | 17.93 | 18.12 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 140 |  | 12 | 1.2 | 14.4 | 25.57 | 25.84 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 140 |  | 32 | 0.6 | 19.2 | 20.22 | 20.43 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 140 |  | 32 | 0.6 | 19.2 | 21.7 | 21.94 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 140 |  | 32 | 0.6 | 19.2 | 24.68 | 24.94 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 140 |  | 48 | 1.2 | 57.6 | 22.4 | 22.64 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 140 |  | 64 | 1.2 | 76.8 | 21.26 | 21.49 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 140 |  | 96 | 0.6 | 57.6 | 20 | 20.21 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 140 |  | 24 | 1.2 | 28.8 | 22 | 22.24 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 140 |  | 48 | 1.2 | 57.6 | 20.01 | 20.23 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 150 |  | 2 | 1 | 2 | 21.49 | 21.73 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 150 |  | 6 | 0.6 | 3.6 | 35.82 | 36.22 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 150 |  | 1 | 5 | 5 | 20.77 | 21 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 150 |  | 1 | 10 | 10 | 20.77 | 21 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 150 |  | 12 | 1.2 | 14.4 | 29.62 | 29.95 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 150 |  | 32 | 0.6 | 19.2 | 23.42 | 23.68 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 150 |  | 32 | 0.6 | 19.2 | 25.15 | 25.42 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 150 |  | 32 | 0.6 | 19.2 | 28.59 | 28.91 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 150 |  | 48 | 1.2 | 57.6 | 25.95 | 26.24 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 150 |  | 64 | 1.2 | 76.8 | 24.63 | 24.9 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 150 |  | 96 | 0.6 | 57.6 | 23.17 | 23.42 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 150 |  | 24 | 1.2 | 28.8 | 25.49 | 25.77 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 150 |  | 48 | 1.2 | 57.6 | 23.19 | 23.44 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 100 | Sn | 2 | 1 | 2 | 0.79 | 0.77 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 100 | Sn | 6 | 0.6 | 3.6 | 1.31 | 1.29 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 100 | Sn | 1 | 5 | 5 | 0.76 | 0.75 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 100 | Sn | 1 | 10 | 10 | 0.76 | 0.75 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 100 | Sn | 12 | 1.2 | 14.4 | 1.08 | 1.07 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 100 | Sn | 32 | 0.6 | 19.2 | 0.86 | 0.84 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 100 | Sn | 32 | 0.6 | 19.2 | 0.92 | 0.9 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 100 | Sn | 32 | 0.6 | 19.2 | 1.05 | 1.03 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 100 | Sn | 48 | 1.2 | 57.6 | 0.95 | 0.93 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 100 | Sn | 64 | 1.2 | 76.8 | 0.9 | 0.89 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 100 | Sn | 96 | 0.6 | 57.6 | 0.85 | 0.83 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 100 | Sn | 24 | 1.2 | 28.8 | 0.93 | 0.92 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 100 | Sn | 48 | 1.2 | 57.6 | 0.85 | 0.83 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 120 | Sn | 2 | 1 | 2 | 1.92 | 1.92 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 120 | Sn | 6 | 0.6 | 3.6 | 3.21 | 3.2 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 120 | Sn | 1 | 5 | 5 | 1.86 | 1.85 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 120 | Sn | 1 | 10 | 10 | 1.86 | 1.85 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 120 | Sn | 12 | 1.2 | 14.4 | 2.65 | 2.65 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 120 | Sn | 32 | 0.6 | 19.2 | 2.1 | 2.09 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 120 | Sn | 32 | 0.6 | 19.2 | 2.25 | 2.25 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 120 | Sn | 32 | 0.6 | 19.2 | 2.56 | 2.55 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 120 | Sn | 48 | 1.2 | 57.6 | 2.32 | 2.32 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 120 | Sn | 64 | 1.2 | 76.8 | 2.2 | 2.2 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 120 | Sn | 96 | 0.6 | 57.6 | 2.07 | 2.07 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 120 | Sn | 24 | 1.2 | 28.8 | 2.28 | 2.28 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 120 | Sn | 48 | 1.2 | 57.6 | 2.07 | 2.07 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 150 | Sn | 2 | 1 | 2 | 4.99 | 5.01 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 150 | Sn | 6 | 0.6 | 3.6 | 8.31 | 8.35 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 150 | Sn | 1 | 5 | 5 | 4.82 | 4.84 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 150 | Sn | 1 | 10 | 10 | 4.82 | 4.84 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 150 | Sn | 12 | 1.2 | 14.4 | 6.87 | 6.9 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 150 | Sn | 32 | 0.6 | 19.2 | 5.43 | 5.46 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 150 | Sn | 32 | 0.6 | 19.2 | 5.83 | 5.86 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 150 | Sn | 32 | 0.6 | 19.2 | 6.63 | 6.66 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 150 | Sn | 48 | 1.2 | 57.6 | 6.02 | 6.05 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 150 | Sn | 64 | 1.2 | 76.8 | 5.71 | 5.74 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 150 | Sn | 96 | 0.6 | 57.6 | 5.38 | 5.4 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard Z-sharp | 150 | Sn | 24 | 1.2 | 28.8 | 5.91 | 5.94 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Standard | 150 | Sn | 48 | 1.2 | 57.6 | 5.38 | 5.4 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 70 |  | 2 | 1 | 2 | 2.28 | 2.25 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 70 |  | 6 | 0.6 | 3.6 | 3.8 | 3.76 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 70 |  | 1 | 5 | 5 | 2.2 | 2.18 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 70 |  | 1 | 10 | 10 | 2.2 | 2.18 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 70 |  | 12 | 1.2 | 14.4 | 3.14 | 3.11 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 70 |  | 32 | 0.6 | 19.2 | 2.48 | 2.46 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 70 |  | 32 | 0.6 | 19.2 | 2.67 | 2.64 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 70 |  | 32 | 0.6 | 19.2 | 3.03 | 3 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 70 |  | 48 | 1.2 | 57.6 | 2.75 | 2.72 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 70 |  | 64 | 1.2 | 76.8 | 2.61 | 2.58 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 70 |  | 96 | 0.6 | 57.6 | 2.46 | 2.43 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 70 |  | 24 | 1.2 | 28.8 | 2.7 | 2.64 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 70 |  | 48 | 1.2 | 57.6 | 2.46 | 2.43 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 80 |  | 2 | 1 | 2 | 3.67 | 3.6 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 80 |  | 6 | 0.6 | 3.6 | 6.12 | 5.99 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 80 |  | 1 | 5 | 5 | 3.55 | 3.48 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 80 |  | 1 | 10 | 10 | 3.55 | 3.48 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 80 |  | 12 | 1.2 | 14.4 | 5.06 | 4.96 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 80 |  | 32 | 0.6 | 19.2 | 4 | 3.92 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 80 |  | 32 | 0.6 | 19.2 | 4.29 | 4.21 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 80 |  | 32 | 0.6 | 19.2 | 4.88 | 4.78 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 80 |  | 48 | 1.2 | 57.6 | 4.43 | 4.34 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 80 |  | 64 | 1.2 | 76.8 | 4.21 | 4.12 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 80 |  | 96 | 0.6 | 57.6 | 3.96 | 3.88 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 80 |  | 24 | 1.2 | 28.8 | 4.35 | 4.27 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 80 |  | 48 | 1.2 | 57.6 | 3.96 | 3.88 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 90 |  | 2 | 1 | 2 | 5.43 | 5.29 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 90 |  | 6 | 0.6 | 3.6 | 9.05 | 8.82 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 90 |  | 1 | 5 | 5 | 5.25 | 5.11 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 90 |  | 1 | 10 | 10 | 52.25 | 5.11 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 90 |  | 12 | 1.2 | 14.4 | 7.49 | 7.29 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 90 |  | 32 | 0.6 | 19.2 | 5.92 | 5.76 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 90 |  | 32 | 0.6 | 19.2 | 6.35 | 6.19 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 90 |  | 32 | 0.6 | 19.2 | 7.23 | 7.04 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 90 |  | 48 | 1.2 | 57.6 | 6.56 | 6.39 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 90 |  | 64 | 1.2 | 76.8 | 6.22 | 6.06 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 90 |  | 96 | 0.6 | 57.6 | 5.86 | 5.7 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 90 |  | 24 | 1.2 | 28.8 | 6.44 | 6.27 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 90 |  | 48 | 1.2 | 57.6 | 5.86 | 5.71 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 100 |  | 2 | 1 | 2 | 7.45 | 7.23 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 100 |  | 6 | 0.6 | 3.6 | 12.43 | 12.04 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 100 |  | 1 | 5 | 5 | 7.2 | 6.98 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 100 |  | 1 | 10 | 10 | 7.2 | 6.98 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 100 |  | 12 | 1.2 | 14.4 | 10.26 | 9.96 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 100 |  | 32 | 0.6 | 19.2 | 8.12 | 7.87 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 100 |  | 32 | 0.6 | 19.2 | 8.71 | 8.45 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 100 |  | 32 | 0.6 | 19.2 | 9.91 | 9.61 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 100 |  | 48 | 1.2 | 57.6 | 8.99 | 8.72 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 100 |  | 64 | 1.2 | 76.8 | 8.53 | 8.28 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 100 |  | 96 | 0.6 | 57.6 | 8.03 | 7.79 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 100 |  | 24 | 1.2 | 28.8 | 8.83 | 8.57 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 100 |  | 48 | 1.2 | 57.6 | 8.03 | 7.79 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 110 |  | 2 | 1 | 2 | 9.66 | 9.37 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 110 |  | 6 | 0.6 | 3.6 | 16.11 | 15.62 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 110 |  | 1 | 5 | 5 | 9.34 | 9.06 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 110 |  | 1 | 10 | 10 | 9.34 | 9.06 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 110 |  | 12 | 1.2 | 14.4 | 13.32 | 12.92 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 110 |  | 32 | 0.6 | 19.2 | 10.53 | 10.22 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 110 |  | 32 | 0.6 | 19.2 | 11.31 | 10.97 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 110 |  | 32 | 0.6 | 19.2 | 12.86 | 12.47 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 110 |  | 48 | 1.2 | 57.6 | 11.67 | 11.32 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 110 |  | 64 | 1.2 | 76.8 | 11.07 | 10.74 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 110 |  | 96 | 0.6 | 57.6 | 10.42 | 10.1 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 110 |  | 24 | 1.2 | 28.8 | 11.46 | 11.12 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 110 |  | 48 | 1.2 | 57.6 | 10.43 | 10.11 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 120 |  | 2 | 1 | 2 | 12.1 | 11.74 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 120 |  | 6 | 0.6 | 3.6 | 20.16 | 19.57 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 120 |  | 1 | 5 | 5 | 11.69 | 11.34 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 120 |  | 1 | 10 | 10 | 11.69 | 11.34 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 120 |  | 12 | 1.2 | 14.4 | 16.67 | 16.18 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 120 |  | 32 | 0.6 | 19.2 | 13.18 | 12.79 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 120 |  | 32 | 0.6 | 19.2 | 14.15 | 13.73 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 120 |  | 32 | 0.6 | 19.2 | 16.09 | 15.62 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 120 |  | 48 | 1.2 | 57.6 | 14.6 | 14.17 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 120 |  | 64 | 1.2 | 76.8 | 13.86 | 13.45 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 120 |  | 96 | 0.6 | 57.6 | 13.04 | 12.65 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 120 |  | 24 | 1.2 | 28.8 | 14.34 | 13.92 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 120 |  | 48 | 1.2 | 57.6 | 13.05 | 12.66 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 130 |  | 2 | 1 | 2 | 16.63 | 14.19 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 130 |  | 6 | 0.6 | 3.6 | 24.38 | 23.66 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 130 |  | 1 | 5 | 5 | 14.13 | 13.71 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 130 |  | 1 | 10 | 10 | 14.13 | 13.71 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 130 |  | 12 | 1.2 | 14.4 | 20.16 | 19.56 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 130 |  | 32 | 0.6 | 19.2 | 15.94 | 15.47 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 130 |  | 32 | 0.6 | 19.2 | 17.12 | 16.6 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 130 |  | 32 | 0.6 | 19.2 | 19.46 | 18.88 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 130 |  | 48 | 1.2 | 57.6 | 17.66 | 17.13 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 130 |  | 64 | 1.2 | 76.8 | 16.76 | 16.6 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 130 |  | 96 | 0.6 | 57.6 | 15.77 | 15.3 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 130 |  | 24 | 1.2 | 28.8 | 17.35 | 16.83 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 130 |  | 48 | 1.2 | 57.6 | 15.78 | 15.31 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 140 |  | 2 | 1 | 2 | 17.3 | 16.79 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 140 |  | 6 | 0.6 | 3.6 | 28.83 | 27.98 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 140 |  | 1 | 5 | 5 | 16.71 | 16.22 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 140 |  | 1 | 10 | 10 | 16.71 | 16.22 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 140 |  | 12 | 1.2 | 14.4 | 23.84 | 23.14 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 140 |  | 32 | 0.6 | 19.2 | 18.85 | 18.3 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 140 |  | 32 | 0.6 | 19.2 | 20.24 | 19.64 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 140 |  | 32 | 0.6 | 19.2 | 23.01 | 22.33 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 140 |  | 48 | 1.2 | 57.6 | 20.88 | 20.27 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 140 |  | 64 | 1.2 | 76.8 | 19.82 | 19.24 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 140 |  | 96 | 0.6 | 57.6 | 18.64 | 19.1 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 140 |  | 24 | 1.2 | 28.8 | 20.51 | 19.91 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 140 |  | 48 | 1.2 | 57.6 | 18.66 | 18.11 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 150 |  | 2 | 1 | 2 | 20.1 | 19.51 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 150 |  | 6 | 0.6 | 3.6 | 33.49 | 32.51 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 150 |  | 1 | 5 | 5 | 19.42 | 18.85 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 150 |  | 1 | 10 | 10 | 19.42 | 18.85 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 150 |  | 12 | 1.2 | 14.4 | 27.7 | 26.88 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 150 |  | 32 | 0.6 | 19.2 | 21.9 | 21.26 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 150 |  | 32 | 0.6 | 19.2 | 23.51 | 22.82 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 150 |  | 32 | 0.6 | 19.2 | 26.73 | 2595 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 150 |  | 48 | 1.2 | 57.6 | 24.26 | 23.55 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 150 |  | 64 | 1.2 | 76.8 | 23.03 | 22.35 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 150 |  | 96 | 0.6 | 57.6 | 21.66 | 21.03 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 150 |  | 24 | 1.2 | 28.8 | 23.83 | 23.13 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 150 |  | 48 | 1.2 | 57.6 | 21.68 | 21.04 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 100 | Sn | 2 | 1 | 2 | 0.74 | 0.71 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 100 | Sn | 6 | 0.6 | 3.6 | 1.23 | 1.18 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 100 | Sn | 1 | 5 | 5 | 0.71 | 0.68 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 100 | Sn | 1 | 10 | 10 | 0.71 | 0.68 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 100 | Sn | 12 | 1.2 | 14.4 | 1.02 | 0.98 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 100 | Sn | 32 | 0.6 | 19.2 | 1 | 0.77 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 100 | Sn | 32 | 0.6 | 19.2 | 0.86 | 0.83 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 100 | Sn | 32 | 0.6 | 19.2 | 98 | 0.94 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 100 | Sn | 48 | 1.2 | 57.6 | 0.89 | 0.85 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 100 | Sn | 64 | 1.2 | 76.8 | 0.85 | 0.81 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 100 | Sn | 96 | 0.6 | 57.6 | 0.8 | 0.76 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 100 | Sn | 24 | 1.2 | 28.8 | 0.88 | 0.84 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 100 | Sn | 48 | 1.2 | 57.6 | 0.8 | 0.76 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 120 | Sn | 2 | 1 | 2 | 1.83 | 1.77 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 120 | Sn | 6 | 0.6 | 3.6 | 3.04 | 2.95 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 120 | Sn | 1 | 5 | 5 | 1.76 | 1.71 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 120 | Sn | 1 | 10 | 10 | 1.76 | 1.71 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 120 | Sn | 12 | 1.2 | 14.4 | 2.52 | 2.44 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 120 | Sn | 32 | 0.6 | 19.2 | 1.99 | 1.93 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 120 | Sn | 32 | 0.6 | 19.2 | 2.14 | 2.07 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 120 | Sn | 32 | 0.6 | 19.2 | 2.43 | 2.35 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 120 | Sn | 48 | 1.2 | 57.6 | 2.2 | 2.14 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 120 | Sn | 64 | 1.2 | 76.8 | 2.09 | 2.03 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 120 | Sn | 96 | 0.6 | 57.6 | 1.97 | 1.91 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 120 | Sn | 24 | 1.2 | 28.8 | 2.17 | 2.1 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 120 | Sn | 48 | 1.2 | 57.6 | 1.97 | 1.91 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 150 | Sn | 2 | 1 | 2 | 4.78 | 4.66 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 150 | Sn | 6 | 0.6 | 3.6 | 7.97 | 7.77 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 150 | Sn | 1 | 5 | 5 | 4.62 | 4.5 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 150 | Sn | 1 | 10 | 10 | 4.62 | 4.5 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 150 | Sn | 12 | 1.2 | 14.4 | 6.59 | 6.42 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 150 | Sn | 32 | 0.6 | 19.2 | 5.21 | 5.08 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 150 | Sn | 32 | 0.6 | 19.2 | 5.59 | 5.45 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 150 | Sn | 32 | 0.6 | 19.2 | 6.36 | 6.2 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 150 | Sn | 48 | 1.2 | 57.6 | 5.77 | 5.63 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 150 | Sn | 64 | 1.2 | 76.8 | 5.48 | 5.34 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 150 | Sn | 96 | 0.6 | 57.6 | 5.15 | 5.02 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow Z-sharp | 150 | Sn | 24 | 1.2 | 28.8 | 567 | 5.53 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 16 | Narrow | 150 | Sn | 48 | 1.2 | 57.6 | 5.16 | 5.03 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 70 |  | 2 | 1 | 2 | 0.58 | 1.09 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 70 |  | 6 | 0.6 | 3.6 | 0.97 | 1.81 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 70 |  | 1 | 5 | 5 | 0.56 | 1.05 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 70 |  | 1 | 10 | 10 | 0.56 | 1.05 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 70 |  | 12 | 1.2 | 14.4 | 0.8 | 1.5 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 70 |  | 32 | 0.6 | 19.2 | 0.64 | 1.18 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 70 |  | 32 | 0.6 | 19.2 | 0.68 | 1.27 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 70 |  | 32 | 0.6 | 19.2 | 0.78 | 1.44 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 70 |  | 48 | 1.2 | 57.6 | 0.7 | 1.31 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 70 |  | 64 | 1.2 | 76.8 | 0.67 | 1.24 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 70 |  | 96 | 0.6 | 57.6 | 0.63 | 1.17 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 70 |  | 24 | 1.2 | 28.8 | 0.69 | 1.29 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 70 |  | 48 | 1.2 | 57.6 | 0.63 | 1.17 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 80 |  | 2 | 1 | 2 | 1.03 | 1.78 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 80 |  | 6 | 0.6 | 3.6 | 1.72 | 2.97 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 80 |  | 1 | 5 | 5 | 1 | 1.72 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 80 |  | 1 | 10 | 10 | 1 | 1.72 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 80 |  | 12 | 1.2 | 14.4 | 1.42 | 2.46 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 80 |  | 32 | 0.6 | 19.2 | 1.12 | 1.94 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 80 |  | 32 | 0.6 | 19.2 | 1.21 | 2.08 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 80 |  | 32 | 0.6 | 19.2 | 1.37 | 2.37 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 80 |  | 48 | 1.2 | 57.6 | 1.24 | 2.15 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 80 |  | 64 | 1.2 | 76.8 | 1.18 | 2.04 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 80 |  | 96 | 0.6 | 57.6 | 1.11 | 1.92 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 80 |  | 24 | 1.2 | 28.8 | 1.22 | 2.11 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 80 |  | 48 | 1.2 | 57.6 | 1.11 | 1.92 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 90 |  | 2 | 1 | 2 | 1.62 | 2.68 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 90 |  | 6 | 0.6 | 3.6 | 2.69 | 4.46 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 90 |  | 1 | 5 | 5 | 1.56 | 2.59 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 90 |  | 1 | 10 | 10 | 1.56 | 2.59 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 90 |  | 12 | 1.2 | 14.4 | 2.23 | 3.69 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 90 |  | 32 | 0.6 | 19.2 | 1.76 | 2.92 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 90 |  | 32 | 0.6 | 19.2 | 1.89 | 3.13 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 90 |  | 32 | 0.6 | 19.2 | 2.15 | 3.56 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 90 |  | 48 | 1.2 | 57.6 | 1.95 | 3.23 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 90 |  | 64 | 1.2 | 76.8 | 1.85 | 3.07 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 90 |  | 96 | 0.6 | 57.6 | 1.7 | 2.89 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 90 |  | 24 | 1.2 | 28.8 | 1.92 | 3.18 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 90 |  | 48 | 1.2 | 57.6 | 1.74 | 2.89 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 100 |  | 2 | 1 | 2 | 2.32 | 3.72 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 100 |  | 6 | 0.6 | 3.6 | 3.86 | 6.2 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 100 |  | 1 | 5 | 5 | 2.24 | 3.59 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 100 |  | 1 | 10 | 10 | 2.24 | 3.59 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 100 |  | 12 | 1.2 | 14.4 | 3.19 | 5.13 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 100 |  | 32 | 0.6 | 19.2 | 2.52 | 4.05 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 100 |  | 32 | 0.6 | 19.2 | 2.71 | 4.35 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 100 |  | 32 | 0.6 | 19.2 | 3.08 | 4.95 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 100 |  | 48 | 1.2 | 57.6 | 2.8 | 4.49 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 100 |  | 64 | 1.2 | 76.8 | 2.65 | 4.26 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 100 |  | 96 | 0.6 | 57.6 | 2.5 | 4.01 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 100 |  | 24 | 1.2 | 28.8 | 2.75 | 4.41 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 100 |  | 48 | 1.2 | 57.6 | 2.5 | 4.01 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 110 |  | 2 | 1 | 2 | 3.11 | 4.84 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 110 |  | 6 | 0.6 | 3.6 | 5.19 | 8.07 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 110 |  | 1 | 5 | 5 | 3.01 | 4.68 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 110 |  | 1 | 10 | 10 | 3.01 | 4.68 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 110 |  | 12 | 1.2 | 14.4 | 4.29 | 6.67 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 110 |  | 32 | 0.6 | 19.2 | 3.39 | 5.28 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 110 |  | 32 | 0.6 | 19.2 | 3.64 | 5.66 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 110 |  | 32 | 0.6 | 19.2 | 4.14 | 6.44 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 110 |  | 48 | 1.2 | 57.6 | 3.76 | 5.85 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 110 |  | 64 | 1.2 | 76.8 | 3.57 | 5.55 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 110 |  | 96 | 0.6 | 57.6 | 3.36 | 5.22 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 110 |  | 24 | 1.2 | 28.8 | 3.69 | 5.74 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 110 |  | 48 | 1.2 | 57.6 | 3.36 | 5.22 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 120 |  | 2 | 1 | 2 | 4.02 | 6.2 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 120 |  | 6 | 0.6 | 3.6 | 6.69 | 10.34 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 120 |  | 1 | 5 | 5 | 3.88 | 5.99 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 120 |  | 1 | 10 | 10 | 3.88 | 5.99 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 120 |  | 12 | 1.2 | 14.4 | 5.53 | 8.55 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 120 |  | 32 | 0.6 | 19.2 | 4.38 | 6.76 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 120 |  | 32 | 0.6 | 19.2 | 4.7 | 7.26 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 120 |  | 32 | 0.6 | 19.2 | 5.34 | 8.25 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 120 |  | 48 | 1.2 | 57.6 | 4.85 | 7.49 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 120 |  | 64 | 1.2 | 76.8 | 4.6 | 7.11 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 120 |  | 96 | 0.6 | 57.6 | 4.33 | 6.69 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 120 |  | 24 | 1.2 | 28.8 | 4.76 | 7.36 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 120 |  | 48 | 1.2 | 57.6 | 4.33 | 6.69 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 130 |  | 2 | 1 | 2 | 4.97 | 7.58 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 130 |  | 6 | 0.6 | 3.6 | 8.28 | 12.63 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 130 |  | 1 | 5 | 5 | 4.8 | 7.32 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 130 |  | 1 | 10 | 10 | 4.8 | 7.32 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 130 |  | 12 | 1.2 | 14.4 | 6.84 | 10.44 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 130 |  | 32 | 0.6 | 19.2 | 5.41 | 8.26 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 130 |  | 32 | 0.6 | 19.2 | 5.81 | 8.86 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 130 |  | 32 | 0.6 | 19.2 | 6.6 | 10.08 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 130 |  | 48 | 1.2 | 57.6 | 5.99 | 9.15 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 130 |  | 64 | 1.2 | 76.8 | 5.69 | 8.67 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 130 |  | 96 | 0.6 | 57.6 | 5.35 | 8.17 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 130 |  | 24 | 1.2 | 28.8 | 5.89 | 8.98 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 130 |  | 48 | 1.2 | 57.6 | 5.36 | 8.17 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 140 |  | 2 | 1 | 2 | 5.98 | 9.05 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 140 |  | 6 | 0.6 | 3.6 | 9.97 | 15.08 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 140 |  | 1 | 5 | 5 | 5.78 | 8.74 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 140 |  | 1 | 10 | 10 | 5.78 | 8.74 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 140 |  | 12 | 1.2 | 14.4 | 8.25 | 12.47 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 140 |  | 32 | 0.6 | 19.2 | 6.52 | 9.86 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 140 |  | 32 | 0.6 | 19.2 | 7 | 10.58 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 140 |  | 32 | 0.6 | 19.2 | 7.96 | 12.03 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 140 |  | 48 | 1.2 | 57.6 | 7.23 | 10.92 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 140 |  | 64 | 1.2 | 76.8 | 6.86 | 10.36 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 140 |  | 96 | 0.6 | 57.6 | 6.45 | 9.75 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 140 |  | 24 | 1.2 | 28.8 | 7.1 | 10.73 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 140 |  | 48 | 1.2 | 57.6 | 6.46 | 9.76 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 150 |  | 2 | 1 | 2 | 7.07 | 10.61 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 150 |  | 6 | 0.6 | 3.6 | 11.79 | 17.68 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 150 |  | 1 | 5 | 5 | 6.83 | 10.25 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 150 |  | 1 | 10 | 10 | 6.83 | 10.25 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 150 |  | 12 | 1.2 | 14.4 | 9.75 | 14.62 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 150 |  | 32 | 0.6 | 19.2 | 7.71 | 11.56 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 150 |  | 32 | 0.6 | 19.2 | 8.25 | 12.41 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 150 |  | 32 | 0.6 | 19.2 | 9.41 | 14.11 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 150 |  | 48 | 1.2 | 57.6 | 8.54 | 12.81 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 150 |  | 64 | 1.2 | 76.8 | 8.11 | 12.15 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 150 |  | 96 | 0.6 | 57.6 | 8.62 | 11.43 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 150 |  | 24 | 1.2 | 28.8 | 8.39 | 12.58 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 150 |  | 48 | 1.2 | 57.6 | 7.63 | 11.44 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 100 | Sn | 2 | 1 | 2 | 0.28 | 0.39 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 100 | Sn | 6 | 0.6 | 3.6 | 0.47 | 0.66 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 100 | Sn | 1 | 5 | 5 | 0.27 | 0.38 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 100 | Sn | 1 | 10 | 10 | 0.27 | 0.38 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 100 | Sn | 12 | 1.2 | 14.4 | 0.38 | 0.54 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 100 | Sn | 32 | 0.6 | 19.2 | 0.3 | 0.43 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 100 | Sn | 32 | 0.6 | 19.2 | 0.33 | 0.46 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 100 | Sn | 32 | 0.6 | 19.2 | 0.37 | 0.52 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 100 | Sn | 48 | 1.2 | 57.6 | 0.34 | 0.47 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 100 | Sn | 64 | 1.2 | 76.8 | 0.32 | 0.45 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 100 | Sn | 96 | 0.6 | 57.6 | 0.3 | 0.42 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 100 | Sn | 24 | 1.2 | 28.8 | 0.33 | 0.47 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 100 | Sn | 48 | 1.2 | 57.6 | 0.3 | 0.42 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 120 | Sn | 2 | 1 | 2 | 0.71 | 1 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 120 | Sn | 6 | 0.6 | 3.6 | 1.19 | 1.67 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 120 | Sn | 1 | 5 | 5 | 0.69 | 0.97 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 120 | Sn | 1 | 10 | 10 | 0.69 | 0.97 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 120 | Sn | 12 | 1.2 | 14.4 | 0.99 | 1.38 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 120 | Sn | 32 | 0.6 | 19.2 | 0.78 | 1.09 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 120 | Sn | 32 | 0.6 | 19.2 | 0.84 | 1.17 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 120 | Sn | 32 | 0.6 | 19.2 | 0.95 | 1.33 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 120 | Sn | 48 | 1.2 | 57.6 | 0.86 | 1.21 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 120 | Sn | 64 | 1.2 | 76.8 | 0.82 | 1.15 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 120 | Sn | 96 | 0.6 | 57.6 | 0.77 | 1.08 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 120 | Sn | 24 | 1.2 | 28.8 | 0.85 | 1.19 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 120 | Sn | 48 | 1.2 | 57.6 | 0.77 | 1.08 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 150 | Sn | 2 | 1 | 2 | 1.96 | 2.7 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 150 | Sn | 6 | 0.6 | 3.6 | 3.27 | 4.51 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 150 | Sn | 1 | 5 | 5 | 1.89 | 2.61 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 150 | Sn | 1 | 10 | 10 | 1.89 | 2.61 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 150 | Sn | 12 | 1.2 | 14.4 | 2.7 | 3.73 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 150 | Sn | 32 | 0.6 | 19.2 | 2.14 | 2.95 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 150 | Sn | 32 | 0.6 | 19.2 | 2.29 | 3.16 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 150 | Sn | 32 | 0.6 | 19.2 | 2.61 | 3.6 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 150 | Sn | 48 | 1.2 | 57.6 | 2.37 | 3.26 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 150 | Sn | 64 | 1.2 | 76.8 | 2.25 | 3.1 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 150 | Sn | 96 | 0.6 | 57.6 | 2.11 | 2.92 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard Z-sharp | 150 | Sn | 24 | 1.2 | 28.8 | 2.32 | 3.21 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Standard | 150 | Sn | 48 | 1.2 | 57.6 | 2.11 | 2.92 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 70 |  | 2 | 1 | 2 | 0.51 | 0.83 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 70 |  | 6 | 0.6 | 3.6 | 0.84 | 1.39 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 70 |  | 1 | 5 | 5 | 0.49 | 0.81 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 70 |  | 1 | 10 | 10 | 0.49 | 0.81 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 70 |  | 12 | 1.2 | 14.4 | 0.7 | 1.15 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 70 |  | 32 | 0.6 | 19.2 | 0.55 | 0.91 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 70 |  | 32 | 0.6 | 19.2 | 0.59 | 0.98 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 70 |  | 32 | 0.6 | 19.2 | 0.67 | 1.11 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 70 |  | 48 | 1.2 | 57.6 | 0.61 | 1.01 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 70 |  | 64 | 1.2 | 76.8 | 0.58 | 0.96 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 70 |  | 96 | 0.6 | 57.6 | 0.54 | 0.9 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 70 |  | 24 | 1.2 | 28.8 | 0.6 | 0.99 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 70 |  | 48 | 1.2 | 57.6 | 0.55 | 0.9 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 80 |  | 2 | 1 | 2 | 0.9 | 1.36 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 80 |  | 6 | 0.6 | 3.6 | 1.49 | 2.27 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 80 |  | 1 | 5 | 5 | 0.87 | 1.32 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 80 |  | 1 | 10 | 10 | 0.87 | 1.32 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 80 |  | 12 | 1.2 | 14.4 | 1.23 | 1.88 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 80 |  | 32 | 0.6 | 19.2 | 0.98 | 1.48 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 80 |  | 32 | 0.6 | 19.2 | 1.05 | 1.59 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 80 |  | 32 | 0.6 | 19.2 | 1.19 | 1.81 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 80 |  | 48 | 1.2 | 57.6 | 1.08 | 1.64 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 80 |  | 64 | 1.2 | 76.8 | 1.03 | 1.56 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 80 |  | 96 | 0.6 | 57.6 | 0.97 | 1.47 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 80 |  | 24 | 1.2 | 28.8 | 1.06 | 1.61 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 80 |  | 48 | 1.2 | 57.6 | 0.97 | 1.47 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 90 |  | 2 | 1 | 2 | 1.4 | 2.1 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 90 |  | 6 | 0.6 | 3.6 | 2.34 | 3.49 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 90 |  | 1 | 5 | 5 | 1.35 | 2.02 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 90 |  | 1 | 10 | 10 | 1.35 | 2.02 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 90 |  | 12 | 1.2 | 14.4 | 1.93 | 2.89 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 90 |  | 32 | 0.6 | 19.2 | 1.53 | 2.28 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 90 |  | 32 | 0.6 | 19.2 | 1.64 | 2.45 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 90 |  | 32 | 0.6 | 19.2 | 1.87 | 2.79 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 90 |  | 48 | 1.2 | 57.6 | 1.69 | 2.53 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 90 |  | 64 | 1.2 | 76.8 | 1.61 | 2.4 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 90 |  | 96 | 0.6 | 57.6 | 1.51 | 2.26 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 90 |  | 24 | 1.2 | 28.8 | 1.66 | 2.49 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 90 |  | 48 | 1.2 | 57.6 | 1.51 | 2.26 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 100 |  | 2 | 1 | 2 | 2.04 | 2.94 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 100 |  | 6 | 0.6 | 3.6 | 3.41 | 4.9 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 100 |  | 1 | 5 | 5 | 1.98 | 2.84 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 100 |  | 1 | 10 | 10 | 1.98 | 2.84 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 100 |  | 12 | 1.2 | 14.4 | 2.82 | 4.05 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 100 |  | 32 | 0.6 | 19.2 | 2.23 | 3.2 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 100 |  | 32 | 0.6 | 19.2 | 2.39 | 3.44 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 100 |  | 32 | 0.6 | 19.2 | 2.72 | 3.91 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 100 |  | 48 | 1.2 | 57.6 | 2.47 | 3.55 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 100 |  | 64 | 1.2 | 76.8 | 2.34 | 3.37 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 100 |  | 96 | 0.6 | 57.6 | 2.2 | 3.17 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 100 |  | 24 | 1.2 | 28.8 | 2.42 | 3.48 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 100 |  | 48 | 1.2 | 57.6 | 2.21 | 3.17 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 110 |  | 2 | 1 | 2 | 2.75 | 3.89 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 110 |  | 6 | 0.6 | 3.6 | 4.59 | 6.48 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 110 |  | 1 | 5 | 5 | 2.66 | 3.76 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 110 |  | 1 | 10 | 10 | 2.66 | 3.76 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 110 |  | 12 | 1.2 | 14.4 | 3.79 | 5.36 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 110 |  | 32 | 0.6 | 19.2 | 3 | 4.24 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 110 |  | 32 | 0.6 | 19.2 | 3.22 | 4.55 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 110 |  | 32 | 0.6 | 19.2 | 3.66 | 5.17 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 110 |  | 48 | 1.2 | 57.6 | 3.32 | 4.69 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 110 |  | 64 | 1.2 | 76.8 | 3.5 | 4.45 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 110 |  | 96 | 0.6 | 57.6 | 2.97 | 4.19 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 110 |  | 24 | 1.2 | 28.8 | 3.26 | 4.61 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 110 |  | 48 | 1.2 | 57.6 | 2.97 | 4.19 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 120 |  | 2 | 1 | 2 | 3.56 | 4.95 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 120 |  | 6 | 0.6 | 3.6 | 5.94 | 8.26 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 120 |  | 1 | 5 | 5 | 3.44 | 4.79 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 120 |  | 1 | 10 | 10 | 3.44 | 4.79 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 120 |  | 12 | 1.2 | 14.4 | 4.91 | 6.83 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 120 |  | 32 | 0.6 | 19.2 | 3.88 | 5.4 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 120 |  | 32 | 0.6 | 19.2 | 4.17 | 5.8 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 120 |  | 32 | 0.6 | 19.2 | 4.74 | 6.59 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 120 |  | 48 | 1.2 | 57.6 | 4.3 | 5.98 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 120 |  | 64 | 1.2 | 76.8 | 4.08 | 5.68 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 120 |  | 96 | 0.6 | 57.6 | 3.84 | 5.34 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 120 |  | 24 | 1.2 | 28.8 | 4.23 | 5.87 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 120 |  | 48 | 1.2 | 57.6 | 3.85 | 5.34 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 130 |  | 2 | 1 | 2 | 4.42 | 6.08 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 130 |  | 6 | 0.6 | 3.6 | 7.37 | 10.13 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 130 |  | 1 | 5 | 5 | 4.27 | 5.87 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 130 |  | 1 | 10 | 10 | 4.27 | 5.87 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 130 |  | 12 | 1.2 | 14.4 | 609 | 8.38 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 130 |  | 32 | 0.6 | 19.2 | 4.82 | 6.63 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 130 |  | 32 | 0.6 | 19.2 | 5.17 | 7.11 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 130 |  | 32 | 0.6 | 19.2 | 5.88 | 8.09 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 130 |  | 48 | 1.2 | 57.6 | 5.34 | 7.34 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 130 |  | 64 | 1.2 | 76.8 | 5.07 | 6.97 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 130 |  | 96 | 0.6 | 57.6 | 4.77 | 6.55 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 130 |  | 24 | 1.2 | 28.8 | 5.24 | 7.21 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 130 |  | 48 | 1.2 | 57.6 | 4.77 | 6.56 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 140 |  | 2 | 1 | 2 | 5.34 | 7.28 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 140 |  | 6 | 0.6 | 3.6 | 8.9 | 12.14 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 140 |  | 1 | 5 | 5 | 5.16 | 7.04 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 140 |  | 1 | 10 | 10 | 5.16 | 7.04 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 140 |  | 12 | 1.2 | 14.4 | 7.36 | 10.04 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 140 |  | 32 | 0.6 | 19.2 | 5.82 | 7.94 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 140 |  | 32 | 0.6 | 19.2 | 6.25 | 8.52 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 140 |  | 32 | 0.6 | 19.2 | 7.1 | 9.69 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 140 |  | 48 | 1.2 | 57.6 | 6.45 | 8.79 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 140 |  | 64 | 1.2 | 76.8 | 6.12 | 8.35 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 140 |  | 96 | 0.6 | 57.6 | 5.76 | 7.85 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 140 |  | 24 | 1.2 | 28.8 | 6.33 | 8.64 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 140 |  | 48 | 1.2 | 57.6 | 5.76 | 7.86 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 150 |  | 2 | 1 | 2 | 6.33 | 8.56 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 150 |  | 6 | 0.6 | 3.6 | 10.56 | 14.27 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 150 |  | 1 | 5 | 5 | 6.12 | 8.27 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 150 |  | 1 | 10 | 10 | 6.12 | 8.27 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 150 |  | 12 | 1.2 | 14.4 | 8.73 | 11.8 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 150 |  | 32 | 0.6 | 19.2 | 6.9 | 9.33 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 150 |  | 32 | 0.6 | 19.2 | 7.41 | 10.02 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 150 |  | 32 | 0.6 | 19.2 | 8.42 | 11.39 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 150 |  | 48 | 1.2 | 57.6 | 7.65 | 10.34 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 150 |  | 64 | 1.2 | 76.8 | 7.26 | 9.81 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 150 |  | 96 | 0.6 | 57.6 | 6.83 | 9.23 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 150 |  | 24 | 1.2 | 28.8 | 7.51 | 10.16 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 150 |  | 48 | 1.2 | 57.6 | 6.83 | 9.24 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 100 | Sn | 2 | 1 | 2 | 0.25 | 0.32 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 100 | Sn | 6 | 0.6 | 3.6 | 0.42 | 0.53 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 100 | Sn | 1 | 5 | 5 | 0.24 | 0.31 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 100 | Sn | 1 | 10 | 10 | 0.24 | 0.31 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 100 | Sn | 12 | 1.2 | 14.4 | 0.35 | 0.44 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 100 | Sn | 32 | 0.6 | 19.2 | 0.27 | 0.35 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 100 | Sn | 32 | 0.6 | 19.2 | 0.29 | 0.38 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 100 | Sn | 32 | 0.6 | 19.2 | 0.33 | 0.43 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 100 | Sn | 48 | 1.2 | 57.6 | 0.3 | 0.39 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 100 | Sn | 64 | 1.2 | 76.8 | 0.29 | 0.37 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 100 | Sn | 96 | 0.6 | 57.6 | 0.27 | 0.35 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 100 | Sn | 24 | 1.2 | 28.8 | 0.3 | 0.38 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 100 | Sn | 48 | 1.2 | 57.6 | 0.27 | 0.35 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 120 | Sn | 2 | 1 | 2 | 0.65 | 0.83 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 120 | Sn | 6 | 0.6 | 3.6 | 1.08 | 1.38 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 120 | Sn | 1 | 5 | 5 | 0.62 | 0.8 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 120 | Sn | 1 | 10 | 10 | 0.62 | 0.8 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 120 | Sn | 12 | 1.2 | 14.4 | 0.89 | 1.14 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 120 | Sn | 32 | 0.6 | 19.2 | 0.7 | 0.9 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 120 | Sn | 32 | 0.6 | 19.2 | 0.76 | 0.97 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 120 | Sn | 32 | 0.6 | 19.2 | 0.86 | 1.1 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 120 | Sn | 48 | 1.2 | 57.6 | 0.78 | 1 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 120 | Sn | 64 | 1.2 | 76.8 | 0.74 | 0.95 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 120 | Sn | 96 | 0.6 | 57.6 | 0.7 | 0.89 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 120 | Sn | 24 | 1.2 | 28.8 | 0.77 | 0.98 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 120 | Sn | 48 | 1.2 | 57.6 | 0.7 | 0.89 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 150 | Sn | 2 | 1 | 2 | 1.78 | 2.26 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 150 | Sn | 6 | 0.6 | 3.6 | 2.97 | 3.77 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 150 | Sn | 1 | 5 | 5 | 1.72 | 2.18 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 150 | Sn | 1 | 10 | 10 | 1.72 | 2.18 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 150 | Sn | 12 | 1.2 | 14.4 | 2.46 | 3.12 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 150 | Sn | 32 | 0.6 | 19.2 | 1.94 | 2.46 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 150 | Sn | 32 | 0.6 | 19.2 | 2.09 | 2.65 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 150 | Sn | 32 | 0.6 | 19.2 | 2.37 | 3.01 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 150 | Sn | 48 | 1.2 | 57.6 | 2.15 | 2.73 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 150 | Sn | 64 | 1.2 | 76.8 | 2.04 | 2.59 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 150 | Sn | 96 | 0.6 | 57.6 | 1.92 | 2.44 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow Z-sharp | 150 | Sn | 24 | 1.2 | 28.8 | 2.12 | 2.68 | Somatom Force | Siemens
Somatom Force | 4 | 1 | 32 | Narrow | 150 | Sn | 48 | 1.2 | 57.6 | 1.92 | 2.44 | Somatom Force | Siemens
Somatom Flash A tube | 3 | 1 | 16 | Wide | 70 |  | 2 | 1 | 2 | 2.9 | 3.3 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 16 | Wide | 70 |  | 6 | 0.6 | 3.6 | 4.7 | 5.3 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 16 | Wide | 70 |  | 8 | 0.6 | 2.4 | 5 | 5.6 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 16 | Wide | 70 |  | 1 | 5 | 5 | 2.6 | 2.9 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 16 | Wide | 70 |  | 10 | 0.6 | 6 | 3.4 | 3.9 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 16 | Wide | 70 |  | 1 | 10 | 10 | 2.6 | 2.9 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 16 | Wide | 70 |  | 12 | 1.2 | 14.4 | 3.3 | 3.8 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 16 | Wide | 70 |  | 20 | 0.6 | 12 | 3.5 | 4 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 16 | Wide | 70 |  | 32 | 0.6 | 19.2 | 3.3 | 3.7 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 16 | Wide | 70 |  | 32 | 1.2 | 38.4 | 2.9 | 3.3 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 16 | Wide | 70 |  | 64 | 0.6 | 38.4 | 2.8 | 3.2 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 16 | Wide | 80 |  | 2 | 1 | 2 | 4.6 | 5 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 16 | Wide | 80 |  | 6 | 0.6 | 3.6 | 7.4 | 8.2 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 16 | Wide | 80 |  | 8 | 0.6 | 2.4 | 7.9 | 8.6 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 16 | Wide | 80 |  | 1 | 5 | 5 | 4.1 | 4.5 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 16 | Wide | 80 |  | 10 | 0.6 | 6 | 5.4 | 6 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 16 | Wide | 80 |  | 1 | 10 | 10 | 4.1 | 4.5 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 16 | Wide | 80 |  | 12 | 1.2 | 14.4 | 5.3 | 5.8 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 16 | Wide | 80 |  | 20 | 0.6 | 12 | 5.6 | 6.1 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 16 | Wide | 80 |  | 32 | 0.6 | 19.2 | 5.2 | 5.7 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 16 | Wide | 80 |  | 32 | 1.2 | 38.4 | 4.6 | 5.1 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 16 | Wide | 80 |  | 64 | 0.6 | 38.4 | 4.5 | 5 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 16 | Wide | 100 |  | 2 | 1 | 2 | 9.2 | 9.8 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 16 | Wide | 100 |  | 6 | 0.6 | 3.6 | 14.9 | 15.9 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 16 | Wide | 100 |  | 8 | 0.6 | 2.4 | 15.8 | 16.8 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 16 | Wide | 100 |  | 1 | 5 | 5 | 8.2 | 8.8 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 16 | Wide | 100 |  | 10 | 0.6 | 6 | 10.9 | 11.6 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 16 | Wide | 100 |  | 1 | 10 | 10 | 8.2 | 8.8 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 16 | Wide | 100 |  | 12 | 1.2 | 14.4 | 10.6 | 11.3 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 16 | Wide | 100 |  | 20 | 0.6 | 12 | 11.2 | 11.9 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 16 | Wide | 100 |  | 32 | 0.6 | 19.2 | 10.4 | 11 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 16 | Wide | 100 |  | 32 | 1.2 | 38.4 | 9.2 | 9.8 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 16 | Wide | 100 |  | 64 | 0.6 | 38.4 | 9.1 | 9.6 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 16 | Wide | 120 |  | 2 | 1 | 2 | 14.8 | 15.5 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 16 | Wide | 120 |  | 6 | 0.6 | 3.6 | 23.9 | 25.1 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 16 | Wide | 120 |  | 8 | 0.6 | 2.4 | 25.4 | 26.6 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 16 | Wide | 120 |  | 1 | 5 | 5 | 13.3 | 13.9 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 16 | Wide | 120 |  | 10 | 0.6 | 6 | 17.6 | 18.4 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 16 | Wide | 120 |  | 1 | 10 | 10 | 13.3 | 13.9 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 16 | Wide | 120 |  | 12 | 1.2 | 14.4 | 17 | 17.8 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 16 | Wide | 120 |  | 20 | 0.6 | 12 | 17.9 | 18.8 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 16 | Wide | 120 |  | 32 | 0.6 | 19.2 | 16.7 | 17.5 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 16 | Wide | 120 |  | 32 | 1.2 | 38.4 | 14.8 | 15.6 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 16 | Wide | 120 |  | 64 | 0.6 | 38.4 | 14.6 | 15.3 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 16 | Wide | 140 |  | 2 | 1 | 2 | 22 | 22.9 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 16 | Wide | 140 |  | 6 | 0.6 | 3.6 | 35.6 | 37.1 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 16 | Wide | 140 |  | 8 | 0.6 | 2.4 | 37.8 | 39.3 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 16 | Wide | 140 |  | 1 | 5 | 5 | 19.7 | 20.5 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 16 | Wide | 140 |  | 10 | 0.6 | 6 | 26.2 | 27.2 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 16 | Wide | 140 |  | 1 | 10 | 10 | 19.7 | 20.5 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 16 | Wide | 140 |  | 12 | 1.2 | 14.4 | 25.3 | 26.3 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 16 | Wide | 140 |  | 20 | 0.6 | 12 | 26.7 | 27.8 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 16 | Wide | 140 |  | 32 | 0.6 | 19.2 | 24.9 | 25.9 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 16 | Wide | 140 |  | 32 | 1.2 | 38.4 | 22.1 | 23 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 16 | Wide | 140 |  | 64 | 0.6 | 38.4 | 21.8 | 22.6 | Somatom Flash A tube | Siemens
Somatom Flash B tube | 17 | 1 | 16 | Wide | 140 | Sn | 2 | 1 | 2 | 6.6 | 6.6 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 16 | Wide | 140 | Sn | 6 | 0.6 | 3.6 | 10.6 | 10.7 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 16 | Wide | 140 | Sn | 8 | 0.6 | 2.4 | 11.3 | 11.4 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 16 | Wide | 140 | Sn | 1 | 5 | 5 | 5.9 | 5.9 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 16 | Wide | 140 | Sn | 10 | 0.6 | 6 | 7.8 | 7.9 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 16 | Wide | 140 | Sn | 1 | 10 | 10 | 5.9 | 5.9 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 16 | Wide | 140 | Sn | 12 | 1.2 | 14.4 | 7.5 | 7.6 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 16 | Wide | 140 | Sn | 20 | 0.6 | 12 | 8 | 8 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 16 | Wide | 140 | Sn | 32 | 0.6 | 19.2 | 7.4 | 7.5 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 16 | Wide | 140 | Sn | 32 | 1.2 | 38.4 | 6.6 | 6.6 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 16 | Wide | 140 | Sn | 64 | 0.6 | 38.4 | 6.5 | 6.5 | Somatom Flash B tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Standard | 70 |  | 2 | 1 | 2 | 0.6 | 1.3 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Standard | 70 |  | 6 | 0.6 | 3.6 | 1 | 2.1 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Standard | 70 |  | 8 | 0.6 | 2.4 | 1.1 | 2.2 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Standard | 70 |  | 1 | 5 | 5 | 0.6 | 1.1 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Standard | 70 |  | 10 | 0.6 | 6 | 0.8 | 1.5 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Standard | 70 |  | 1 | 10 | 10 | 0.6 | 1.1 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Standard | 70 |  | 12 | 1.2 | 14.4 | 0.7 | 1.5 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Standard | 70 |  | 20 | 0.6 | 12 | 0.8 | 1.5 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Standard | 70 |  | 32 | 0.6 | 19.2 | 0.7 | 1.4 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Standard | 70 |  | 32 | 1.2 | 38.4 | 0.6 | 1.3 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Standard | 70 |  | 64 | 0.6 | 38.4 | 0.6 | 1.2 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Standard | 80 |  | 2 | 1 | 2 | 1.1 | 2 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Standard | 80 |  | 6 | 0.6 | 3.6 | 1.8 | 3.2 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Standard | 80 |  | 8 | 0.6 | 2.4 | 1.9 | 3.4 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Standard | 80 |  | 1 | 5 | 5 | 1 | 1.8 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Standard | 80 |  | 10 | 0.6 | 6 | 1.3 | 2.4 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Standard | 80 |  | 1 | 10 | 10 | 1 | 1.8 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Standard | 80 |  | 12 | 1.2 | 14.4 | 1.3 | 2.3 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Standard | 80 |  | 20 | 0.6 | 12 | 1.3 | 2.4 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Standard | 80 |  | 32 | 0.6 | 19.2 | 1.2 | 2.2 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Standard | 80 |  | 32 | 1.2 | 38.4 | 1.1 | 2 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Standard | 80 |  | 64 | 0.6 | 38.4 | 1.1 | 2 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Standard | 100 |  | 2 | 1 | 2 | 2.5 | 4.2 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Standard | 100 |  | 6 | 0.6 | 3.6 | 4.1 | 6.8 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Standard | 100 |  | 8 | 0.6 | 2.4 | 4.3 | 7.2 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Standard | 100 |  | 1 | 5 | 5 | 2.3 | 3.8 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Standard | 100 |  | 10 | 0.6 | 6 | 3 | 5 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Standard | 100 |  | 1 | 10 | 10 | 2.3 | 3.8 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Standard | 100 |  | 12 | 1.2 | 14.4 | 2.9 | 4.8 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Standard | 100 |  | 20 | 0.6 | 12 | 3.1 | 5.1 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Standard | 100 |  | 32 | 0.6 | 19.2 | 2.9 | 4.7 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Standard | 100 |  | 32 | 1.2 | 38.4 | 2.5 | 4.2 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Standard | 100 |  | 64 | 0.6 | 38.4 | 2.5 | 4.1 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Standard | 120 |  | 2 | 1 | 2 | 4.4 | 6.9 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Standard | 120 |  | 6 | 0.6 | 3.6 | 7.1 | 11.1 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Standard | 120 |  | 8 | 0.6 | 2.4 | 7.5 | 11.7 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Standard | 120 |  | 1 | 5 | 5 | 3.9 | 6.2 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Standard | 120 |  | 10 | 0.6 | 6 | 5.2 | 8.1 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Standard | 120 |  | 1 | 10 | 10 | 3.9 | 6.2 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Standard | 120 |  | 12 | 1.2 | 14.4 | 5 | 7.9 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Standard | 120 |  | 20 | 0.6 | 12 | 5.3 | 8.3 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Standard | 120 |  | 32 | 0.6 | 19.2 | 4.9 | 7.7 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Standard | 120 |  | 32 | 1.2 | 38.4 | 4.4 | 6.9 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Standard | 120 |  | 64 | 0.6 | 38.4 | 4.3 | 6.7 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Standard | 140 |  | 2 | 1 | 2 | 6.5 | 10 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Standard | 140 |  | 6 | 0.6 | 3.6 | 10.6 | 16.2 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Standard | 140 |  | 8 | 0.6 | 2.4 | 11.2 | 17.1 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Standard | 140 |  | 1 | 5 | 5 | 5.9 | 9 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Standard | 140 |  | 10 | 0.6 | 6 | 7.8 | 11.8 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Standard | 140 |  | 1 | 10 | 10 | 5.9 | 9 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Standard | 140 |  | 12 | 1.2 | 14.4 | 7.5 | 11.5 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Standard | 140 |  | 20 | 0.6 | 12 | 7.9 | 12.1 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Standard | 140 |  | 32 | 0.6 | 19.2 | 7.4 | 11.2 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Standard | 140 |  | 32 | 1.2 | 38.4 | 6.6 | 10 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Standard | 140 |  | 64 | 0.6 | 38.4 | 6.5 | 9.8 | Somatom Flash A tube | Siemens
Somatom Flash B tube | 17 | 1 | 32 | Wide | 70 |  | 2 | 1 | 2 | 0.7 | 1.7 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 32 | Wide | 70 |  | 6 | 0.6 | 3.6 | 1.2 | 2.7 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 32 | Wide | 70 |  | 8 | 0.6 | 2.4 | 1.2 | 2.9 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 32 | Wide | 70 |  | 1 | 5 | 5 | 0.6 | 1.5 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 32 | Wide | 70 |  | 10 | 0.6 | 6 | 0.9 | 2 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 32 | Wide | 70 |  | 1 | 10 | 10 | 0.6 | 1.5 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 32 | Wide | 70 |  | 12 | 1.2 | 14.4 | 0.8 | 2 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 32 | Wide | 70 |  | 20 | 0.6 | 12 | 0.9 | 2.1 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 32 | Wide | 70 |  | 32 | 0.6 | 19.2 | 0.8 | 1.9 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 32 | Wide | 70 |  | 32 | 1.2 | 38.4 | 0.7 | 1.7 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 32 | Wide | 70 |  | 64 | 0.6 | 38.4 | 0.7 | 1.7 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 32 | Wide | 80 |  | 2 | 1 | 2 | 1.2 | 2.6 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 32 | Wide | 80 |  | 6 | 0.6 | 3.6 | 2 | 4.3 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 32 | Wide | 80 |  | 8 | 0.6 | 2.4 | 2.1 | 4.5 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 32 | Wide | 80 |  | 1 | 5 | 5 | 1.1 | 2.4 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 32 | Wide | 80 |  | 10 | 0.6 | 6 | 1.5 | 3.1 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 32 | Wide | 80 |  | 1 | 10 | 10 | 1.1 | 2.4 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 32 | Wide | 80 |  | 12 | 1.2 | 14.4 | 1.4 | 3 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 32 | Wide | 80 |  | 20 | 0.6 | 12 | 1.5 | 3.2 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 32 | Wide | 80 |  | 32 | 0.6 | 19.2 | 1.4 | 2.9 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 32 | Wide | 80 |  | 32 | 1.2 | 38.4 | 1.2 | 2.6 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 32 | Wide | 80 |  | 64 | 0.6 | 38.4 | 1.2 | 2.6 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 32 | Wide | 100 |  | 2 | 1 | 2 | 2.7 | 5.2 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 32 | Wide | 100 |  | 6 | 0.6 | 3.6 | 4.4 | 8.4 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 32 | Wide | 100 |  | 8 | 0.6 | 2.4 | 4.7 | 8.9 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 32 | Wide | 100 |  | 1 | 5 | 5 | 2.5 | 4.6 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 32 | Wide | 100 |  | 10 | 0.6 | 6 | 3.3 | 6.1 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 32 | Wide | 100 |  | 1 | 10 | 10 | 2.5 | 4.6 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 32 | Wide | 100 |  | 12 | 1.2 | 14.4 | 3.2 | 6 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 32 | Wide | 100 |  | 20 | 0.6 | 12 | 3.3 | 6.3 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 32 | Wide | 100 |  | 32 | 0.6 | 19.2 | 3.1 | 5.8 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 32 | Wide | 100 |  | 32 | 1.2 | 38.4 | 2.8 | 5.2 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 32 | Wide | 100 |  | 64 | 0.6 | 38.4 | 2.7 | 5.1 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 32 | Wide | 120 |  | 2 | 1 | 2 | 4.7 | 8.3 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 32 | Wide | 120 |  | 6 | 0.6 | 3.6 | 7.6 | 13.4 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 32 | Wide | 120 |  | 8 | 0.6 | 2.4 | 8 | 14.2 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 32 | Wide | 120 |  | 1 | 5 | 5 | 4.2 | 7.4 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 32 | Wide | 120 |  | 10 | 0.6 | 6 | 5.6 | 9.9 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 32 | Wide | 120 |  | 1 | 10 | 10 | 4.2 | 7.4 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 32 | Wide | 120 |  | 12 | 1.2 | 14.4 | 5.4 | 9.6 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 32 | Wide | 120 |  | 20 | 0.6 | 12 | 5.7 | 10 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 32 | Wide | 120 |  | 32 | 0.6 | 19.2 | 5.3 | 9.4 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 32 | Wide | 120 |  | 32 | 1.2 | 38.4 | 4.7 | 8.3 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 32 | Wide | 120 |  | 64 | 0.6 | 38.4 | 4.6 | 8.2 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 32 | Wide | 140 |  | 2 | 1 | 2 | 7.3 | 12.4 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 32 | Wide | 140 |  | 6 | 0.6 | 3.6 | 11.8 | 20.1 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 32 | Wide | 140 |  | 8 | 0.6 | 2.4 | 12.6 | 21.2 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 32 | Wide | 140 |  | 1 | 5 | 5 | 6.6 | 11.1 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 32 | Wide | 140 |  | 10 | 0.6 | 6 | 8.7 | 14.7 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 32 | Wide | 140 |  | 1 | 10 | 10 | 6.6 | 11.1 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 32 | Wide | 140 |  | 12 | 1.2 | 14.4 | 8.4 | 14.3 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 32 | Wide | 140 |  | 20 | 0.6 | 12 | 8.9 | 15 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 32 | Wide | 140 |  | 32 | 0.6 | 19.2 | 8.3 | 14 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 32 | Wide | 140 |  | 32 | 1.2 | 38.4 | 7.3 | 12.5 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 32 | Wide | 140 |  | 64 | 0.6 | 38.4 | 7.2 | 12.2 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 32 | Wide | 140 | Sn | 2 | 1 | 2 | 2.5 | 3.7 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 32 | Wide | 140 | Sn | 6 | 0.6 | 3.6 | 4.1 | 6 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 32 | Wide | 140 | Sn | 8 | 0.6 | 2.4 | 4.3 | 6.3 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 32 | Wide | 140 | Sn | 1 | 5 | 5 | 2.3 | 3.3 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 32 | Wide | 140 | Sn | 10 | 0.6 | 6 | 3 | 4.4 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 32 | Wide | 140 | Sn | 1 | 10 | 10 | 2.3 | 3.3 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 32 | Wide | 140 | Sn | 12 | 1.2 | 14.4 | 2.9 | 4.2 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 32 | Wide | 140 | Sn | 20 | 0.6 | 12 | 3.1 | 4.5 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 32 | Wide | 140 | Sn | 32 | 0.6 | 19.2 | 2.8 | 4.1 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 32 | Wide | 140 | Sn | 32 | 1.2 | 38.4 | 2.5 | 3.7 | Somatom Flash B tube | Siemens
Somatom Flash B tube | 17 | 1 | 32 | Wide | 140 | Sn | 64 | 0.6 | 38.4 | 2.5 | 3.6 | Somatom Flash B tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Narrow | 70 |  | 2 | 1 | 2 | 0.6 | 1 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Narrow | 70 |  | 6 | 0.6 | 3.6 | 0.9 | 1.6 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Narrow | 70 |  | 8 | 0.6 | 2.4 | 1 | 1.7 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Narrow | 70 |  | 1 | 5 | 5 | 0.5 | 0.9 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Narrow | 70 |  | 10 | 0.6 | 6 | 0.7 | 1.2 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Narrow | 70 |  | 1 | 10 | 10 | 0.5 | 0.9 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Narrow | 70 |  | 12 | 1.2 | 14.4 | 0.6 | 1.1 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Narrow | 70 |  | 20 | 0.6 | 12 | 0.7 | 1.2 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Narrow | 70 |  | 32 | 0.6 | 19.2 | 0.6 | 1.1 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Narrow | 70 |  | 32 | 1.2 | 38.4 | 0.6 | 1 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Narrow | 70 |  | 64 | 0.6 | 38.4 | 0.6 | 1 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Narrow | 80 |  | 2 | 1 | 2 | 1 | 1.6 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Narrow | 80 |  | 6 | 0.6 | 3.6 | 1.6 | 2.6 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Narrow | 80 |  | 8 | 0.6 | 2.4 | 1.7 | 2.7 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Narrow | 80 |  | 1 | 5 | 5 | 0.9 | 1.4 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Narrow | 80 |  | 10 | 0.6 | 6 | 1.2 | 1.9 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Narrow | 80 |  | 1 | 10 | 10 | 0.9 | 1.4 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Narrow | 80 |  | 12 | 1.2 | 14.4 | 1.1 | 1.9 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Narrow | 80 |  | 20 | 0.6 | 12 | 1.2 | 1.9 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Narrow | 80 |  | 32 | 0.6 | 19.2 | 1.1 | 1.8 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Narrow | 80 |  | 32 | 1.2 | 38.4 | 1 | 1.6 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Narrow | 80 |  | 64 | 0.6 | 38.4 | 1 | 1.6 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Narrow | 100 |  | 2 | 1 | 2 | 2.2 | 3.4 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Narrow | 100 |  | 6 | 0.6 | 3.6 | 3.6 | 5.5 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Narrow | 100 |  | 8 | 0.6 | 2.4 | 3.8 | 5.8 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Narrow | 100 |  | 1 | 5 | 5 | 2 | 3 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Narrow | 100 |  | 10 | 0.6 | 6 | 2.6 | 4 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Narrow | 100 |  | 1 | 10 | 10 | 2 | 3 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Narrow | 100 |  | 12 | 1.2 | 14.4 | 2.6 | 3.9 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Narrow | 100 |  | 20 | 0.6 | 12 | 2.7 | 4.1 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Narrow | 100 |  | 32 | 0.6 | 19.2 | 2.5 | 3.8 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Narrow | 100 |  | 32 | 1.2 | 38.4 | 2.2 | 3.4 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Narrow | 100 |  | 64 | 0.6 | 38.4 | 2.2 | 3.3 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Narrow | 120 |  | 2 | 1 | 2 | 3.9 | 5.6 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Narrow | 120 |  | 6 | 0.6 | 3.6 | 6.3 | 9.1 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Narrow | 120 |  | 8 | 0.6 | 2.4 | 6.7 | 9.6 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Narrow | 120 |  | 1 | 5 | 5 | 3.5 | 5.1 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Narrow | 120 |  | 10 | 0.6 | 6 | 4.6 | 6.7 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Narrow | 120 |  | 1 | 10 | 10 | 3.5 | 5.1 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Narrow | 120 |  | 12 | 1.2 | 14.4 | 4.5 | 6.5 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Narrow | 120 |  | 20 | 0.6 | 12 | 4.7 | 6.8 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Narrow | 120 |  | 32 | 0.6 | 19.2 | 4.4 | 6.3 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Narrow | 120 |  | 32 | 1.2 | 38.4 | 3.9 | 5.7 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Narrow | 120 |  | 64 | 0.6 | 38.4 | 3.8 | 5.5 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Narrow | 140 |  | 2 | 1 | 2 | 6.1 | 8.7 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Narrow | 140 |  | 6 | 0.6 | 3.6 | 9.9 | 14.1 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Narrow | 140 |  | 8 | 0.6 | 2.4 | 10.5 | 14.8 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Narrow | 140 |  | 1 | 5 | 5 | 5.5 | 7.8 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Narrow | 140 |  | 10 | 0.6 | 6 | 7.3 | 10.3 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Narrow | 140 |  | 1 | 10 | 10 | 5.5 | 7.8 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Narrow | 140 |  | 12 | 1.2 | 14.4 | 7.1 | 10 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Narrow | 140 |  | 20 | 0.6 | 12 | 7.4 | 10.5 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Narrow | 140 |  | 32 | 0.6 | 19.2 | 6.9 | 9.7 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Narrow | 140 |  | 32 | 1.2 | 38.4 | 6.2 | 8.7 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Narrow | 140 |  | 64 | 0.6 | 38.4 | 6 | 8.5 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Narrow | 140 | Sn | 2 | 1 | 2 | 2.2 | 2.9 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Narrow | 140 | Sn | 6 | 0.6 | 3.6 | 3.6 | 4.6 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Narrow | 140 | Sn | 8 | 0.6 | 2.4 | 3.8 | 4.9 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Narrow | 140 | Sn | 1 | 5 | 5 | 2 | 2.6 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Narrow | 140 | Sn | 10 | 0.6 | 6 | 2.6 | 3.4 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Narrow | 140 | Sn | 1 | 10 | 10 | 2 | 2.6 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Narrow | 140 | Sn | 12 | 1.2 | 14.4 | 2.5 | 3.3 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Narrow | 140 | Sn | 20 | 0.6 | 12 | 2.7 | 3.5 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Narrow | 140 | Sn | 32 | 0.6 | 19.2 | 2.5 | 3.2 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Narrow | 140 | Sn | 32 | 1.2 | 38.4 | 2.2 | 2.9 | Somatom Flash A tube | Siemens
Somatom Flash A tube | 3 | 1 | 32 | Narrow | 140 | Sn | 64 | 0.6 | 38.4 | 2.2 | 2.8 | Somatom Flash A tube | Siemens
Definition AS | 2 | 1 | 16 | Standard | 70 |  | 2 | 1 | 2 |  |  | Definition AS | Siemens
[7 empty] | 6 | 0.6 | 3.6
[7 empty] | 8 | 0.6 | 4.8
[7 empty] | 16 | 0.6 | 4.8
[7 empty] | 1 | 5 | 5
[7 empty] | 10 | 0.6 | 6
[7 empty] | 6 | 1.2 | 7.2
[7 empty] | 1 | 10 | 10
[7 empty] | 20 | 0.6 | 12 | 1.2
[7 empty] | 12 | 1.2 | 14.4
[7 empty] | 32 | 1.2 | 38.4
[7 empty] | 64 | 0.6 | 38.4

## Pro.Pulse
16 cm phantom, Standard shaped filter (Routine head)
|  | Tube A | [5 empty] | Tube B
| Collimation | 1x5 | 1x10 | 3x3 | 3x5 | 32x0.6 | 64x0.6 | 32x0.6 | 64x0.6
| Acquisition | 1x5 | 1x10 | 3x3 | 3x5 | 32x0.6 | 64x0.6 | 32x0.6 | 64x0.6
kV
70 | Center | 4.07 | 4.01 | 4.45 | 4.04 | 4.73 | 4.18 |  | 4.4
| CTDIw | 4.39 | 4.32 | 4.8 | 4.35 | 5.1 | 4.52 |  | 4.74
| k | 1.0786 | 1.0773 | 1.0787 | 1.0767 | 1.0782 | 1.0813 |  | 1.0773
80 | Center | 6.44 | 6.34 | 7.04 | 6.39 | 7.49 | 6.61 |  | 6.92
| CTDIw | 6.81 | 6.69 | 7.43 | 6.74 | 7.9 | 7 |  | 7.32
| k | 1.0575 | 1.0552 | 1.0554 | 1.0548 | 1.0547 | 1.059 |  | 1.0578
90 | Center | 9.29 | 9.13 | 10.14 | 9.2 | 10.77 | 9.51 |  | 9.92
| CTDIw | 9.69 | 9.53 | 10.59 | 9.6 | 11.25 | 9.96 |  | 10.38
| k | 1.0431 | 1.0438 | 1.0444 | 1.0435 | 1.0446 | 1.0473 |  | 1.0464
100 | Center | 12.52 | 12.31 | 13.68 | 12.42 | 14.53 | 12.81 |  | 13.32
| CTDIw | 12.96 | 12.75 | 14.17 | 12.85 | 15.04 | 13.3 |  | 13.84
| k | 1.0351 | 1.0357 | 1.0358 | 1.0346 | 1.0351 | 1.0383 |  | 1.039
110 | Center | 15.99 | 15.79 | 17.54 | 15.9 | 18.64 | 16.42 |  | 17.02
| CTDIw | 16.57 | 16.3 | 18.11 | 16.4 | 19.21 | 16.98 |  | 17.62
| k | 1.0363 | 1.0323 | 1.0325 | 1.0314 | 1.0306 | 1.0341 |  | 1.0353
120 | Center | 19.92 | 19.6 | 21.77 | 19.71 | 23.08 | 20.34 |  | 21.08
| CTDIw | 20.5 | 20.13 | 22.37 | 20.26 | 23.73 | 20.96 |  | 21.72
| k | 1.0291 | 1.027 | 1.0276 | 1.0279 | 1.0282 | 1.0305 |  | 1.0304
130 | Center | 24.1 | 23.69 | 26.32 | 23.81 | 27.89 | 24.52 |  | 25.36
| CTDIw | 24.72 | 24.26 | 26.95 | 24.41 | 28.57 | 25.22 |  | 26.1
| k | 1.0257 | 1.0241 | 1.0239 | 1.0252 | 1.0244 | 1.0285 |  | 1.0292
140 | Center | 28.32 | 27.94 | 31.04 | 28.13 | 23.89 | 29 |  | 30
| CTDIw | 29.08 | 28.6 | 31.78 | 28.77 | 33.69 | 29.75 |  | 30.77
| k | 1.0268 | 1.0236 | 1.0238 | 1.0228 | 1.4102 | 1.0259 |  | 1.0257
100Sn | Center |  |  | 2.11 | 1.92 |  | 2 | 2.41 | 2.08
| CTDIw |  |  | 2.13 | 1.93 |  | 2.02 | 2.42 | 2.11
| k |  |  | 1.0095 | 1.0052 |  | 1.01 | 1.0041 | 1.0144
110Sn | Center |  |  | 3.26 | 2.97 |  | 3.1 | 3.71 | 3.22
| CTDIw |  |  | 3.3 | 3 |  | 3.14 | 3.7 | 3.26
| k |  |  | 1.0123 | 1.0101 |  | 1.0129 | 0.9973 | 1.0124
120Sn | Center |  |  | 4.72 | 4.29 |  | 4.47 | 5.35 | 4.63
| CTDIw |  |  | 4.77 | 4.33 |  | 4.53 | 5.41 | 4.69
| k |  |  | 1.0106 | 1.0093 |  | 1.0134 | 1.0112 | 1.013
130Sn | Center |  |  | 6.47 | 5.87 |  | 6.11 | 7.33 | 6.33
| CTDIw |  |  | 6.55 | 5.95 |  | 6.21 | 7.41 | 6.43
| k |  |  | 1.0124 | 1.0136 |  | 1.0164 | 1.0109 | 1.0158
140Sn | Center |  |  | 8.5 | 7.73 |  | 8.04 | 9.62 | 8.32
| CTDIw |  |  | 8.63 | 7.83 |  | 8.19 | 9.75 | 8.46
| k |  |  | 1.0153 | 1.0129 |  | 1.0187 | 1.0135 | 1.0168
16 cm phantom, Narrow shaped filter
|  | Tube A | [5 empty] | Tube B
| Collimation | 1x5 | 1x10 | 3x3 | 3x5 | 32x0.6 | 64x0.6 | 32x0.6 | 64x0.6
| Acquisition | 1x5 | 1x10 | 3x3 | 3x5 | 32x0.6 | 64x0.6 | 32x0.6 | 64x0.6
kV
70 | Center | 3.64 | 3.61 | 4.01 | 3.63 | 4.24 | 3.74 |  | 3.93
| CTDIw | 3.68 | 3.63 | 4.03 | 3.65 | 4.27 | 3.77 |  | 3.97
| k | 1.011 | 1.0055 | 1.005 | 1.0055 | 1.0071 | 1.008 |  | 1.0102
80 | Center | 5.83 | 5.76 | 6.39 | 5.8 | 6.79 | 5.98 |  | 6.25
| CTDIw | 5.8 | 5.72 | 6.36 | 5.76 | 6.74 | 5.95 |  | 6.23
| k | 0.99485 | 0.99306 | 0.99531 | 0.9931 | 0.99264 | 0.99498 |  | 0.9968
90 | Center | 8.5 | 8.35 | 9.28 | 8.41 | 9.84 | 8.66 |  | 9.04
| CTDIw | 8.39 | 8.25 | 9.17 | 8.3 | 9.72 | 8.58 |  | 8.95
| k | 0.98706 | 0.98802 | 0.98815 | 0.98692 | 0.9878 | 0.99076 |  | 0.99004
100 | Center | 11.51 | 11.35 | 12.61 | 11.4 | 13.35 | 11.73 |  | 12.22
| CTDIw | 11.33 | 11.16 | 12.4 | 11.2 | 13.12 | 11.57 |  | 12.05
| k | 0.98436 | 0.98326 | 0.98335 | 0.98246 | 0.98277 | 0.98636 |  | 0.98609
110 | Center | 16.9 | 16.65 | 16.28 | 14.72 | 17.19 | 15.1 |  | 15.7
| CTDIw | 16.62 | 14.38 | 15.98 | 14.44 | 16.88 | 14.88 |  | 15.46
| k | 0.98343 | 0.86366 | 0.98157 | 0.98098 | 0.98197 | 0.98543 |  | 0.98471
120 | Center | 18.54 | 18.2 | 20.22 | 18.27 | 21.41 | 18.8 |  | 19.52
| CTDIw | 18.19 | 17.86 | 19.84 | 17.92 | 20.99 | 18.49 |  | 19.18
| k | 0.98112 | 0.98132 | 0.98121 | 0.98084 | 0.98038 | 0.98351 |  | 0.98258
130 | Center | 22.44 | 22.05 | 24.49 | 22.15 | 25.92 | 22.75 |  | 23.57
| CTDIw | 22.01 | 21.61 | 24.01 | 21.7 | 25.41 | 22.37 |  | 23.18
| k | 0.98084 | 0.98005 | 0.9804 | 0.97968 | 0.98032 | 0.9833 |  | 0.98345
140 | Center | 26.59 | 26.16 | 29.06 | 26.26 | 30.73 | 26.99 |  | 27.92
| CTDIw | 26.09 | 25.66 | 28.53 | 25.74 | 30.1 | 26.53 |  | 27.46
| k | 0.9812 | 0.98089 | 0.98176 | 0.9802 | 0.9795 | 0.98296 |  | 0.98352
100Sn | Center |  |  | 1.99 | 1.8 |  | 1.88 | 2.26 | 1.96
| CTDIw |  |  | 1.94 | 1.76 |  | 1.84 | 2.2 | 1.91
| k |  |  | 0.97487 | 0.97778 |  | 0.97872 | 0.97345 | 0.97449
110Sn | Center |  |  | 3.1 | 2.81 |  | 2.92 | 3.51 | 3.04
| CTDIw |  |  | 3.03 | 2.74 |  | 2.86 | 3.43 | 2.97
| k |  |  | 0.97742 | 0.97509 |  | 0.97945 | 0.97721 | 0.97697
120Sn | Center |  |  | 4.48 | 4.06 |  | 4.23 | 5.08 | 4.39
| CTDIw |  |  | 4.39 | 3.98 |  | 4.15 | 4.97 | 4.31
| k |  |  | 0.97991 | 0.9803 |  | 0.98109 | 0.97835 | 0.98178
130Sn | Center |  |  | 6.15 | 5.58 |  | 5.8 | 6.96 | 6.01
| CTDIw |  |  | 6.04 | 5.48 |  | 5.71 | 6.83 | 5.92
| k |  |  | 0.98211 | 0.98208 |  | 0.98448 | 0.98132 | 0.98502
140Sn | Center |  |  | 8.13 | 7.36 |  | 7.64 | 9.17 | 7.92
| CTDIw |  |  | 8.01 | 7.25 |  | 7.55 | 9.02 | 7.81
| k |  |  | 0.98524 | 0.98505 |  | 0.98822 | 0.98364 | 0.98611
32 cm phantom, Standard shaped filter (Routine body)
|  | Tube A | [5 empty] | Tube B
| Collimation | 1x5 | 1x10 | 3x3 | 3x5 | 32x0.6 | 64x0.6 | 32x0.6 | 64x0.6
| Acquisition | 1x5 | 1x10 | 3x3 | 3x5 | 32x0.6 | 64x0.6 | 32x0.6 | 64x0.6
kV
70 | Center | 0.91 | 0.9 | 1 | 0.9 | 1.05 | 0.93 |  | 0.98
| CTDIw | 1.83 | 1.79 | 1.98 | 1.79 | 2.1 | 1.87 |  | 1.96
| k | 2.011 | 1.9889 | 1.98 | 1.9889 | 2 | 2.0108 |  | 2
80 | Center | 1.59 | 1.57 | 1.74 | 1.57 | 1.84 | 1.63 |  | 1.7
| CTDIw | 2.96 | 2.89 | 3.21 | 2.9 | 3.39 | 3.01 |  | 3.14
| k | 1.8616 | 1.8408 | 1.8448 | 1.8471 | 1.8424 | 1.8466 |  | 1.8471
90 | Center | 2.46 | 2.42 | 2.68 | 2.43 | 2.85 | 2.51 |  | 2.96
| CTDIw | 4.35 | 4.25 | 4.72 | 4.26 | 4.98 | 4.42 |  | 4.59
| k | 1.7683 | 1.7562 | 1.7612 | 1.7531 | 1.7474 | 1.761 |  | 1.5507
100 | Center | 3.49 | 3.43 | 3.82 | 3.45 | 4.03 | 3.56 |  | 3.7
| CTDIw | 5.96 | 5.81 | 6.46 | 5.84 | 6.82 | 6.06 |  | 6.28
| k | 1.7077 | 1.6939 | 1.6911 | 1.6928 | 1.6923 | 1.7022 |  | 1.6973
110 | Center | 4.69 | 4.58 | 5.09 | 4.62 | 5.4 | 4.77 |  | 4.94
| CTDIw | 7.79 | 7.59 | 8.44 | 7.62 | 8.91 | 7.9 |  | 8.16
| k | 1.661 | 1.6572 | 1.6582 | 1.6494 | 1.65 | 1.6562 |  | 1.6518
120 | Center | 5.98 | 5.87 | 6.52 | 5.9 | 6.92 | 6.11 |  | 6.31
| CTDIw | 9.8 | 9.54 | 10.6 | 9.57 | 11.19 | 9.92 |  | 10.24
| k | 1.6388 | 1.6252 | 1.6258 | 1.622 | 1.6171 | 1.6236 |  | 1.6228
130 | Center | 7.41 | 7.27 | 8.08 | 7.33 | 8.56 | 7.57 |  | 7.81
| CTDIw | 12 | 11.68 | 12.97 | 1.71 | 13.68 | 12.12 |  | 12.49
| k | 1.6194 | 1.6066 | 1.6052 | 0.23329 | 1.5981 | 1.6011 |  | 1.5992
140 | Center | 8.95 | 8.77 | 9.75 | 8.83 | 10.34 | 9.13 |  | 9.44
| CTDIw | 14.39 | 13.96 | 15.51 | 14 | 16.35 | 14.47 |  | 14.91
| k | 1.6078 | 1.5918 | 1.5908 | 1.5855 | 1.5812 | 1.5849 |  | 1.5794
100Sn | Center |  |  | 0.73 | 0.66 |  | 0.69 | 0.83 | 0.71
| CTDIw |  |  | 1.1 | 1 |  | 1.04 | 1.24 | 1.08
| k |  |  | 1.5068 | 1.5152 |  | 1.5072 | 1.494 | 1.5211
110Sn | Center |  |  | 1.16 | 1.06 |  | 1.1 | 1.32 | 1.14
| CTDIw |  |  | 1.74 | 1.57 |  | 1.65 | 1.95 | 1.7
| k |  |  | 1.5 | 1.4811 |  | 1.5 | 1.4773 | 1.4912
120Sn | Center |  |  | 1.72 | 1.56 |  | 1.63 | 1.94 | 1.68
| CTDIw |  |  | 2.54 | 2.3 |  | 2.41 | 2.85 | 2.48
| k |  |  | 1.4767 | 1.4744 |  | 1.4785 | 1.4691 | 1.4762
130Sn | Center |  |  | 2.39 | 2.17 |  | 2.27 | 2.7 | 2.34
| CTDIw |  |  | 3.53 | 3.19 |  | 3.34 | 3.94 | 3.43
| k |  |  | 1.477 | 1.47 |  | 1.4714 | 1.4593 | 1.4658
140Sn | Center |  |  | 3.18 | 2.89 |  | 3.02 | 3.6 | 3.12
| CTDIw |  |  | 4.7 | 4.25 |  | 4.44 | 5.24 | 4.56
| k |  |  | 1.478 | 1.4706 |  | 1.4702 | 1.4556 | 1.4615
32 cm phantom, Narrow shaped filter
|  | Tube A | [5 empty] | Tube B
| Collimation | 1x5 | 1x10 | 3x3 | 3x5 | 32x0.6 | 64x0.6 | 32x0.6 | 64x0.6
| Acquisition | 1x5 | 1x10 | 3x3 | 3x5 | 32x0.6 | 64x0.6 | 32x0.6 | 64x0.6
kV
70 | Center | 0.79 | 0.79 | 0.86 | 0.79 | 0.91 | 0.8 |  | 0.85
| CTDIw | 1.38 | 1.36 | 1.51 | 1.36 | 1.59 | 1.41 |  | 1.48
| k | 1.7468 | 1.7215 | 1.7558 | 1.7215 | 1.7473 | 1.7625 |  | 1.7412
80 | Center | 1.39 | 1.37 | 1.52 | 1.37 | 1.61 | 1.42 |  | 1.48
| CTDIw | 2.26 | 2.22 | 2.47 | 2.23 | 2.61 | 2.31 |  | 2.41
| k | 1.6259 | 1.6204 | 1.625 | 1.6277 | 1.6211 | 1.6268 |  | 1.6284
90 | Center | 2.17 | 2.13 | 2.36 | 2.13 | 2.5 | 2.2 |  | 2.3
| CTDIw | 3.37 | 3.3 | 3.67 | 3.31 | 3.87 | 3.43 |  | 3.57
| k | 1.553 | 1.5493 | 1.5551 | 1.554 | 1.548 | 1.5591 |  | 1.5522
100 | Center | 3.09 | 3.04 | 3.38 | 3.05 | 3.57 | 3.13 |  | 3.27
| CTDIw | 4.66 | 4.57 | 5.08 | 4.58 | 5.36 | 4.73 |  | 4.92
| k | 1.5081 | 1.5033 | 1.503 | 1.5016 | 1.5014 | 1.5112 |  | 1.5046
110 | Center | 4.15 | 4.08 | 4.54 | 4.1 | 4.79 | 4.22 |  | 4.38
| CTDIw | 6.14 | 6.01 | 6.68 | 6.03 | 7.04 | 6.22 |  | 6.45
| k | 1.4795 | 1.473 | 1.4714 | 1.4707 | 1.4697 | 1.4739 |  | 1.4726
120 | Center | 5.35 | 5.25 | 5.83 | 5.26 | 6.17 | 5.42 |  | 5.62
| CTDIw | 7.78 | 7.6 | 8.45 | 7.62 | 8.91 | 7.87 |  | 8.14
| k | 1.4542 | 1.4476 | 1.4494 | 1.4487 | 1.4441 | 1.452 |  | 1.4484
130 | Center | 6.67 | 6.54 | 7.26 | 6.55 | 7.68 | 6.74 |  | 6.98
| CTDIw | 9.57 | 9.36 | 10.4 | 9.37 | 10.95 | 9.66 |  | 9.98
| k | 1.4348 | 1.4312 | 1.4325 | 1.4305 | 1.4258 | 1.4332 |  | 1.4298
140 | Center | 8.03 | 7.9 | 8.78 | 7.93 | 9.3 | 8.18 |  | 8.46
| CTDIw | 11.5 | 11.25 | 12.5 | 11.25 | 13.15 | 11.61 |  | 11.98
| k | 1.4321 | 1.4241 | 1.4237 | 1.4187 | 1.414 | 1.4193 |  | 1.4161
100Sn | Center |  |  | 0.66 | 0.59 |  | 0.62 | 0.74 | 0.64
| CTDIw |  |  | 0.9 | 0.81 |  | 0.85 | 1.01 | 0.88
| k |  |  | 1.3636 | 1.3729 |  | 1.371 | 1.3649 | 1.375
110Sn | Center |  |  | 1.05 | 0.95 |  | 0.99 | 1.19 | 1.03
| CTDIw |  |  | 1.42 | 1.29 |  | 1.34 | 1.6 | 1.39
| k |  |  | 1.3524 | 1.3579 |  | 1.3535 | 1.3445 | 1.3495
120Sn | Center |  |  | 1.56 | 1.41 |  | 1.47 | 1.76 | 1.53
| CTDIw |  |  | 2.1 | 1.9 |  | 1.98 | 2.35 | 2.04
| k |  |  | 1.3462 | 1.3475 |  | 1.3469 | 1.3352 | 2.13
130Sn | Center |  |  | 2.18 | 1.98 |  | 2.06 | 2.46 | 2.13
| CTDIw |  |  | 2.92 | 2.64 |  | 2.76 | 3.27 | 2.84
| k |  |  | 1.3394 | 1.3333 |  | 1.3398 | 1.3293 | 3.79
140Sn | Center |  |  | 2.92 | 2.64 |  | 2.75 | 3.3 | 2.85
| CTDIw |  |  | 3.91 | 3.53 |  | 6.68 | 4.37 | 3.79
| k |  |  | 1.339 | 1.3371 |  | 2.4291 | 1.3242 | 1.3298

## Neurologica BodyTom
16 cm phantom (Adult head)
| Collimation | 16x0.6 | 7.9 | 9.46
| Total | 9.6 | 16.23 | 18.1
| Acquisition | 16x0.6 | 24.62 | 27.07
| Shaped filter |  | 35 | 37.74
kV |  |  | 1.49 | 3.15
80 | Center | 7.9 | 3.08 | 5.85
| CTDIw | 9.46 | 5.53 | 9.23
| k | 1.1975 | 8.26 | 13.13
100 | Center | 16.23
| CTDIw | 18.1
| k | 1.1152
120 | Center | 24.62
| CTDIw | 27.07
| k | 1.0995
140 | Center | 35
| CTDIw | 37.74
| k | 1.0783
32 cm phantom (Adult body)
| Collimation | 64x0.6
| Total | 38.4
| Acquisition | 64x0.6
Shaped filter
kV
80 | Center | 1.49
| CTDIw | 3.15
| k | 2.1141
100 | Center | 3.08
| CTDIw | 5.85
| k | 1.8994
120 | Center | 5.53
| CTDIw | 9.23
| k | 1.6691
140 | Center | 8.26
| CTDIw | 13.13
| k | 1.5896

## Symbia Pro.specta Q3
16 cm phantom (Adult head) | [6 empty] | 5.33 | 5.92
| Collimation | 1x5 | 1x10 | 3x3 | 16x0.7 | 16x0.7 (Sn) | 5.36 | 5.93
| Total | 5 | 10 | 9 | 11.2 | 11.2 | 5.96 | 6.59
| Acquisition | 1x5 | 1x10 | 3x3 | 16x0.7 | 16x0.7 (Sn) | 6.38 | 7.08
Shaped filter
kV | [6 empty] | 13.33 | 14.3
80 | Center | 5.33 | 5.36 | 5.96 | 6.38 |  | 13.39 | 14.32
| CTDIw | 5.92 | 5.93 | 6.59 | 7.08 |  | 14.87 | 15.92
| k | 1.1107 | 1.1063 | 1.1057 | 1.1097 |  | 15.98 | 17.1
110 | Center | 13.33 | 13.39 | 14.87 | 15.98 | 2.99 | 2.99 | 3.08
| CTDIw | 14.3 | 14.32 | 15.92 | 17.1 | 3.08 | 20.27 | 21.52
| k | 1.0728 | 1.0695 | 1.0706 | 1.0701 | 1.0301 | 20.27 | 21.52
130 | Center | 20.27 | 20.27 | 22.52 | 24.24 | 5.86 | 22.52 | 23.91
| CTDIw | 21.52 | 21.52 | 23.91 | 25.7 | 6.05 | 24.24 | 25.7
| k | 1.0617 | 1.0617 | 1.0617 | 1.0602 | 1.0324 | 5.86 | 6.05
[7 empty] | 1.36 | 2.64
32 cm phantom (Adult body) | [6 empty] | 1.36 | 2.64
| Collimation | 1x5 | 1x10 | 3x3 | 16x0.7 | 16x0.7 (Sn) | 1.51 | 2.94
| Total | 5 | 10 | 9 | 11.2 | 11.2 | 1.62 | 3.15
| Acquisition | 1x5 | 1x10 | 3x3 | 16x0.7 | 16x0.7 (Sn)
| Shaped filter | [5 empty] | 3.98 | 6.82
kV | [6 empty] | 3.97 | 6.83
80 | Center | 1.36 | 1.36 | 1.51 | 1.62 |  | 4.41 | 7.59
| CTDIw | 2.64 | 2.64 | 2.94 | 3.15 |  | 4.75 | 8.13
| k | 1.9412 | 1.9412 | 1.947 | 1.9444 |  | 1.09 | 1.65
110 | Center | 3.98 | 3.97 | 4.41 | 4.75 | 1.09 | 6.4 | 10.57
| CTDIw | 6.82 | 6.83 | 7.59 | 8.13 | 1.65 | 6.39 | 10.57
| k | 1.7136 | 1.7204 | 1.7211 | 1.7116 | 1.5138 | 7.1 | 11.74
130 | Center | 6.4 | 6.39 | 7.1 | 7.63 | 2.23 | 7.63 | 12.57
| CTDIw | 10.57 | 10.57 | 11.74 | 12.57 | 3.32 | 2.23 | 3.32
| k | 1.6516 | 1.6541 | 1.6535 | 1.6474 | 1.4888

## Intevo Bold
16 cm phantom (Routine head)
| Collimation | 4x0.6 | 12x0.6 | 16x0.6 | 1x10 | 12x1.2 | 2x8 | 16x1.2
| Total | 2.4 | 7.2 | 9.6 | 10 | 14.4 | 16 | 19.2
| Acquisition | 4x0.6 | 12x0.6 | 16x0.6 | 1x10 | 12x1.2 | 2x8 | 16x1.2
Shaped filter
kV
80 | Center | 7.6 | 7.7 | 6.64 | 5.55 | 6.04 | 5.43 | 6.13
| CTDIw | 8.12 | 8.2 | 7.19 | 5.9 | 6.45 | 5.81 | 6.56
| k | 1.0684 | 1.0649 | 1.0828 | 1.0631 | 1.0679 | 1.07 | 1.0701
110 | Center | 18 | 17.63 | 15.57 | 12.69 | 13.88 | 12.49 | 14.14
| CTDIw | 18.65 | 18.31 | 16.18 | 13.18 | 14.43 | 12.99 | 14.7
| k | 1.0361 | 1.0386 | 1.0392 | 1.0386 | 1.0396 | 1.04 | 1.0396
130 | Center | 25.78 | 25.93 | 22.91 | 18.67 | 20.42 | 18.38 | 20.81
| CTDIw | 26.52 | 26.72 | 23.62 | 19.24 | 21.05 | 18.94 | 21.45
| k | 1.0287 | 1.0305 | 1.031 | 1.0305 | 1.0309 | 1.0305 | 1.0308
32 cm phantom (Adult body)
| Collimation | 4x0.6 | 12x0.6 | 16x0.6 | 1x10 | 12x1.2 | 2x8 | 16x1.2
| Total | 2.4 | 7.2 | 9.6 | 10 | 14.4 | 16 | 19.2
| Acquisition | 4x0.6 | 12x0.6 | 16x0.6 | 1x10 | 12x1.2 | 2x8 | 16x1.2
Shaped filter
kV
80 | Center | 1.88 | 1.9 | 1.68 | 1.37 | 1.5 | 1.35 | 1.52
| CTDIw | 3.49 | 3.54 | 3.12 | 2.55 | 2.78 | 2.5 | 2.9
| k | 1.8564 | 1.8632 | 1.8571 | 1.8613 | 1.8533 | 1.8519 | 1.9079
110 | Center | 5.24 | 5.12 | 4.52 | 3.68 | 4.05 | 3.64 | 4.12
| CTDIw | 8.63 | 8.5 | 7.51 | 6.12 | 6.69 | 6.02 | 6.82
| k | 1.6469 | 1.6602 | 1.6615 | 1.663 | 1.6519 | 1.6538 | 1.6553
130 | Center | 7.95 | 8.02 | 7.09 | 5.78 | 6.31 | 5.68 | 6.41
| CTDIw | 12.63 | 12.79 | 11.29 | 9.21 | 10.04 | 9.04 | 10.21
| k | 1.5887 | 1.5948 | 1.5924 | 1.5934 | 1.5911 | 1.5915 | 1.5928

## Biograph Vision
16 cm phantom, Standard shaped filter (Routine head)
| Collimation | 2x1 | 6x0.6 | 8x0.6 | 8x0.6 | 1x5 | 10x0.6 | 6x1.2 | 1x10 | 20x0.6 | 12x1.2 | 64x0.6 | 64x0.6
| Total | 2 | 3.6 | 4.8 | 4.8 | 5 | 6 | 7.2 | 10 | 12 | 14.4 | 19.2 | 38.4
| Acquisition | 2x1 | 6x0.6 | 8x0.6 | 16x0.6 | 1x5 | 20x0.6 | 6x1.2 | 1x10 | 40x0.6 | 12x1.2 | 32x1.2 | 128x0.6
|  |  | Topogram | UHR | UHR, z-sharp |  | z-sharp |  |  | z-sharp |  |  | z-sharp
kV
70 | Center | 2.6 | 4.22 | 3.88 | 4.58 | 2.34 | 3.17 | 3.3 | 2.34 | 3.23 | 3 | 3.01 | 2.87
| CTDIw | 2.74 | 4.43 | 4.07 | 4.78 | 2.46 | 3.31 | 3.35 | 2.46 | 3.38 | 3.15 | 3.15 | 3.01
| k | 1.0538 | 1.0498 | 1.049 | 1.0437 | 1.0513 | 1.0442 | 1.0152 | 1.0513 | 1.0464 | 1.05 | 1.0465 | 1.0488
80 | Center | 4.18 | 6.76 | 6.22 | 7.24 | 3.75 | 5.02 | 5.39 | 3.75 | 5.12 | 4.81 | 4.77 | 4.6
| CTDIw | 4.36 | 7.06 | 6.49 | 7.46 | 3.91 | 5.17 | 5.62 | 3.91 | 5.27 | 5.02 | 4.91 | 4.8
| k | 1.0431 | 1.0444 | 1.0434 | 1.0304 | 1.0427 | 1.0299 | 1.0427 | 1.0427 | 1.0293 | 1.0437 | 1.0294 | 1.0435
100 | Center | 8.44 | 13.67 | 12.56 | 14.69 | 7.57 | 10.19 | 10.89 | 7.57 | 10.38 | 9.71 | 8.48 | 8.57
| CTDIw | 8.65 | 14.01 | 12.88 | 14.93 | 7.76 | 10.35 | 11.16 | 7.76 | 10.55 | 9.96 | 8.69 | 8.71
| k | 1.0249 | 1.0249 | 1.0255 | 1.0163 | 1.0251 | 1.0157 | 1.0248 | 1.0251 | 1.0164 | 1.0257 | 1.0248 | 1.0163
120 | Center | 13.85 | 22.42 | 20.6 | 24.33 | 12.42 | 16.87 | 17.86 | 12.42 | 17.19 | 15.93 | 13.91 | 14.19
| CTDIw | 14.11 | 22.85 | 21 | 24.63 | 12.66 | 17.08 | 18.2 | 12.66 | 17.41 | 16.24 | 14.17 | 14.37
| k | 1.0188 | 1.0192 | 1.0194 | 1.0123 | 1.0193 | 1.0124 | 1.019 | 1.0193 | 1.0128 | 1.0195 | 1.0187 | 1.0127
140 | Center | 20.24 | 32.77 | 30.12 | 36.16 | 18.16 | 25.07 | 26.11 | 18.16 | 25.55 | 23.3 | 20.33 | 21.09
| CTDIw | 20.6 | 33.35 | 30.65 | 36.6 | 18.48 | 25.4 | 26.56 | 18.48 | 25.89 | 23.7 | 20.69 | 21.37
| k | 1.0178 | 1.0177 | 1.0176 | 1.0122 | 1.0176 | 1.0132 | 1.0172 | 1.0176 | 1.0133 | 1.0172 | 1.0177 | 1.0133
16 cm phantom, Narrow shaped filter (Pediatric body)
| Collimation | 2x1 | 6x0.6 | 8x0.6 | 8x0.6 | 1x5 | 10x0.6 | 6x1.2 | 1x10 | 20x0.6 | 12x1.2 | 32x0.6 | 64x0.6
| Total | 2 | 3.6 | 4.8 | 4.8 | 5 | 6 | 7.2 | 10 | 12 | 14.4 | 38.4 | 38.4
| Acquisition | 2x1 | 6x0.6 | 8x0.6 | 16x0.6 | 1x5 | 20x0.6 | 6x1.2 | 1x10 | 40x0.6 | 12x1.2 | 32x1.2 | 28.x.6
| Shaped filter |  |  |  | z-sharp |  | z-sharp |  |  | z-sharp |  |  | z-sharp
kV
70 | Center | 2.36 | 3.83 | 3.52 | 4.12 | 2.12 | 2.86 | 3.05 | 2.12 | 2.91 | 2.72 | 2.38 | 2.4
| CTDIw | 2.36 | 3.81 | 3.51 | 4.07 | 2.11 | 2.82 | 3.04 | 2.11 | 2.88 | 2.71 | 2.37 | 2.37
| k | 1 | 0.99478 | 0.99716 | 0.98786 | 0.99528 | 0.98601 | 0.99672 | 0.99528 | 0.98969 | 0.99632 | 0.9958 | 0.9875
80 | Center | 3.81 | 6.17 | 5.67 | 6.58 | 3.42 | 4.56 | 4.92 | 3.42 | 4.65 | 4.39 | 3.83 | 3.84
| CTDIw | 3.78 | 6.12 | 5.63 | 6.44 | 3.39 | 4.47 | 4.88 | 3.39 | 4.55 | 4.35 | 3.8 | 3.76
| k | 0.99213 | 0.9919 | 0.99295 | 0.97872 | 0.99123 | 0.98026 | 0.99187 | 0.99123 | 0.97849 | 0.99089 | 0.99217 | 0.97917
100 | Center | 7.81 | 12.65 | 11.63 | 13.53 | 7.01 | 9.38 | 10.08 | 7.01 | 9.56 | 8.99 | 7.85 | 7.89
| CTDIw | 7.69 | 12.45 | 11.44 | 13.14 | 6.9 | 9.11 | 9.92 | 6.9 | 9.29 | 8.85 | 7.72 | 7.67
| k | 0.98464 | 0.98419 | 0.98366 | 0.97118 | 0.98431 | 0.97122 | 0.98413 | 0.98431 | 0.97176 | 0.98443 | 0.98344 | 0.97212
120 | Center | 12.93 | 20.93 | 19.23 | 22.53 | 11.6 | 15.62 | 16.67 | 11.6 | 15.92 | 14.87 | 12.98 | 13.14
| CTDIw | 12.68 | 20.52 | 18.86 | 21.91 | 11.37 | 15.19 | 16.35 | 11.37 | 15.48 | 14.59 | 12.73 | 12.78
| k | 0.98067 | 0.98041 | 0.98076 | 0.97248 | 0.98017 | 0.97247 | 0.9808 | 0.98017 | 0.97236 | 0.98117 | 0.98074 | 0.9726
140 | Center | 18.83 | 30.49 | 28.02 | 33.71 | 16.89 | 23.38 | 24.29 | 16.89 | 23.82 | 21.67 | 18.91 | 19.67
| CTDIw | 18.5 | 29.95 | 27.53 | 32.85 | 16.6 | 22.77 | 23.86 | 16.6 | 23.21 | 21.29 | 18.58 | 19.16
| k | 0.98247 | 0.98229 | 0.98251 | 0.97449 | 0.98283 | 0.97391 | 0.9823 | 0.98283 | 0.97439 | 0.98246 | 0.98255 | 0.97407
32 cm phantom, Standard shaped filter (Adult Body)
| Collimation | 2x1 | 6x0.6 | 8x0.6 | 8x0.6 | 1x5 | 10x0.6 | 6x1.2 | 1x10 | 20x0.6 | 12x1.2 | 32x0.6 | 64x0.6
| Total | 2 | 3.6 | 4.8 | 4.8 | 5 | 6 | 7.2 | 10 | 12 | 14.4 | 38.4 | 38.4
| Acquisition | 2x1 | 6x0.6 | 8x0.6 | 16x0.6 | 1x5 | 20x0.6 | 6x1.2 | 1x10 | 40x0.6 | 12x1.2 | 32x1.2 | 28.x.6
| Shaped filter |  |  |  | z-sharp |  | z-sharp |  |  | z-sharp |  |  | z-sharp
kV
70 | Center | 0.61 | 0.99 | 0.91 | 1.07 | 0.55 | 0.74 | 0.79 | 0.55 | 0.75 | 0.7 | 0.61 | 0.62
| CTDIw | 1.13 | 1.83 | 1.68 | 1.97 | 1.01 | 1.37 | 1.46 | 1.01 | 1.4 | 1.3 | 1.14 | 1.15
| k | 1.8525 | 1.8485 | 1.8462 | 1.8411 | 1.8364 | 1.8514 | 1.8481 | 1.8364 | 1.8667 | 1.8571 | 1.8689 | 1.8548
80 | Center | 1.07 | 1.73 | 1.59 | 1.83 | 0.96 | 1.27 | 1.38 | 0.96 | 1.29 | 1.23 | 1.08 | 1.07
| CTDIw | 1.86 | 3.01 | 2.77 | 3.18 | 1.67 | 2.21 | 2.4 | 1.67 | 2.25 | 2.14 | 1.87 | 1.86
| k | 1.7383 | 1.7399 | 1.7421 | 1.7377 | 1.7396 | 1.7402 | 1.7391 | 1.7396 | 1.7442 | 1.7398 | 1.7315 | 1.7383
100 | Center | 2.43 | 3.94 | 3.62 | 4.19 | 2.18 | 2.91 | 3.14 | 2.18 | 2.96 | 2.8 | 2.44 | 2.45
| CTDIw | 3.9 | 6.32 | 5.81 | 6.75 | 3.5 | 4.68 | 5.03 | 3.5 | 4.77 | 4.49 | 3.92 | 3.94
| k | 1.6049 | 1.6041 | 1.605 | 1.611 | 1.6055 | 1.6082 | 1.6019 | 1.6055 | 1.6115 | 1.6036 | 1.6066 | 1.6082
120 | Center | 4.28 | 6.93 | 6.37 | 7.5 | 3.84 | 5 | 5.52 | 3.84 | 5.3 | 4.93 | 4.3 | 4.37
| CTDIw | 6.61 | 10.7 | 9.84 | 11.55 | 5.93 | 8.01 | 8.53 | 5.93 | 8.16 | 7.61 | 6.64 | 6.74
| k | 1.5444 | 1.544 | 1.5447 | 1.54 | 1.5443 | 1.602 | 1.5453 | 1.5443 | 1.5396 | 1.5436 | 1.5442 | 1.5423
140 | Center | 6.57 | 10.64 | 9.78 | 11.59 | 5.89 | 8.03 | 8.47 | 5.89 | 8.19 | 7.56 | 6.6 | 6.76
| CTDIw | 9.94 | 16.09 | 14.79 | 17.62 | 8.91 | 12.22 | 12.81 | 8.91 | 12.45 | 11.43 | 9.98 | 10.28
| k | 1.5129 | 1.5122 | 1.5123 | 1.5203 | 1.5127 | 1.5218 | 1.5124 | 1.5127 | 1.5201 | 1.5119 | 1.5121 | 1.5207
32 cm phantom, Narrow shaped filter (Cardiac, Ped Body)
| Collimation | 2x1 | 6x0.6 | 8x0.6 | 8x0.6 | 1x5 | 10x0.6 | 6x1.2 | 1x10 | 20x0.6 | 12x1.2 | 32x0.6 | 64x0.6
| Total | 2 | 3.6 | 4.8 | 4.8 | 5 | 6 | 7.2 | 10 | 12 | 14.4 | 38.4 | 38.4
| Acquisition | 2x1 | 6x0.6 | 8x0.6 | 16x0.6 | 1x5 | 20x0.6 | 6x1.2 | 1x10 | 40x0.6 | 12x1.2 | 32x1.2 | 28.x.6
| Shaped filter |  |  |  | z-sharp |  | z-sharp |  |  | z-sharp |  |  | z-sharp
kV
70 | Center | 0.54 | 0.87 | 0.8 | 0.93 | 0.48 | 0.65 | 0.69 | 0.48 | 0.66 | 0.62 | 0.54 | 0.55
| CTDIw | 0.9 | 1.45 | 1.33 | 1.54 | 0.8 | 1.07 | 1.16 | 0.8 | 1.09 | 1.03 | 0.9 | 0.9
| k | 1.6667 | 1.6667 | 1.6625 | 1.6559 | 1.6667 | 1.6462 | 1.6812 | 1.6667 | 1.6515 | 1.6613 | 1.6667 | 1.6364
80 | Center | 0.94 | 1.52 | 1.39 | 1.62 | 0.84 | 1.12 | 1.21 | 0.84 | 1.15 | 1.08 | 0.94 | 0.95
| CTDIw | 1.46 | 2.36 | 2.17 | 2.48 | 1.31 | 1.72 | 1.88 | 1.31 | 1.76 | 1.68 | 1.47 | 1.45
| k | 1.5532 | 1.5526 | 1.5612 | 1.5309 | 1.5595 | 1.5357 | 1.5537 | 1.5595 | 1.5304 | 1.5556 | 1.5638 | 1.5263
100 | Center | 2.17 | 3.51 | 3.22 | 3.75 | 1.94 | 2.6 | 2.79 | 1.94 | 2.65 | 2.49 | 2.17 | 2.19
| CTDIw | 3.12 | 5.06 | 4.65 | 5.34 | 2.8 | 3.7 | 4.03 | 2.8 | 3.78 | 3.59 | 3.14 | 3.12
| k | 1.4378 | 1.4416 | 1.4441 | 1.424 | 1.4433 | 1.4231 | 1.4444 | 1.4433 | 1.4264 | 1.4418 | 1.447 | 1.4247
120 | Center | 3.84 | 6.22 | 5.72 | 6.67 | 3.45 | 4.62 | 4.95 | 3.45 | 4.71 | 4.42 | 3.86 | 3.89
| CTDIw | 5.36 | 8.68 | 7.98 | 9.21 | 4.81 | 6.39 | 6.91 | 4.81 | 6.51 | 6.17 | 5.39 | 5.38
| k | 1.3958 | 1.3955 | 1.3951 | 1.3808 | 1.3942 | 1.3831 | 1.396 | 1.3942 | 1.3822 | 1.3959 | 1.3964 | 1.383
140 | Center | 5.89 | 9.53 | 8.76 | 10.45 | 5.28 | 7.25 | 7.59 | 5.28 | 7.38 | 6.77 | 5.91 | 6.1
| CTDIw | 8.03 | 13 | 11.95 | 14.2 | 7.2 | 9.84 | 10.36 | 7.2 | 10.03 | 9.24 | 8.07 | 8.28
| k | 1.3633 | 1.3641 | 1.3642 | 1.3589 | 1.3636 | 1.3572 | 1.365 | 1.3636 | 1.3591 | 1.3648 | 1.3655 | 1.3574

## X.cite
16 cm phantom, Standard shaped filter (Routine head)
| Collimation | 6x0.6 | 1x5 | 1x10 | 3x3 | 3x5 | 64x0.6 | 64x0.6 | 128x0.6
| Acquisition | 6x0.6 | 1x5 | 1x10 | 3x3 | 3x5 | 32x0.6 | 64x0.6 | 64x0.6
kV
70 | Center |  |  | 3.16 |  |  | 4.47 | 3.93 | 3.91
| CTDIw |  |  | 3.4 |  |  | 4.8 | 4.24 | 4.21
| k |  |  | 1.0759 |  |  | 1.0738 | 1.0789 | 1.0767
80 | Center |  |  | 4.87 |  |  | 6.88 | 6.05 | 6.02
| CTDIw |  |  | 5.14 |  |  | 7.25 | 6.42 | 6.37
| k |  |  | 1.0554 |  |  | 1.0538 | 1.0612 | 1.0581
90 | Center |  |  | 6.88 |  |  | 9.73 | 8.56 | 8.52
| CTDIw |  |  | 7.19 |  |  | 10.15 | 8.97 | 8.9
| k |  |  | 1.0451 |  |  | 1.0432 | 1.0479 | 1.0446
100 | Center |  |  | 9.15 |  |  | 12.95 | 11.38 | 11.33
| CTDIw |  |  | 9.49 |  |  | 13.39 | 11.84 | 11.75
| k |  |  | 1.0372 |  |  | 1.034 | 1.0404 | 1.0371
110 | Center |  |  | 11.63 |  |  | 16.45 | 14.46 | 14.4
| CTDIw |  |  | 12.01 |  |  | 16.94 | 14.98 | 14.86
| k |  |  | 1.0327 |  |  | 1.0298 | 1.036 | 1.0319
120 | Center |  |  | 14.33 |  |  | 20.26 | 17.81 | 17.73
| CTDIw |  |  | 14.71 |  |  | 20.75 | 18.36 | 18.21
| k |  |  | 1.0265 |  |  | 1.0242 | 1.0309 | 1.0271
130 | Center |  |  | 17.15 |  |  | 24.27 | 21.33 | 21.24
| CTDIw |  |  | 17.6 |  |  | 24.83 | 21.96 | 21.79
| k |  |  | 1.0262 |  |  | 1.0231 | 1.0295 | 1.0259
140 | Center |  |  | 20.15 |  |  | 28.51 | 25.06 | 24.95
| CTDIw |  |  | 20.65 |  |  | 29.14 | 25.77 | 25.57
| k |  |  | 1.0248 |  |  | 1.0221 | 1.0283 | 1.0248
150 | Center |  |  | 23.36 |  |  | 33.04 | 29.05 | 28.92
| CTDIw |  |  | 23.89 |  |  | 33.7 | 29.81 | 29.57
| k |  |  | 1.0227 |  |  | 1.02 | 1.0262 | 1.0225
100SnLD | Center | 2.7 |  |  |  |  | 2.08 | 1.83 | 1.82
| CTDIw | 2.69 |  |  |  |  | 2.08 | 1.84 | 1.83
| k | 0.9963 |  |  |  |  | 1 | 1.0055 | 1.0055
110SnLD | Center | [5 empty] | 3.19 | 2.8 | 2.79
| CTDIw | [5 empty] | 3.18 | 2.821 | 2.8
| k | [5 empty] | 0.99687 | 1.0075 | 1.0036
120SnLD | Center | [5 empty] | 4.55 | 4 | 3.97
| CTDIw | [5 empty] | 4.55 | 4.03 | 4
| k | [5 empty] | 1 | 1.0075 | 1.0076
130SnLD | Center | [5 empty] | 6.17 | 5.43 | 5.39
| CTDIw | [5 empty] | 6.19 | 5.49 | 5.44
| k | [5 empty] | 1.0032 | 1.011 | 1.0093
140SnLD | Center | 10.39 |  |  |  |  | 8.05 | 7.09 | 7.04
| CTDIw | 10.47 |  |  |  |  | 8.1 | 7.18 | 7.12
| k | 1.0077 |  |  |  |  | 1.0062 | 1.0127 | 1.0114
150SnLD | Center | [5 empty] | 10.19 | 8.97 | 8.91
| CTDIw | [5 empty] | 10.28 | 9.11 | 9.04
| k | [5 empty] | 1.0088 | 1.0156 | 1.0146
150SnDE | Center | [5 empty] | 6.01 | 5.3 | 5.27
| CTDIw | [5 empty] | 6.1 | 5.41 | 6.36
| k | [5 empty] | 1.015 | 1.0208 | 1.2068
120AuSn | Center | [6 empty] | 4.09
| CTDIw | [6 empty] | 4.11
| k | [6 empty] | 1.0049
140AuSn | Center | [6 empty] | 6.73
| CTDIw | [6 empty] | 6.8
| k | [6 empty] | 1.0104
16 cm phantom, Narrow shaped filter
| Collimation | 6x0.6 | 1x5 | 1x10 | 3x3 | 3x5 | 64x0.6 | 64x0.6 | 128x0.6
| Acquisition | 6x0.6 | 1x5 | 1x10 | 3x3 | 3x5 | 32x0.6 | 64x0.6 | 64x0.6
kV
70 | Center |  | 3.1 | 2.82 | 3.93 | 3.75 | 3.98 | 3.5 | 3.48
| CTDIw |  | 3.12 | 2.84 | 4 | 3.81 | 4.01 | 3.55 | 3.52
| k |  | 1.0065 | 1.0071 | 1.0178 | 1.016 | 1.0075 | 1.0143 | 1.0115
80 | Center |  | 4.83 | 4.38 | 6.13 | 5.85 | 6.19 | 5.45 | 5.42
| CTDIw |  | 4.8 | 4.37 | 6.15 | 5.86 | 6.16 | 5.46 | 5.41
| k |  | 0.99379 | 0.99772 | 1.0033 | 1.0017 | 0.99515 | 1.0018 | 0.99815
90 | Center |  | 6.89 | 6.26 | 8.75 | 8.34 | 8.84 | 7.78 | 7.74
| CTDIw |  | 6.8 | 6.19 | 8.72 | 8.3 | 8.73 | 7.74 | 7.67
| k |  | 0.98694 | 0.98882 | 0.99657 | 0.9952 | 0.98756 | 0.99486 | 0.99096
100 | Center |  | 9.22 | 8.36 | 11.69 | 11.15 | 11.81 | 10.4 | 10.35
| CTDIw |  | 9.06 | 8.24 | 11.61 | 11.06 | 11.63 | 10.3 | 10.21
| k |  | 0.98265 | 0.98565 | 0.99316 | 0.99193 | 0.98476 | 0.99038 | 0.98647
110 | Center |  | 11.77 | 10.69 | 14.94 | 14.23 | 15.1 | 13.29 | 13.23
| CTDIw |  | 11.53 | 10.5 | 14.79 | 14.09 | 14.81 | 13.13 | 13.01
| k |  | 0.97961 | 0.98223 | 0.98996 | 0.99016 | 0.98079 | 0.98796 | 0.98337
120 | Center |  | 14.55 | 13.2 | 18.44 | 17.59 | 18.64 | 16.41 | 16.33
| CTDIw |  | 14.23 | 12.95 | 18.24 | 17.38 | 18.26 | 16.19 | 16.04
| k |  | 0.97801 | 0.98106 | 0.98915 | 0.98806 | 0.97961 | 0.98659 | 0.98224
130 | Center |  | 17.51 | 15.87 | 22.17 | 21.17 | 22.41 | 19.73 | 19.64
| CTDIw |  | 17.11 | 15.57 | 21.93 | 20.9 | 21.96 | 19.46 | 19.29
| k |  | 0.97716 | 0.9811 | 0.98917 | 0.98725 | 0.97992 | 0.98632 | 0.98218
140 | Center |  | 20.65 | 18.71 | 26.14 | 24.97 | 26.42 | 23.26 | 23.15
| CTDIw |  | 20.18 | 18.36 | 25.86 | 24.65 | 25.89 | 22.95 | 22.74
| k |  | 0.97724 | 0.98129 | 0.98929 | 0.98718 | 0.97994 | 0.98667 | 0.98229
150 | Center |  | 23.98 | 21.71 | 30.33 | 28.92 | 30.66 | 26.99 | 26.86
| CTDIw |  | 23.43 | 21.31 | 30.02 | 28.6 | 30.06 | 26.64 | 26.4
| k |  | 0.97706 | 0.98158 | 0.98978 | 0.98893 | 0.98043 | 0.98703 | 0.98287
100SnLD | Center | 2.53 |  |  | 1094 | 1.84 | 1.96 | 1.72 | 1.71
| CTDIw | 2.45 |  |  | 1.89 | 1.8 | 1.9 | 1.68 | 1.66
| k | 0.96838 |  |  | 0.0017276 | 0.97826 | 0.96939 | 0.97674 | 0.97076
110SnLD | Center |  |  |  | 2.97 | 2.82 | 3 | 2.64 | 2.62
| CTDIw |  |  |  | 2.91 | 2.77 | 2.92 | 2.58 | 2.56
| k |  |  |  | 0.9798 | 0.98227 | 0.97333 | 0.97727 | 0.9771
120SnLD | Center |  |  |  | 4.25 | 4.04 | 4.3 | 3.78 | 3.76
| CTDIw |  |  |  | 4.17 | 3.97 | 4.19 | 3.7 | 3.67
| k |  |  |  | 0.98118 | 0.98267 | 0.97442 | 0.97884 | 0.97606
130SnLD | Center |  |  |  | 5.77 | 5.49 | 5.84 | 4.14 | 5.11
| CTDIw |  |  |  | 5.69 | 5.42 | 5.71 | 5.05 | 5.01
| k |  |  |  | 0.98614 | 0.98725 | 0.97774 | 1.2198 | 0.98043
140SnLD | Center | 9.87 |  |  | 7.55 | 7.18 | 7.64 | 6.73 | 6.68
| CTDIw | 9.68 |  |  | 7.48 | 7.12 | 7.5 | 6.63 | 6.58
| k | 0.98075 |  |  | 0.99073 | 0.99164 | 0.98168 | 0.98514 | 0.98503
150SnLD | Center |  |  |  | 9.58 | 9.11 | 9.69 | 8.53 | 8.47
| CTDIw |  |  |  | 9.51 | 9.05 | 9.55 | 8.43 | 8.37
| k |  |  |  | 0.99269 | 0.99341 | 0.98555 | 0.98828 | 0.98819
150SnDE | Center | [5 empty] | 5.73 | 5.06 | 5.02
| CTDIw | [5 empty] | 5.67 | 5.02 | 4.98
| k | [5 empty] | 0.98953 | 0.99209 | 0.99203
120AuSn | Center
CTDIw
k
140AuSn | Center
CTDIw
k
32 cm phantom, Standard shaped filter
| Collimation | 6x0.6 | 1x5 | 1x10 | 3x3 | 3x5 | 64x0.6 | 64x0.6 | 128x0.6
| Acquisition | 6x0.6 | 1x5 | 1x10 | 3x3 | 3x5 | 32x0.6 | 64x0.6 | 64x0.6
kV
70 | Center |  |  | 0.69 |  |  | 0.97 | 0.86 | 0.85
| CTDIw |  |  | 1.37 |  |  | 1.94 | 1.72 | 1.7
| k |  |  | 1.9855 |  |  | 2 | 2 | 2
80 | Center |  |  | 1.18 |  |  | 1.66 | 1.46 | 1.45
| CTDIw |  |  | 2.16 |  |  | 3.05 | 2.7 | 2.68
| k |  |  | 1.8305 |  |  | 1.8373 | 1.8493 | 1.8483
90 | Center |  |  | 1.79 |  |  | 2.52 | 2.21 | 2.2
| CTDIw |  |  | 3.11 |  |  | 4.39 | 3.89 | 3.86
| k |  |  | 1.7374 |  |  | 1.7421 | 1.7602 | 1.7545
100 | Center |  |  | 2.51 |  |  | 3.54 | 3.1 | 3.09
| CTDIw |  |  | 4.21 |  |  | 5.95 | 5.27 | 5.23
| k |  |  | 1.6773 |  |  | 1.6808 | 1.7 | 1.6926
110 | Center |  |  | 3.33 |  |  | 4.68 | 4.12 | 4.09
| CTDIw |  |  | 5.44 |  |  | 7.69 | 6.8 | 6.75
| k |  |  | 1.6336 |  |  | 1.6432 | 1.6505 | 1.6504
120 | Center |  |  | 4.23 |  |  | 5.97 | 5.24 | 5.2
| CTDIw |  |  | 6.8 |  |  | 9.61 | 8.49 | 8.42
| k |  |  | 1.6076 |  |  | 1.6097 | 1.6202 | 1.6192
130 | Center |  |  | 5.22 |  |  | 7.37 | 6.46 | 6.42
| CTDIw |  |  | 8.26 |  |  | 11.66 | 10.31 | 10.23
| k |  |  | 1.5824 |  |  | 1.5821 | 1.596 | 1.5935
140 | Center |  |  | 6.29 |  |  | 8.87 | 7.8 | 7.73
| CTDIw |  |  | 9.82 |  |  | 13.88 | 12.27 | 12.15
| k |  |  | 1.5612 |  |  | 1.5648 | 1.5731 | 1.5718
150 | Center |  |  | 7.43 |  |  | 10.49 | 9.23 | 9.15
| CTDIw |  |  | 11.51 |  |  | 16.23 | 14.37 | 14.23
| k |  |  | 1.5491 |  |  | 1.5472 | 1.5569 | 1.5552
100SnLD | Center | 0.92 |  |  |  |  | 0.71 | 0.63 | 0.62
| CTDIw | 1.35 |  |  |  |  | 1.05 | 0.93 | 0.92
| k | 1.4674 |  |  |  |  | 1.4789 | 1.4762 | 1.4839
110SnLD | Center | [5 empty] | 1.12 | 0.99 | 0.98
| CTDIw | [5 empty] | 1.63 | 1.44 | 1.43
| k | [5 empty] | 1.4554 | 1.4545 | 1.4592
120SnLD | Center | [5 empty] | 1.63 | 1.44 | 1.43
| CTDIw | [5 empty] | 2.36 | 2.09 | 2.07
| k | [5 empty] | 1.4479 | 1.4514 | 1.4476
130SnLD | Center | [5 empty] | 2.26 | 1.99 | 1.98
| CTDIw | [5 empty] | 3.25 | 2.88 | 2.85
| k | [5 empty] | 1.4381 | 1.4472 | 1.4394
140SnLD | Center | 3.86 |  |  |  |  | 3 | 2.63 | 2.62
| CTDIw | 5.54 |  |  |  |  | 4.3 | 3.81 | 3.77
| k | 1.4352 |  |  |  |  | 1.4333 | 1.4487 | 1.4389
150SnLD | Center | [5 empty] | 3.84 | 3.38 | 3.36
| CTDIw | [5 empty] | 5.5 | 4.87 | 4.83
| k | [5 empty] | 1.4323 | 1.4408 | 1.4375
150SnDE | Center | [5 empty] | 2.34 | 2.06 | 2.06
| CTDIw | [5 empty] | 3.32 | 2.95 | 2.93
| k | [5 empty] | 1.4188 | 1.432 | 1.4223
120AuSn | Center | [6 empty] | 1.38
| CTDIw | [6 empty] | 2.06
| k | [6 empty] | 1.4928
140AuSn | Center | [6 empty] | 2.38
| CTDIw | [6 empty] | 3.5
| k | [6 empty] | 1.4706
32 cm phantom, Narrow shaped filter
| Collimation | 6x0.6 | 1x5 | 1x10 | 3x3 | 3x5 | 64x0.6 | 64x0.6 | 128x0.6
| Acquisition | 6x0.6 | 1x5 | 1x10 | 3x3 | 3x5 | 32x0.6 | 64x0.6 | 64x0.6
kV
70 | Center |  | 0.65 | 0.6 | 0.83 | 0.79 | 0.84 | 0.74 | 0.73
| CTDIw |  | 1.15 | 1.05 | 1.45 | 1.4 | 1.48 | 1.31 | 1.3
| k |  | 1.7692 | 1.75 | 1.747 | 1.7722 | 1.7619 | 1.7703 | 1.7808
80 | Center |  | 1.12 | 1.03 | 1.42 | 1.36 | 1.44 | 1.27 | 1.26
| CTDIw |  | 1.83 | 1.67 | 2.31 | 2.22 | 2.35 | 2.09 | 2.07
| k |  | 1.6339 | 1.6214 | 1.6268 | 1.6324 | 1.6319 | 1.6457 | 1.6429
90 | Center |  | 1.72 | 1.57 | 2.18 | 2.08 | 2.21 | 1.94 | 1.92
| CTDIw |  | 2.66 | 2.44 | 3.37 | 3.24 | 3.43 | 3.04 | 3.01
| k |  | 1.5465 | 1.5541 | 1.5459 | 1.5577 | 1.552 | 1.567 | 1.5677
100 | Center |  | 2.42 | 2.21 | 3.08 | 2.93 | 3.12 | 2.74 | 2.71
| CTDIw |  | 3.63 | 3.32 | 4.59 | 4.42 | 4.68 | 4.15 | 4.1
| k |  | 1.5 | 1.5023 | 1.4903 | 1.5085 | 1.5 | 1.5146 | 1.5129
110 | Center |  | 3.21 | 2.94 | 4.09 | 3.89 | 4.14 | 3.65 | 3.61
| CTDIw |  | 4.71 | 4.31 | 5.96 | 5.74 | 6.08 | 5.39 | 5.33
| k |  | 1.4673 | 1.466 | 1.4572 | 1.4756 | 1.4686 | 1.4767 | 1.4765
120 | Center |  | 4.11 | 3.76 | 5.22 | 4.97 | 5.29 | 4.65 | 4.61
| CTDIw |  | 5.92 | 5.41 | 7.49 | 7.2 | 7.63 | 6.76 | 6.69
| k |  | 1.4404 | 1.4388 | 1.4349 | 1.4487 | 1.4423 | 1.4538 | 1.4512
130 | Center |  | 5.08 | 4.65 | 6.47 | 6.14 | 6.55 | 5.76 | 5.71
| CTDIw |  | 7.22 | 6.61 | 9.16 | 8.78 | 9.31 | 8.25 | 8.17
| k |  | 1.4213 | 1.4215 | 1.4158 | 1.43 | 1.4214 | 1.4323 | 1.4308
140 | Center |  | 6.14 | 5.62 | 7.81 | 7.43 | 7.92 | 6.96 | 6.9
| CTDIw |  | 8.63 | 7.91 | 10.94 | 10.5 | 11.13 | 9.86 | 9.76
| k |  | 1.4055 | 1.4075 | 1.4008 | 1.4132 | 1.4053 | 1.4167 | 1.4145
150 | Center |  | 7.29 | 6.66 | 9.26 | 8.8 | 9.38 | 8.26 | 8.18
| CTDIw |  | 10.15 | 9.29 | 12.85 | 12.32 | 13.08 | 11.6 | 11.47
| k |  | 1.3923 | 1.3949 | 1.3877 | 1.4 | 1.3945 | 1.4044 | 1.4022
100SnLD | Center | 0.83 |  |  | 0.63 | 0.6 | 0.64 | 0.56 | 0.56
| CTDIw | 1.1 |  |  | 0.84 | 0.81 | 0.85 | 0.76 | 0.75
| k | 1.3253 |  |  | 1.3333 | 1.35 | 1.3281 | 1.3571 | 1.3393
110SnLD | Center |  |  |  | 1 | 0.95 | 1.01 | 0.89 | 0.88
| CTDIw |  |  |  | 1.31 | 1.26 | 1.33 | 1.18 | 1.17
| k |  |  |  | 1.31 | 1.3263 | 1.3168 | 1.3258 | 1.3295
120SnLD | Center |  |  |  | 1.46 | 1.39 | 1.48 | 1.3 | 1.29
| CTDIw |  |  |  | 1.91 | 1.84 | 1.94 | 1.75 | 1.71
| k |  |  |  | 1.3082 | 1.3237 | 1.3108 | 1.3462 | 1.3256
130SnLD | Center |  |  |  | 2.02 | 1.93 | 2.05 | 1.8 | 1.79
| CTDIw |  |  |  | 2.63 | 2.53 | 2.68 | 2.38 | 2.36
| k |  |  |  | 1.302 | 1.3109 | 1.3073 | 1.3222 | 1.3184
140SnLD | Center |  |  |  | 2.69 | 2.56 | 2.73 | 2.4 | 2.38
| CTDIw |  |  |  | 3.5 | 3.36 | 3.56 | 3.16 | 3.13
| k |  |  |  | 1.3011 | 1.3125 | 1.304 | 1.3167 | 1.3151
150SnLD | Center |  |  |  | 3.46 | 3.3 | 3.51 | 3.08 | 3.06
| CTDIw |  |  |  | 4.49 | 4.31 | 4.57 | 4.06 | 4.02
| k |  |  |  | 1.2977 | 1.3061 | 1.302 | 1.3182 | 1.3137
150SnDE | Center | [5 empty] | 2.14 | 1.89 | 1.88
| CTDIw | [5 empty] | 2.78 | 2.47 | 2.45
| k | [5 empty] | 1.2991 | 1.3069 | 1.3032
120AuSn | Center
CTDIw
k
140AuSn | Center
CTDIw
k

## Confidence 20
16 cm phantom, Standard shaped filter (Routine head)
| Collimation | 2x1 | 1x5 | 10x0.6 | 6x1.2 | 1x10 | 20x0.6 | 12x1.2 | 16x1.2
| Total | 2 | 5 | 6 | 7.2 | 10 | 12 | 14.4 | 19.2
| Acquisition | 2x1 | 1x5 | 10x0.6 | 6x1.2 | 1x10 | 20x0.6 | 12x1.2 | 16x1.2
kV
70 | Center | 2.86 | 2.57 | 3.54 | 3.69 | 2.57 | 3.48 | 3.29 | 3.15
| CTDIw | 3.06 | 2.74 | 3.79 | 3.94 | 2.74 | 3.71 | 3.52 | 3.37
| k | 1.0699 | 1.0661 | 1.0706 | 1.0678 | 1.0661 | 1.0661 | 1.0699 | 1.0698
80 | Center | 4.57 | 4.1 | 5.65 | 5.89 | 4.1 | 5.54 | 5.25 | 5.03
| CTDIw | 4.77 | 4.28 | 5.91 | 6.16 | 4.28 | 5.8 | 5.49 | 5.26
| k | 1.0438 | 1.0439 | 1.046 | 1.0458 | 1.0439 | 1.0469 | 1.0457 | 1.0457
100 | Center | 8.84 | 7.93 | 10.94 | 11.4 | 7.93 | 10.73 | 10.17 | 9.73
| CTDIw | 9.02 | 8.1 | 11.17 | 11.64 | 8.1 | 10.96 | 10.39 | 9.94
| k | 1.0204 | 1.0214 | 1.021 | 1.0211 | 1.0214 | 1.0214 | 1.0216 | 1.0216
120 | Center | 14.89 | 13.36 | 18.44 | 19.21 | 13.36 | 18.08 | 17.14 | 16.4
| CTDIw | 15.07 | 13.52 | 18.66 | 19.43 | 13.52 | 18.3 | 17.34 | 16.59
| k | 1.0121 | 1.012 | 1.0119 | 1.0115 | 1.012 | 1.0122 | 1.0117 | 1.0116
140 | Center | 21.22 | 19.04 | 26.27 | 27.37 | 19.04 | 25.77 | 24.42 | 23.37
| CTDIw | 21.38 | 19.18 | 26.47 | 27.57 | 19.18 | 15.96 | 24.6 | 23.54
| k | 1.0075 | 1.0074 | 1.0076 | 1.0073 | 1.0074 | 0.61932 | 1.0074 | 1.0073
16 cm phantom, Narrow shaped filter (Pediatric body)
| Collimation | 2x1 | 1x5 | 10x0.6 | 6x1.2 | 1x10 | 20x0.6 | 12x1.2 | 16x1.2
| Total | 2 | 5 | 6 | 7.2 | 10 | 12 | 14.4 | 19.2
| Acquisition | 2x1 | 1x5 | 10x0.6 | 6x1.2 | 1x10 | 20x0.6 | 12x1.2 | 16x1.2
kV
70 | Center | 2.54 | 2.28 | 3.15 | 3.28 | 2.28 | 3.09 | 2.93 | 2.8
| CTDIw | 2.55 | 2.28 | 3.15 | 3.28 | 2.28 | 3.09 | 2.93 | 2.8
| k | 1.0039 | 1 | 1 | 1 | 1 | 1 | 1 | 1
80 | Center | 4.11 | 3.69 | 5.09 | 5.3 | 3.69 | 4.99 | 4.73 | 4.53
| CTDIw | 4.05 | 3.64 | 5.02 | 5.23 | 3.64 | 4.92 | 4.66 | 5.26
| k | 0.9854 | 0.98645 | 0.98625 | 0.98679 | 0.98645 | 0.98597 | 0.9852 | 1.1611
100 | Center | 8.04 | 7.21 | 9.95 | 10.37 | 7.21 | 9.76 | 9.25 | 8.85
| CTDIw | 7.82 | 7.02 | 9.68 | 10.09 | 7.02 | 9.5 | 9 | 8.61
| k | 0.97264 | 0.97365 | 0.97286 | 0.973 | 0.97365 | 0.97336 | 0.97297 | 0.97288
120 | Center | 11.68 | 12.28 | 16.94 | 17.65 | 12.28 | 16.62 | 15.75 | 15.07
| CTDIw | 13.28 | 11.92 | 16.44 | 17.13 | 11.92 | 16.13 | 15.28 | 14.63
| k | 1.137 | 0.97068 | 0.97048 | 0.97054 | 0.97068 | 0.97052 | 0.97016 | 0.9708
140 | Center | 19.6 | 17.58 | 24.27 | 25.28 | 17.58 | 23.8 | 22.56 | 21.58
| CTDIw | 19.01 | 17.05 | 23.53 | 24.51 | 17.05 | 23.08 | 21.87 | 20.93
| k | 0.9699 | 0.96985 | 0.96951 | 0.96954 | 0.96985 | 0.96975 | 0.96941 | 0.96988
32 cm phantom, Standard shaped filter (Adult body)
| Collimation | 2x1 | 1x5 | 10x0.6 | 6x1.2 | 1x10 | 20x0.6 | 12x1.2 | 16x1.2
| Total | 2 | 5 | 6 | 7.2 | 10 | 12 | 14.4 | 19.2
| Acquisition | 2x1 | 1x5 | 10x0.6 | 6x1.2 | 1x10 | 20x0.6 | 12x1.2 | 16x1.2
kV
70 | Center | 0.68 | 0.61 | 0.84 | 0.88 | 0.61 | 0.83 | 0.79 | 0.75
| CTDIw | 1.37 | 1.22 | 1.69 | 1.76 | 1.22 | 1.66 | 1.57 | 1.59
| k | 2.0147 | 2 | 2.0119 | 2 | 2 | 2 | 1.9873 | 2.12
80 | Center | 1.18 | 1.06 | 1.46 | 1.52 | 1.06 | 1.43 | 1.36 | 1.3
| CTDIw | 2.18 | 1.96 | 2.7 | 2.81 | 1.96 | 2.65 | 2.51 | 2.4
| k | 1.8475 | 1.8491 | 1.8493 | 1.8487 | 1.8491 | 1.8531 | 1.8456 | 1.8462
100 | Center | 2.56 | 2.29 | 3.17 | 3.3 | 2.29 | 3.1 | 2.94 | 2.82
| CTDIw | 4.28 | 3.84 | 5.3 | 5.52 | 3.84 | 5.19 | 4.92 | 4.71
| k | 1.6719 | 1.6769 | 1.6719 | 1.6727 | 1.6769 | 1.6742 | 1.6735 | 1.6702
120 | Center | 4.61 | 4.14 | 5.71 | 5.95 | 4.14 | 5.6 | 5.31 | 5.08
| CTDIw | 7.33 | 6.58 | 9.08 | 9.46 | 6.58 | 8.91 | 8.44 | 8.08
| k | 1.59 | 1.5894 | 1.5902 | 1.5899 | 1.5894 | 1.5911 | 1.5895 | 1.5906
140 | Center | 6.88 | 6.17 | 8.52 | 8.87 | 6.17 | 8.36 | 7.92 | 7.58
| CTDIw | 10.59 | 9.5 | 13.11 | 13.66 | 9.5 | 12.86 | 12.19 | 11.66
| k | 1.5392 | 1.5397 | 1.5387 | 1.54 | 1.5397 | 1.5383 | 1.5391 | 1.5383
32 cm phantom, Narrow shaped filter (Cardiac)
| Collimation | 2x1 | 1x5 | 10x0.6 | 6x1.2 | 1x10 | 20x0.6 | 12x1.2 | 16x1.2
| Total | 2 | 5 | 6 | 7.2 | 10 | 12 | 14.4 | 19.2
| Acquisition | 2x1 | 1x5 | 10x0.6 | 6x1.2 | 1x10 | 20x0.6 | 12x1.2 | 16x1.2
kV
70 | Center | 0.58 | 0.52 | 0.72 | 0.75 | 0.52 | 0.7 | 0.67 | 0.64
| CTDIw | 1.01 | 0.91 | 1.25 | 1.3 | 0.91 | 1.23 | 1.16 | 1.11
| k | 1.7414 | 1.75 | 1.7361 | 1.7333 | 1.75 | 1.7571 | 1.7313 | 1.7344
80 | Center | 1.03 | 0.92 | 1.27 | 1.33 | 0.92 | 1.25 | 1.18 | 1.13
| CTDIw | 1.66 | 1.49 | 2.06 | 2.14 | 1.49 | 2.02 | 1.91 | 1.83
| k | 1.6117 | 1.6196 | 1.622 | 1.609 | 1.6196 | 1.616 | 1.6186 | 1.6195
100 | Center | 2.25 | 2.02 | 2.78 | 2.9 | 2.02 | 2.73 | 2.59 | 2.47
| CTDIw | 3.35 | 3.01 | 4.15 | 4.32 | 3.01 | 4.07 | 3.86 | 3.69
| k | 1.4889 | 1.4901 | 1.4928 | 1.4897 | 1.4901 | 1.4908 | 1.4903 | 1.4939
120 | Center | 4.09 | 3.67 | 5.07 | 5.28 | 3.67 | 4.97 | 4.71 | 4.51
| CTDIw | 5.86 | 5.26 | 7.26 | 7.56 | 5.26 | 7.12 | 6.75 | 6.46
| k | 1.4328 | 1.4332 | 1.432 | 1.4318 | 1.4332 | 1.4326 | 1.4331 | 1.4324
140 | Center | 6.16 | 5.52 | 7.62 | 7.94 | 5.52 | 7.48 | 7.09 | 6.78
| CTDIw | 8.6 | 7.71 | 10.64 | 11.09 | 7.71 | 10.44 | 9.89 | 9.47
| k | 1.3961 | 1.3967 | 1.3963 | 1.3967 | 1.3967 | 1.3957 | 1.3949 | 1.3968

## Confidence 64
16 cm phantom, Standard shaped filter (Routine head)
| Collimation | 2x1 | 1x5 | 10x0.6 | 6x1.2 | 1x10 | 20x0.6 | 12x1.2 | 32x0.6 | 16x1.2
| Total | 2 | 5 | 6 | 7.2 | 10 | 12 | 14.4 | 19.2 | 19.2
| Acquisition | 2x1 | 1x5 | 20x0.6 | 6x1.2 | 1x10 | 40x0.6 | 12x1.2 | 64x0.6 | 16x1.2
kV
70 | Center | 2.86 | 2.57 | 3.41 | 3.69 | 2.57 | 3.48 | 3.29 | 3.24 | 3.15
| CTDIw | 3.06 | 2.74 | 3.63 | 3.94 | 2.74 | 3.7 | 3.52 | 3.44 | 3.37
| k | 1.0699 | 1.0661 | 1.0645 | 1.0678 | 1.0661 | 1.0632 | 1.0699 | 1.0617 | 1.0698
80 | Center | 4.57 | 4.1 | 5.44 | 5.89 | 4.1 | 5.55 | 5.25 | 5.17 | 5.03
| CTDIw | 4.77 | 4.28 | 5.67 | 6.16 | 4.28 | 5.78 | 5.49 | 5.38 | 5.26
| k | 1.0438 | 1.0439 | 1.0423 | 1.0458 | 1.0439 | 1.0414 | 1.0457 | 1.0406 | 1.0457
100 | Center | 8.84 | 7.93 | 10.53 | 11.4 | 7.93 | 10.73 | 10.17 | 10 | 9.73
| CTDIw | 9.02 | 8.1 | 10.7 | 11.64 | 8.1 | 10.91 | 10.39 | 10.16 | 9.94
| k | 1.0204 | 1.0214 | 1.0161 | 1.0211 | 1.0214 | 1.0168 | 1.0216 | 1.016 | 1.0216
120 | Center | 14.89 | 13.36 | 17.69 | 19.21 | 13.36 | 18.03 | 17.14 | 16.8 | 16.4
| CTDIw | 15.07 | 13.52 | 17.84 | 19.43 | 13.52 | 18.18 | 17.34 | 16.94 | 16.59
| k | 1.0121 | 1.012 | 1.0085 | 1.0115 | 1.012 | 1.0083 | 1.0117 | 1.0083 | 1.0116
140 | Center | 21.22 | 19.04 | 25.23 | 27.37 | 19.04 | 25.72 | 24.42 | 23.96 | 23.37
| CTDIw | 21.38 | 19.18 | 25.35 | 27.57 | 19.18 | 25.83 | 24.6 | 24.07 | 23.54
| k | 1.0075 | 1.0074 | 1.0048 | 1.0073 | 1.0074 | 1.0043 | 1.0074 | 1.0046 | 1.0073
16 cm phantom, Narrow shaped filter (Pediatric body)
| Collimation | 2x1 | 1x5 | 10x0.6 | 6x1.2 | 1x10 | 20x0.6 | 12x1.2 | 32x0.6 | 16x1.2
| Total | 2 | 5 | 6 | 7.2 | 10 | 12 | 14.4 | 19.2 | 19.2
| Acquisition | 2x1 | 1x5 | 20x0.6 | 6x1.2 | 1x10 | 40x0.6 | 12x1.2 | 64x0.6 | 16x1.2
kV
70 | Center | 2.54 | 2.28 | 3.02 | 3.28 | 2.28 | 3.08 | 2.93 | 2.87 | 2.8
| CTDIw | 2.55 | 2.28 | 3.01 | 3.28 | 2.28 | 3.07 | 2.93 | 2.86 | 2.8
| k | 1.0039 | 1 | 0.99669 | 1 | 1 | 0.99675 | 1 | 0.99652 | 1
80 | Center | 4.11 | 3.69 | 4.91 | 5.3 | 3.69 | 5 | 4.73 | 4.66 | 4.53
| CTDIw | 4.05 | 3.64 | 4.8 | 5.23 | 3.64 | 4.89 | 4.66 | 4.56 | 5.26
| k | 0.9854 | 0.98645 | 0.9776 | 0.98679 | 0.98645 | 0.978 | 0.9852 | 0.97854 | 1.1611
100 | Center | 8.04 | 7.21 | 9.58 | 10.37 | 7.21 | 9.76 | 9.25 | 9.09 | 8.85
| CTDIw | 7.82 | 7.02 | 9.26 | 10.09 | 7.02 | 9.43 | 9 | 8.79 | 8.61
| k | 0.97264 | 0.97365 | 0.9666 | 0.973 | 0.97365 | 0.96619 | 0.97297 | 0.967 | 0.97288
120 | Center | 11.68 | 12.28 | 16.3 | 17.65 | 12.28 | 16.61 | 15.75 | 15.48 | 15.07
| CTDIw | 13.28 | 11.92 | 15.71 | 17.13 | 11.92 | 16.01 | 15.28 | 14.92 | 14.63
| k | 1.137 | 0.97068 | 0.9638 | 0.97054 | 0.97068 | 0.96388 | 0.97016 | 0.96382 | 0.9708
140 | Center | 19.6 | 17.58 | 23.34 | 25.28 | 17.58 | 23.79 | 22.56 | 22.16 | 21.58
| CTDIw | 19.01 | 17.05 | 22.49 | 24.51 | 17.05 | 22.92 | 21.87 | 21.35 | 20.93
| k | 0.9699 | 0.96985 | 0.96358 | 0.96954 | 0.96985 | 0.96343 | 0.96941 | 0.96345 | 0.96988
32 cm phantom, Standard shaped filter (Adult body)
| Collimation | 2x1 | 1x5 | 10x0.6 | 6x1.2 | 1x10 | 20x0.6 | 12x1.2 | 32x0.6 | 16x1.2
| Total | 2 | 5 | 6 | 7.2 | 10 | 12 | 14.4 | 19.2 | 19.2
| Acquisition | 2x1 | 1x5 | 20x0.6 | 6x1.2 | 1x10 | 40x0.6 | 12x1.2 | 64x0.6 | 16x1.2
kV
70 | Center | 0.68 | 0.61 | 0.82 | 0.88 | 0.61 | 0.83 | 0.79 | 0.78 | 0.75
| CTDIw | 1.37 | 1.22 | 1.61 | 1.76 | 1.22 | 1.64 | 1.57 | 1.52 | 1.59
| k | 2.0147 | 2 | 1.9634 | 2 | 2 | 1.9759 | 1.9873 | 1.9487 | 2.12
80 | Center | 1.18 | 1.06 | 1.4 | 1.52 | 1.06 | 1.43 | 1.36 | 1.33 | 1.3
| CTDIw | 2.18 | 1.96 | 2.58 | 2.81 | 1.96 | 2.62 | 2.51 | 2.45 | 2.4
| k | 1.8475 | 1.8491 | 1.8429 | 1.8487 | 1.8491 | 1.8322 | 1.8456 | 1.8421 | 1.8462
100 | Center | 2.56 | 2.29 | 3.04 | 3.3 | 2.29 | 3.1 | 2.94 | 2.89 | 2.82
| CTDIw | 4.28 | 3.84 | 5.05 | 5.52 | 3.84 | 5.15 | 4.92 | 4.8 | 4.71
| k | 1.6719 | 1.6769 | 1.6612 | 1.6727 | 1.6769 | 1.6613 | 1.6735 | 1.6609 | 1.6702
120 | Center | 4.61 | 4.14 | 5.48 | 5.95 | 4.14 | 5.59 | 5.31 | 5.21 | 5.08
| CTDIw | 7.33 | 6.58 | 8.68 | 9.46 | 6.58 | 8.85 | 8.44 | 8.25 | 8.08
| k | 1.59 | 1.5894 | 1.5839 | 1.5899 | 1.5894 | 1.5832 | 1.5895 | 1.5835 | 1.5906
140 | Center | 6.88 | 6.17 | 8.2 | 8.87 | 6.17 | 8.36 | 7.92 | 7.79 | 7.58
| CTDIw | 10.59 | 9.5 | 12.51 | 13.66 | 9.5 | 12.75 | 12.19 | 11.88 | 11.66
| k | 1.5392 | 1.5397 | 1.5256 | 1.54 | 1.5397 | 1.5251 | 1.5391 | 1.525 | 1.5383
32 cm phantom, Narrow shaped filter (Cardiac)
| Collimation | 2x1 | 1x5 | 10x0.6 | 6x1.2 | 1x10 | 20x0.6 | 12x1.2 | 32x0.6 | 16x1.2
| Total | 2 | 5 | 6 | 7.2 | 10 | 12 | 14.4 | 19.2 | 19.2
| Acquisition | 2x1 | 1x5 | 20x0.6 | 6x1.2 | 1x10 | 40x0.6 | 12x1.2 | 64x0.6 | 16x1.2
kV
70 | Center | 0.58 | 0.52 | 0.69 | 0.75 | 0.52 | 0.7 | 0.67 | 0.66 | 0.64
| CTDIw | 1.01 | 0.91 | 1.19 | 1.3 | 0.91 | 1.22 | 1.16 | 1.13 | 1.11
| k | 1.7414 | 1.75 | 1.7246 | 1.7333 | 1.75 | 1.7429 | 1.7313 | 1.7121 | 1.7344
80 | Center | 1.03 | 0.92 | 1.22 | 1.33 | 0.92 | 1.24 | 1.18 | 1.16 | 1.13
| CTDIw | 1.66 | 1.49 | 1.96 | 2.14 | 1.49 | 2 | 1.91 | 1.86 | 1.83
| k | 1.6117 | 1.6196 | 1.6066 | 1.609 | 1.6196 | 1.6129 | 1.6186 | 1.6034 | 1.6195
100 | Center | 2.25 | 2.02 | 2.67 | 2.9 | 2.02 | 2.72 | 2.59 | 2.54 | 2.47
| CTDIw | 3.35 | 3.01 | 3.95 | 4.32 | 3.01 | 4.03 | 3.86 | 3.75 | 3.69
| k | 1.4889 | 1.4901 | 1.4794 | 1.4897 | 1.4901 | 1.4816 | 1.4903 | 1.4764 | 1.4939
120 | Center | 4.09 | 3.67 | 4.87 | 5.28 | 3.67 | 4.96 | 4.71 | 4.62 | 4.51
| CTDIw | 5.86 | 5.26 | 6.89 | 7.56 | 5.26 | 7.02 | 6.75 | 6.54 | 6.46
| k | 1.4328 | 1.4332 | 1.4148 | 1.4318 | 1.4332 | 1.4153 | 1.4331 | 1.4156 | 1.4324
140 | Center | 6.16 | 5.52 | 7.29 | 7.94 | 5.52 | 7.43 | 7.09 | 6.92 | 6.78
| CTDIw | 8.6 | 7.71 | 10.12 | 11.09 | 7.71 | 10.32 | 9.89 | 9.61 | 9.47
| k | 1.3961 | 1.3967 | 1.3882 | 1.3967 | 1.3967 | 1.389 | 1.3949 | 1.3887 | 1.3968

## go.Top
16 cm phantom (Routine head)
| Collimation | 6x0.6 |  | 1x5 |  | 1X10 |  | 3x5 |  |  |  | 64x0.6
| Total | 3.6 |  | 5 |  | 10 |  | 15 |  |  |  | 38.4
| Acquisition | 6x0.6 |  | 1x5 |  | 1X10 |  | 3x5 |  |  |  | 64x0.6
| Shaped filter | Sn | Narrow Sn | Standard | Narrow | Standard | Narrow | Standard | Sn | Narrow | Narrow Sn | Standard | Sn | Narrow | Narrow Sn | AuSn
kV
70 | Center |  |  | 3.7 | 3.3 | 3.7 | 3.3 | 3.7 |  | 3.3 |  | 3.8 |  | 3.4
| CTDIw |  |  | 4 | 3.3 | 3.9 | 3.3 | 4 |  | 3.3 |  | 4.1 |  | 3.4
| k |  |  | 1.0811 | 1 | 1.0541 | 1 | 1.0811 |  | 1 |  | 1.0789 |  | 1
80 | Center |  |  | 5.7 | 5.1 | 5.7 | 5.1 | 5.7 |  | 5.1 |  | 5.9 |  | 5.3
| CTDIw |  |  | 6 | 5.1 | 6 | 5.1 | 6 |  | 5.1 |  | 6.2 |  | 5.3
| k |  |  | 1.0526 | 1 | 1.0526 | 1 | 1.0526 |  | 1 |  | 1.0508 |  | 1
90 | Center |  |  | 8.2 | 7.4 | 8.1 | 7.4 | 8.2 |  | 7.4 |  | 8.4 |  | 7.7
| CTDIw |  |  | 8.5 | 7.3 | 8.4 | 7.2 | 8.5 |  | 7.3 |  | 8.8 |  | 7.6
| k |  |  | 1.0366 | 0.98649 | 1.037 | 0.97297 | 1.0366 |  | 0.98649 |  | 1.0476 |  | 0.98701
100 | Center | 1.3 | 1.2 | 11.1 | 10.1 | 11 | 10.1 | 11 | 0.9 | 10.1 | 0.9 | 11.4 | 0.9 | 10.4 | 0.9
| CTDIw | 1.3 | 1.2 | 11.5 | 9.9 | 11.3 | 9.8 | 11.4 | 0.9 | 9.9 | 0.8 | 11.8 | 0.9 | 10.3 | 0.9
| k | 1 | 1 | 1.036 | 0.9802 | 1.0273 | 0.9703 | 1.0364 | 1 | 0.9802 | 0.88889 | 1.0351 | 1 | 0.99038 | 1
110 | Center |  |  | 14.2 | 13.1 | 14.1 | 13 | 14.2 | 1.5 | 13 | 1.4 | 14.6 | 1.6 | 13.5 | 1.5
| CTDIw |  |  | 14.7 | 12.8 | 14.5 | 12.7 | 14.6 | 1.5 | 12.8 | 1.4 | 15.1 | 1.6 | 13.2 | 1.5
| k |  |  | 1.0352 | 0.9771 | 1.0284 | 0.97692 | 1.0282 | 1 | 0.98462 | 1 | 1.0342 | 1 | 0.97778 | 1
120 | Center |  |  | 17.8 | 16.3 | 17.6 | 16.2 | 17.7 | 2.3 | 16.3 | 2.2 | 18.3 | 2.4 | 16.9 | 2.3 | 5.4
| CTDIw |  |  | 18.2 | 16 | 18 | 15.6 | 18.1 | 2.3 | 16 | 2.2 | 18.8 | 2.4 | 16.5 | 2.2 | 5.5
| k |  |  | 1.0225 | 0.9816 | 1.0227 | 0.96296 | 1.0226 | 1 | 0.9816 | 1 | 1.0273 | 1 | 0.97633 | 0.95652 | 1.0185
130 | Center |  |  | 21.3 | 19.6 | 21 | 19.5 | 21.2 | 3.3 | 19.6 | 3.2 | 21.9 | 3.5 | 20.2 | 3.3
| CTDIw |  |  | 21.9 | 19.3 | 21.5 | 19.1 | 21.7 | 3.4 | 19.2 | 3.1 | 22.4 | 3.5 | 19.9 | 3.2
| k |  |  | 1.0282 | 0.98469 | 1.0238 | 0.97949 | 1.0236 | 1.0303 | 0.97959 | 0.96875 | 1.0228 | 1 | 0.98515 | 0.9697
140 | Center | 6.4 | 6.1 | 25 | 23.2 | 24.7 | 23 | 24.9 | 4.6 | 23.2 | 4.4 | 25.8 | 4.7 | 23.9 | 4.5 | 8.7
| CTDIw | 6.5 | 6 | 25.6 | 22.8 | 25.3 | 22.5 | 25.4 | 4.6 | 22.6 | 4.3 | 26.3 | 4.8 | 23.4 | 4.4 | 8.7
| k | 1.0156 | 0.98361 | 1.024 | 0.98276 | 1.0243 | 0.97826 | 1.0201 | 1 | 0.97414 | 0.97727 | 1.0194 | 1.0213 | 0.97908 | 0.97778 | 1
32 cm phantom (Adult body)
| Collimation | 6x0.6 |  | 1x5 |  | 1X10 |  | 3x5 |  |  |  | 64x0.6
| Total | 3.6 |  | 5 |  | 10 |  | 15 |  |  |  | 38.4
| Acquisition | 6x0.6 |  | 1x5 |  | 1X10 |  | 3x5 |  |  |  | 64x0.6
| Shaped filter | Sn | Narrow Sn | Standard | Narrow | Standard | Narrow | Standard | Sn | Narrow | Narrow Sn | Standard | Sn | Narrow | Narrow Sn | AuSn
kV
70 | Center |  |  | 0.9 | 0.7 | 0.8 | 3.3 | 0.8 |  | 0.7 |  | 0.9 |  | 0.8
| CTDIw |  |  | 1.7 | 1.3 | 1.7 | 3.3 | 1.7 |  | 1.3 |  | 1.7 |  | 1.3
| k |  |  | 1.8889 | 1.8571 | 2.125 | 1 | 2.125 |  | 1.8571 |  | 1.8889 |  | 1.625
80 | Center |  |  | 1.4 | 1.2 | 1.4 | 5.1 | 1.4 |  | 1.2 |  | 1.5 |  | 1.3
| CTDIw |  |  | 2.7 | 2 | 2.6 | 5.1 | 2.6 |  | 2 |  | 2.7 |  | 2.1
| k |  |  | 1.9286 | 1.6667 | 1.8571 | 1 | 1.8571 |  | 1.6667 |  | 1.8 |  | 1.6154
90 | Center |  |  | 2.2 | 1.9 | 2.2 | 7.4 | 2.2 |  | 1.9 |  | 2.3 |  | 2
| CTDIw |  |  | 3.9 | 3 | 3.8 | 7.2 | 3.8 |  | 3 |  | 4 |  | 3.1
| k |  |  | 1.7727 | 1.5789 | 1.7273 | 0.97297 | 1.7273 |  | 1.5789 |  | 1.7391 |  | 1.55
100 | Center | 0.5 | 0.4 | 3.2 | 2.8 | 3.1 | 10.1 | 3.1 | 0.3 | 2.8 | 0.3 | 3.2 | 0.3 | 2.9 | 0.3
| CTDIw | 0.7 | 0.6 | 5.4 | 4.2 | 5.2 | 9.8 | 5.3 | 0.5 | 4.1 | 0.4 | 5.4 | 0.5 | 4.3 | 0.4
| k | 1.4 | 1.5 | 1.6875 | 1.5 | 1.6774 | 0.9703 | 1.7097 | 1.6667 | 1.4643 | 1.3333 | 1.6875 | 1.6667 | 1.4828 | 1.3333
110 | Center |  |  | 4.2 | 3.7 | 4.2 | 13 | 4.2 | 0.6 | 3.7 | 0.5 | 4.3 | 0.6 | 3.8 | 0.5
| CTDIw |  |  | 7 | 5.5 | 6.9 | 12.7 | 6.9 | 0.8 | 5.4 | 0.7 | 7.1 | 0.9 | 5.6 | 0.7
| k |  |  | 1.6667 | 1.4865 | 1.6429 | 0.97692 | 1.6429 | 1.3333 | 1.4595 | 1.4 | 1.6512 | 1.5 | 1.4737 | 1.4
120 | Center |  |  | 5.4 | 4.8 | 5.3 | 16.2 | 5.4 | 0.9 | 4.8 | 0.8 | 5.6 | 0.9 | 5 | 0.8 | 1.8
| CTDIw |  |  | 8.9 | 7 | 8.6 | 15.6 | 8.7 | 1.3 | 6.9 | 1.1 | 9 | 1.3 | 7.1 | 1.1 | 2.8
| k |  |  | 1.6481 | 1.4583 | 1.6226 | 0.96296 | 1.6111 | 1.4444 | 1.4375 | 1.375 | 1.6071 | 1.4444 | 1.42 | 1.375 | 1.5556
130 | Center |  |  | 6.7 | 5.9 | 6.6 | 19.5 | 6.6 | 1.3 | 5.9 | 1.2 | 6.8 | 1.3 | 6.1 | 1.2
| CTDIw |  |  | 10.8 | 8.5 | 10.5 | 19.1 | 10.5 | 1.9 | 8.4 | 1.5 | 10.9 | 1.9 | 8.7 | 1.6
| k |  |  | 1.6119 | 1.4407 | 1.5909 | 0.97949 | 1.5909 | 1.4615 | 1.4237 | 1.25 | 1.6029 | 1.4615 | 1.4262 | 1.3333
140 | Center | 2.5 | 2.3 | 8.1 | 7.2 | 7.9 | 23 | 8 | 1.8 | 7.1 | 1.6 | 8.3 | 1.9 | 7.4 | 1.7 | 3.1
| CTDIw | 3.7 | 3 | 12.9 | 10.2 | 12.5 | 22.5 | 12.5 | 2.6 | 10 | 2.1 | 13 | 2.7 | 10.4 | 2.2 | 4.6
| k | 1.48 | 1.3043 | 1.5926 | 1.4167 | 1.5823 | 0.97826 | 1.5625 | 1.4444 | 1.4085 | 1.3125 | 1.5663 | 1.4211 | 1.4054 | 1.2941 | 1.4839

## Definition Edge
16 cm phantom, Standard shaped filter (Routine head)
| Collimation | 2x1 | 6x0.6 | 8x0.6 | 8x0.6 | 1x5 | 10x0.6 | 6x1.2 | 1x10 | 20x0.6 | 12x1.2 | 32x1.2 | 64x0.6
| Total | 2 | 3.6 | 4.8 | 4.8 | 5 | 6 | 7.2 | 10 | 12 | 14.4 | 38.4 | 38.4
| Acquisition | 2x1 | 6x0.6 | 8x0.6 | 16x0.6 | 1x5 | 20x0.6 | 6x1.2 | 1x10 | 40x0.6 | 12x1.2 | 32x1.2 | 28x0.6
| Shaped filter | Standard | Standard | Standard | Standard | Standard | Standard | Standard | Standard | Standard | Standard | Standard | Standard
kV
70 | Center | 2.6 | 4.22 | 3.88 | 4.58 | 2.34 | 3.17 | 3.36 | 2.34 | 3.23 | 3 | 2.62 | 2.67
| CTDIw | 2.74 | 4.43 | 4.07 | 4.78 | 2.46 | 3.31 | 3.53 | 2.46 | 3.38 | 3.15 | 2.75 | 2.79
| k | 1.0538 | 1.0498 | 1.049 | 1.0437 | 1.0513 | 1.0442 | 1.0506 | 1.0513 | 1.0464 | 1.05 | 1.0496 | 1.0449
80 | Center | 4.18 | 6.76 | 6.22 | 7.24 | 3.75 | 5.02 | 5.39 | 3.75 | 5.12 | 4.81 | 4.2 | 4.23
| CTDIw | 4.36 | 7.06 | 6.49 | 7.46 | 3.91 | 5.17 | 5.62 | 3.91 | 5.27 | 5.02 | 4.38 | 4.35
| k | 1.0431 | 1.0444 | 1.0434 | 1.0304 | 1.0427 | 1.0299 | 1.0427 | 1.0427 | 1.0293 | 1.0437 | 1.0429 | 1.0284
100 | Center | 8.44 | 13.67 | 12.56 | 14.69 | 7.57 | 10.19 | 10.89 | 7.57 | 10.38 | 9.71 | 8.48 | 8.57
| CTDIw | 8.65 | 14.01 | 12.88 | 14.93 | 7.76 | 10.35 | 11.16 | 7.76 | 10.55 | 9.96 | 8.69 | 8.71
| k | 1.0249 | 1.0249 | 1.0255 | 1.0163 | 1.0251 | 1.0157 | 1.0248 | 1.0251 | 1.0164 | 1.0257 | 1.0248 | 1.0163
120 | Center | 13.85 | 22.42 | 20.6 | 24.33 | 12.42 | 16.87 | 17.86 | 12.42 | 17.19 | 15.93 | 13.91 | 14.19
| CTDIw | 14.11 | 22.85 | 21 | 24.63 | 12.66 | 17.08 | 18.2 | 12.66 | 17.41 | 16.24 | 14.17 | 14.37
| k | 1.0188 | 1.0192 | 1.0194 | 1.0123 | 1.0193 | 1.0124 | 1.019 | 1.0193 | 1.0128 | 1.0195 | 1.0187 | 1.0127
140 | Center | 20.24 | 32.77 | 30.12 | 36.16 | 18.16 | 25.07 | 26.11 | 18.16 | 25.55 | 23.3 | 20.33 | 21.09
| CTDIw | 20.6 | 33.35 | 30.65 | 36.64 | 18.48 | 25.4 | 26.56 | 18.48 | 25.89 | 23.7 | 20.69 | 21.37
| k | 1.0178 | 1.0177 | 1.0176 | 1.0133 | 1.0176 | 1.0132 | 1.0172 | 1.0176 | 1.0133 | 1.0172 | 1.0177 | 1.0133
32cm phantom, Standard shaped filter (Adult Body)
| Collimation | 2x1 | 6x0.6 | 8x0.6 | 8x0.6 | 1x5 | 10x0.6 | 6x1.2 | 1x10 | 20x0.6 | 12x1.2 | 32x1.2 | 64x0.6
| Total | 2 | 3.6 | 4.8 | 4.8 | 5 | 6 | 7.2 | 10 | 12 | 14.4 | 38.4 | 38.4
| Acquisition | 2x1 | 6x0.6 | 8x0.6 | 16x0.6 | 1x5 | 20x0.6 | 6x1.2 | 1x10 | 40x0.6 | 12x1.2 | 32x1.2 | 28x0.6
| Shaped filter | Standard | Standard | Standard | Standard | Standard | Standard | Standard | Standard | Standard | Standard | Standard | Standard
kV
70 | Center | 0.61 | 0.99 | 0.91 | 1.07 | 0.55 | 0.74 | 0.79 | 0.55 | 0.75 | 0.7 | 0.61 | 0.67
| CTDIw | 1.13 | 1.83 | 1.68 | 1.97 | 1.01 | 1.37 | 1.46 | 1.01 | 1.4 | 1.3 | 1.14 | 1.15
| k | 1.8525 | 1.8485 | 1.8462 | 1.8411 | 1.8364 | 1.8514 | 1.8481 | 1.8364 | 1.8667 | 1.8571 | 1.8689 | 1.7164
80 | Center | 1.07 | 1.73 | 1.59 | 1.83 | 0.96 | 1.27 | 1.38 | 0.96 | 1.29 | 1.23 | 1.08 | 1.07
| CTDIw | 1.86 | 3.01 | 2.77 | 3.18 | 1.67 | 2.21 | 2.4 | 1.67 | 2.25 | 2.14 | 1.87 | 1.86
| k | 1.7383 | 1.7399 | 1.7421 | 1.7377 | 1.7396 | 1.7402 | 1.7391 | 1.7396 | 1.7442 | 1.7398 | 1.7315 | 1.7383
100 | Center | 2.43 | 3.94 | 3.62 | 4.19 | 2.18 | 2.91 | 3.14 | 2.18 | 2.96 | 2.8 | 2.44 | 2.45
| CTDIw | 3.9 | 6.32 | 5.81 | 6.75 | 3.5 | 4.68 | 5.03 | 3.5 | 4.77 | 4.49 | 3.92 | 3.94
| k | 1.6049 | 1.6041 | 1.605 | 1.611 | 1.6055 | 1.6082 | 1.6019 | 1.6055 | 1.6115 | 1.6036 | 1.6066 | 1.6082
120 | Center | 4.28 | 6.93 | 6.37 | 7.5 | 3.84 | 5.2 | 5.52 | 3.84 | 5.3 | 4.93 | 4.3 | 4.37
| CTDIw | 6.61 | 10.7 | 9.84 | 11.55 | 5.93 | 8.01 | 8.53 | 5.93 | 8.16 | 7.61 | 6.64 | 6.74
| k | 1.5444 | 1.544 | 1.5447 | 1.54 | 1.5443 | 1.5404 | 1.5453 | 1.5443 | 1.5396 | 1.5436 | 1.5442 | 1.5423
140 | Center | 6.57 | 10.64 | 9.78 | 11.59 | 5.89 | 8.03 | 8.47 | 5.89 | 8.19 | 7.56 | 6.6 | 6.76
| CTDIw | 9.94 | 16.09 | 14.79 | 17.62 | 8.91 | 12.22 | 12.81 | 8.91 | 12.45 | 11.43 | 9.98 | 10.28
| k | 1.5129 | 1.5122 | 1.5123 | 1.5203 | 1.5127 | 1.5218 | 1.5124 | 1.5127 | 1.5201 | 1.5119 | 1.5121 | 1.5207
16cm phantom, Narrow shaped filter (Pediatric Body)
| Collimation | 2x1 | 6x0.6 | 8x0.6 | 8x0.6 | 1x5 | 10x0.6 | 6x1.2 | 1x10 | 20x0.6 | 12x1.2 | 32x1.2 | 64x0.6
| Total | 2 | 3.6 | 4.8 | 4.8 | 5 | 6 | 7.2 | 10 | 12 | 14.4 | 38.4 | 38.4
| Acquisition | 2x1 | 6x0.6 | 8x0.6 | 16x0.6 | 1x5 | 20x0.6 | 6x1.2 | 1x10 | 40x0.6 | 12x1.2 | 32x1.2 | 28x0.6
| Shaped filter | Narrow | Narrow | Narrow | Narrow | Narrow | Narrow | Narrow | Narrow | Narrow | Narrow | Narrow | Narrow
kV
70 | Center | 2.36 | 3.83 | 3.52 | 4.12 | 2.12 | 2.86 | 3.05 | 2.12 | 2.91 | 2.72 | 2.38 | 2.4
| CTDIw | 2.36 | 3.81 | 3.51 | 4.07 | 2.11 | 2.82 | 3.04 | 2.11 | 2.88 | 2.71 | 2.37 | 2.37
| k | 1 | 0.99478 | 0.99716 | 0.98786 | 0.99528 | 0.98601 | 0.99672 | 0.99528 | 0.98969 | 0.99632 | 0.9958 | 0.9875
80 | Center | 3.81 | 6.17 | 5.67 | 6.58 | 3.42 | 4.56 | 4.92 | 3.42 | 4.65 | 4.39 | 3.83 | 3.84
| CTDIw | 3.78 | 6.12 | 5.63 | 6.44 | 3.39 | 4.47 | 4.88 | 3.39 | 4.55 | 4.39 | 3.8 | 3.76
| k | 0.99213 | 0.9919 | 0.99295 | 0.97872 | 0.99123 | 0.98026 | 0.99187 | 0.99123 | 0.97849 | 1 | 0.99217 | 0.97917
100 | Center | 7.81 | 12.65 | 11.63 | 13.53 | 7.01 | 9.38 | 10.08 | 7.01 | 9.56 | 8.99 | 7.85 | 7.89
| CTDIw | 7.69 | 12.45 | 11.44 | 13.14 | 6.9 | 9.11 | 9.92 | 6.9 | 9.29 | 8.85 | 7.72 | 7.67
| k | 0.98464 | 0.98419 | 0.98366 | 0.97118 | 0.98431 | 0.97122 | 0.98413 | 0.98431 | 0.97176 | 0.98443 | 0.98344 | 0.97212
120 | Center | 12.93 | 20.93 | 19.23 | 22.53 | 11.6 | 15.62 | 16.67 | 11.6 | 15.92 | 14.87 | 12.98 | 13.14
| CTDIw | 12.68 | 20.52 | 18.86 | 21.91 | 11.37 | 16.19 | 16.35 | 11.37 | 15.48 | 14.59 | 12.73 | 12.78
| k | 0.98067 | 0.98041 | 0.98076 | 0.97248 | 0.98017 | 1.0365 | 0.9808 | 0.98017 | 0.97236 | 0.98117 | 0.98074 | 0.9726
140 | Center | 18.83 | 30.49 | 28.02 | 33.71 | 16.89 | 23.38 | 24.29 | 16.89 | 23.82 | 21.67 | 18.91 | 19.67
| CTDIw | 18.5 | 29.95 | 27.53 | 32.85 | 16.6 | 22.77 | 23.86 | 16.6 | 23.21 | 21.29 | 18.58 | 19.16
| k | 0.98247 | 0.98229 | 0.98251 | 0.97449 | 0.98283 | 0.97391 | 0.9823 | 0.98283 | 0.97439 | 0.98246 | 0.98255 | 0.97407
32cm phantom, Narrow shaped filter (Cardiac)
| Collimation | 2x1 | 6x0.6 | 8x0.6 | 8x0.6 | 1x5 | 10x0.6 | 6x1.2 | 1x10 | 20x0.6 | 12x1.2 | 32x1.2 | 64x0.6
| Total | 2 | 3.6 | 4.8 | 4.8 | 5 | 6 | 7.2 | 10 | 12 | 14.4 | 38.4 | 38.4
| Acquisition | 2x1 | 6x0.6 | 8x0.6 | 16x0.6 | 1x5 | 20x0.6 | 6x1.2 | 1x10 | 40x0.6 | 12x1.2 | 32x1.2 | 28x0.6
| Shaped filter | Narrow | Narrow | Narrow | Narrow | Narrow | Narrow | Narrow | Narrow | Narrow | Narrow | Narrow | Narrow
kV
70 | Center | 0.54 | 0.87 | 0.8 | 0.93 | 0.48 | 0.65 | 0.69 | 0.48 | 0.66 | 0.62 | 0.54 | 0.55
| CTDIw | 0.9 | 1.45 | 1.33 | 1.54 | 0.8 | 1.07 | 1.16 | 0.8 | 1.09 | 1.03 | 0.9 | 0.9
| k | 1.6667 | 1.6667 | 1.6625 | 1.6559 | 1.6667 | 1.6462 | 1.6812 | 1.6667 | 1.6515 | 1.6613 | 1.6667 | 1.6364
80 | Center | 0.94 | 1.52 | 1.39 | 1.62 | 0.84 | 1.12 | 1.21 | 0.84 | 1.15 | 1.08 | 0.94 | 0.95
| CTDIw | 1.46 | 2.36 | 2.17 | 2.48 | 1.31 | 1.72 | 1.88 | 1.31 | 1.76 | 1.68 | 1.47 | 1.45
| k | 1.5532 | 1.5526 | 1.5612 | 1.5309 | 1.5595 | 1.5357 | 1.5537 | 1.5595 | 1.5304 | 1.5556 | 1.5638 | 1.5263
100 | Center | 2.17 | 3.51 | 3.22 | 3.75 | 1.94 | 2.6 | 2.79 | 1.94 | 2.65 | 2.49 | 2.17 | 2.19
| CTDIw | 3.12 | 5.06 | 4.65 | 5.34 | 2.8 | 3.7 | 4.03 | 2.8 | 3.78 | 3.59 | 3.14 | 3.12
| k | 1.4378 | 1.4416 | 1.4441 | 1.424 | 1.4433 | 1.4231 | 1.4444 | 1.4433 | 1.4264 | 1.4418 | 1.447 | 1.4247
120 | Center | 3.84 | 6.22 | 5.72 | 6.67 | 3.45 | 4.62 | 4.95 | 3.45 | 4.71 | 4.42 | 3.86 | 3.89
| CTDIw | 5.36 | 8.68 | 7.98 | 9.21 | 4.81 | 6.39 | 6.91 | 4.81 | 6.51 | 6.17 | 5.39 | 5.38
| k | 1.3958 | 1.3955 | 1.3951 | 1.3808 | 1.3942 | 1.3831 | 1.396 | 1.3942 | 1.3822 | 1.3959 | 1.3964 | 1.383
140 | Center | 5.89 | 9.53 | 8.76 | 10.45 | 5.28 | 7.25 | 7.59 | 5.28 | 7.38 | 6.77 | 5.91 | 6.1
| CTDIw | 8.03 | 13 | 11.95 | 14.2 | 7.2 | 9.84 | 10.36 | 7.2 | 10.03 | 9.24 | 8.07 | 8.28
| k | 1.3633 | 1.3641 | 1.3642 | 1.3589 | 1.3636 | 1.3572 | 1.365 | 1.3636 | 1.3591 | 1.3648 | 1.3655 | 1.3574

## Somatom Drive
16 cm phantom, Standard shaped filter (Routine head)
| Collimation | 2x1 | 6x0.6 | 8x0.6 | 1x5 | 10x0.6 | 1x10 | 12x1.2 | 20x0.6 | 32x0.6 | 12x1.2 | 64x0.6
| Total | 2 | 3.6 | 4.8 | 5 | 6 | 10 | 14.4 | 12 | 19.2 | 14.4 | 38.4
| Acquisition | 2x1 | 6x0.6 | 16x0.6 | 1x5 | 20x0.6 | 1x10 | 12x1.2 | 40x0.6 | 64x0.6 | 12x1.2 | 128x0.6
| Shaped filter |  |  | z-sharp |  | z-sharp |  |  | z-sharp | z-sharp
kV
70 | Center | 2.57 | 4.17 | 4.39 | 2.31 | 3.04 | 2.31 | 2.96 | 3.1 | 2.86 | 2.59 | 2.52
| CTDIw | 2.69 | 4.35 | 4.56 | 2.41 | 3.16 | 2.41 | 3.09 | 3.22 | 2.98 | 2.7 | 2.62
| k | 1.0467 | 1.0432 | 1.0387 | 1.0433 | 1.0395 | 1.0433 | 1.0439 | 1.0387 | 1.042 | 1.0425 | 1.0397
80 | Center | 4.12 | 6.66 | 7.03 | 3.69 | 4.87 | 3.69 | 4.74 | 4.97 | 4.58 | 4.13 | 4.04
| CTDIw | 4.24 | 6.87 | 7.21 | 3.8 | 5 | 3.8 | 4.88 | 5.1 | 4.71 | 4.26 | 4.15
| k | 1.0291 | 1.0315 | 1.0256 | 1.0298 | 1.0267 | 1.0298 | 1.0295 | 1.0262 | 1.0284 | 1.0315 | 1.0272
90 | Center | 6.06 | 9.81 | 10.4 | 5.44 | 7.19 | 5.44 | 6.97 | 7.33 | 6.75 | 6.09 | 5.97
| CTDIw | 6.2 | 10 | 10.6 | 5.56 | 7.32 | 5.56 | 7.13 | 7.46 | 6.89 | 6.22 | 6.07
| k | 1.0231 | 1.0194 | 1.0192 | 1.0221 | 1.0181 | 1.0221 | 1.023 | 1.0177 | 1.0207 | 1.0213 | 1.0168
100 | Center | 8.35 | 13.5 | 14.3 | 7.49 | 9.91 | 7.49 | 9.6 | 10.1 | 9.3 | 8.38 | 8.22
| CTDIw | 8.49 | 13.8 | 14.5 | 7.62 | 10 | 7.62 | 9.77 | 10.2 | 9.44 | 8.53 | 8.32
| k | 1.0168 | 1.0222 | 1.014 | 1.0174 | 1.0091 | 1.0174 | 1.0177 | 1.0099 | 1.0151 | 1.0179 | 1.0122
110 | Center | 11.2 | 18.1 | 19.2 | 10.1 | 13.3 | 10.1 | 12.9 | 13.5 | 12.5 | 11.3 | 11
| CTDIw | 11.4 | 18.4 | 19.4 | 10.2 | 13.4 | 10.2 | 13.1 | 13.7 | 12.6 | 11.4 | 11.1
| k | 1.0179 | 1.0166 | 1.0104 | 1.0099 | 1.0075 | 1.0099 | 1.0155 | 1.0148 | 1.008 | 1.0088 | 1.0091
120 | Center | 14 | 22.7 | 23.9 | 12.6 | 16.6 | 12.6 | 16.1 | 16.9 | 15.6 | 14.1 | 13.8
| CTDIw | 14.2 | 23 | 24.1 | 12.7 | 16.7 | 12.7 | 16.3 | 17.1 | 15.8 | 14.2 | 13.9
| k | 1.0143 | 1.0132 | 1.0084 | 1.0079 | 1.006 | 1.0079 | 1.0124 | 1.0118 | 1.0128 | 1.0071 | 1.0072
130 | Center | 16.9 | 27.4 | 29 | 15.2 | 20.1 | 15.2 | 19.5 | 20.5 | 18.9 | 17 | 16.7
| CTDIw | 17.2 | 27.8 | 29.3 | 15.4 | 20.3 | 15.4 | 19.8 | 20.7 | 19.1 | 17.3 | 16.8
| k | 1.0178 | 1.0146 | 1.0103 | 1.0132 | 1.01 | 1.0132 | 1.0154 | 1.0098 | 1.0106 | 1.0176 | 1.006
140 | Center | 20.3 | 32.9 | 34.8 | 18.2 | 24.1 | 18.2 | 23.4 | 24.6 | 22.6 | 20.4 | 20
| CTDIw | 20.5 | 33.3 | 35.1 | 18.4 | 24.3 | 18.4 | 23.6 | 24.8 | 22.9 | 20.9 | 20.2
| k | 1.0099 | 1.0122 | 1.0086 | 1.011 | 1.0083 | 1.011 | 1.0085 | 1.0081 | 1.0133 | 1.0245 | 1.01
Sn 100 | Center | 1.54 | 2.5 | 2.65 | 1.38 | 1.84 | 1.38 | 1.77 | 1.87 | 1.72 | 1.55 | 1.52
| CTDIw | 1.54 | 2.49 | 2.63 | 1.38 | 1.82 | 1.38 | 1.77 | 1.86 | 1.71 | 1.54 | 1.51
| k | 1 | 0.996 | 0.99245 | 1 | 0.98913 | 1 | 1 | 0.99465 | 0.99419 | 0.99355 | 0.99342
Sn 140 | Center | 6.49 | 10.5 | 11.2 | 5.82 | 7.74 | 5.82 | 7.47 | 7.88 | 7.25 | 6.52 | 6.42
| CTDIw | 6.55 | 10.6 | 11.2 | 5.87 | 7.79 | 5.87 | 7.54 | 7.94 | 7.3 | 6.58 | 6.46
| k | 1.0092 | 1.0095 | 1 | 1.0086 | 1.0065 | 1.0086 | 1.0094 | 1.0076 | 1.0069 | 1.0092 | 1.0062
16 cm phantom, Narrow shaped filter (Inner Ear)
| Collimation | 2x1 | 6x0.6 | 8x0.6 | 1x5 | 10x0.6 | 1x10 | 12x1.2 | 20x0.6 | 32x0.6 | 12x1.2 | 64x0.6
| Total | 2 | 3.6 | 4.8 | 5 | 6 | 10 | 14.4 | 12 | 19.2 | 14.4 | 38.4
| Acquisition | 2x1 | 6x0.6 | 16x0.6 | 1x5 | 20x0.6 | 1x10 | 12x1.2 | 40x0.6 | 64x0.6 | 12x1.2 | 128x0.6
| Shaped filter |  |  | z-sharp |  | z-sharp |  |  | z-sharp | z-sharp
kV
70 | Center | 2.32 | 3.75 | 3.95 | 2.08 | 2.74 | 2.08 | 2.67 | 2.79 | 2.58 | 2.33 | 2.27
| CTDIw | 2.29 | 3.71 | 3.88 | 20.6 | 2.69 | 20.6 | 2.64 | 2.74 | 2.54 | 2.3 | 2.23
| k | 0.98707 | 0.98933 | 0.98228 | 9.9038 | 0.98175 | 9.9038 | 0.98876 | 0.98208 | 0.9845 | 0.98712 | 0.98238
80 | Center | 3.74 | 6.05 | 6.38 | 3.35 | 4.43 | 3.35 | 4.3 | 4.51 | 4.16 | 3.76 | 3.67
| CTDIw | 3.67 | 5.94 | 6.23 | 3.29 | 4.32 | 3.29 | 4.22 | 4.4 | 4.07 | 3.68 | 3.58
| k | 0.98128 | 0.98182 | 0.97649 | 0.98209 | 0.97517 | 0.98209 | 0.9814 | 0.97561 | 0.97837 | 0.97872 | 0.97548
90 | Center | 5.54 | 8.97 | 9.47 | 4.97 | 6.57 | 4.97 | 6.38 | 6.69 | 6.17 | 5.56 | 5.45
| CTDIw | 5.42 | 8.77 | 9.2 | 4.86 | 6.38 | 4.86 | 6.24 | 6.5 | 6.01 | 5.44 | 5.29
| k | 0.97834 | 0.9777 | 0.97149 | 0.97787 | 0.97108 | 0.97787 | 0.97806 | 0.9716 | 0.97407 | 0.97842 | 0.97064
100 | Center | 7.66 | 12.4 | 13.1 | 6.87 | 9.08 | 6.87 | 8.81 | 9.25 | 8.53 | 7.69 | 7.53
| CTDIw | 7.48 | 12.1 | 12.7 | 6.71 | 8.81 | 6.71 | 8.6 | 8.98 | 8.3 | 7.51 | 7.31
| k | 0.9765 | 0.97581 | 0.96947 | 0.97671 | 0.97026 | 0.97671 | 0.97616 | 0.97081 | 0.97304 | 0.97659 | 0.97078
110 | Center | 10.3 | 16.7 | 17.6 | 9.26 | 12.2 | 9.26 | 11.9 | 12.5 | 11.5 | 10.4 | 10.1
| CTDIw | 10.1 | 16.3 | 17.1 | 9.03 | 11.9 | 9.03 | 11.6 | 12.1 | 11.2 | 10.1 | 9.83
| k | 0.98058 | 0.97605 | 0.97159 | 0.97516 | 0.97541 | 0.97516 | 0.97479 | 0.968 | 0.97391 | 0.97115 | 0.97327
120 | Center | 12.9 | 20.9 | 22.1 | 11.6 | 15.3 | 11.6 | 14.9 | 15.6 | 14.4 | 13 | 12.7
| CTDIw | 12.6 | 20.4 | 21.4 | 11.3 | 14.9 | 11.3 | 14.5 | 15.1 | 14 | 12.7 | 12.3
| k | 0.97674 | 0.97608 | 0.96833 | 0.97414 | 0.97386 | 0.97414 | 0.97315 | 0.96795 | 0.97222 | 0.97692 | 0.9685
130 | Center | 15.7 | 25.4 | 26.8 | 14.1 | 18.6 | 14.1 | 18.1 | 19 | 17.5 | 15.8 | 15.4
| CTDIw | 15.4 | 24.9 | 26.1 | 13.8 | 18.1 | 13.8 | 17.7 | 18.5 | 17.1 | 15.4 | 15
| k | 0.98089 | 0.98031 | 0.97388 | 0.97872 | 0.97312 | 0.97872 | 0.9779 | 0.97368 | 0.97714 | 0.97468 | 0.97403
140 | Center | 18.8 | 30.5 | 32.3 | 16.9 | 22.4 | 16.9 | 21.7 | 22.8 | 21 | 18.9 | 18.5
| CTDIw | 18.4 | 29.8 | 31.4 | 16.5 | 21.8 | 16.5 | 21.2 | 22.2 | 20.5 | 18.5 | 18
| k | 0.97872 | 0.97705 | 0.97214 | 0.97633 | 0.97321 | 0.97633 | 0.97696 | 0.97368 | 0.97619 | 0.97884 | 0.97297
Sn 100 | Center | 1.45 | 2.35 | 2.49 | 1.3 | 1.73 | 1.3 | 1.67 | 1.76 | 1.62 | 1.46 | 1.43
| CTDIw | 1.4 | 2.27 | 2.4 | 1.26 | 1.66 | 1.26 | 1.61 | 1.69 | 1.56 | 1.41 | 1.38
| k | 0.96552 | 0.96596 | 0.96386 | 0.96923 | 0.95954 | 0.96923 | 0.96407 | 0.96023 | 0.96296 | 0.96575 | 0.96503
Sn 140 | Center | 6.16 | 9.97 | 10.6 | 5.52 | 7.35 | 5.52 | 7.09 | 7.49 | 6.88 | 6.19 | 6.09
| CTDIw | 6.06 | 9.8 | 10.4 | 5.43 | 7.2 | 5.43 | 6.97 | 7.34 | 6.75 | 6.08 | 5.97
| k | 0.98377 | 0.98295 | 0.98113 | 0.9837 | 0.97959 | 0.9837 | 0.98307 | 0.97997 | 0.9811 | 0.98223 | 0.9803
32 cm phantom, Narrow shaped filter (Cardiac, Ped Body, Flash)
| Collimation | 2x1 | 6x0.6 | 8x0.6 | 1x5 | 10x0.6 | 1x10 | 12x1.2 | 20x0.6 | 32x0.6 | 12x1.2 | 64x0.6
| Total | 2 | 3.6 | 4.8 | 5 | 6 | 10 | 14.4 | 12 | 19.2 | 14.4 | 38.4
| Acquisition | 2x1 | 6x0.6 | 16x0.6 | 1x5 | 20x0.6 | 1x10 | 12x1.2 | 40x0.6 | 64x0.6 | 12x1.2 | 128x0.6
| Shaped filter |  |  | z-sharp |  | z-sharp |  |  | z-sharp | z-sharp
kV
70 | Center | 0.5 | 0.81 | 0.86 | 0.45 | 0.6 | 0.45 | 0.58 | 0.61 | 0.56 | 0.5 | 0.49
| CTDIw | 0.82 | 1.33 | 1.39 | 0.74 | 0.97 | 0.74 | 0.94 | 0.98 | 0.91 | 0.82 | 0.8
| k | 1.64 | 1.642 | 1.6163 | 1.6444 | 1.6167 | 1.6444 | 1.6207 | 1.6066 | 1.625 | 1.64 | 1.6327
80 | Center | 0.91 | 1.48 | 1.56 | 0.82 | 1.08 | 0.82 | 1.05 | 1.11 | 1.02 | 0.92 | 0.9
| CTDIw | 1.41 | 2.28 | 2.38 | 1.26 | 1.65 | 1.26 | 1.62 | 1.68 | 1.56 | 1.41 | 1.37
| k | 1.5495 | 1.5405 | 1.5256 | 1.5366 | 1.5278 | 1.5366 | 1.5429 | 1.5135 | 1.5294 | 1.5326 | 1.5222
90 | Center | 1.46 | 2.36 | 2.49 | 1.31 | 1.73 | 1.31 | 1.68 | 1.76 | 1.62 | 1.47 | 1.43
| CTDIw | 2.17 | 3.51 | 3.67 | 1.95 | 2.54 | 1.95 | 2.5 | 2.59 | 2.4 | 2.18 | 2.11
| k | 1.4863 | 1.4873 | 1.4739 | 1.4885 | 1.4682 | 1.4885 | 1.4881 | 1.4716 | 1.4815 | 1.483 | 1.4755
100 | Center | 2.1 | 3.4 | 3.57 | 1.89 | 2.49 | 1.89 | 2.42 | 2.54 | 2.34 | 2.11 | 2.07
| CTDIw | 3.04 | 4.93 | 5.15 | 2.73 | 3.57 | 2.73 | 3.5 | 3.64 | 3.37 | 3.06 | 2.96
| k | 1.4476 | 1.45 | 1.4426 | 1.4444 | 1.4337 | 1.4444 | 1.4463 | 1.4331 | 1.4402 | 1.4502 | 1.43
110 | Center | 2.91 | 4.72 | 4.98 | 2.61 | 3.45 | 2.61 | 3.35 | 3.52 | 3.24 | 2.93 | 2.86
| CTDIw | 4.14 | 6.7 | 7.01 | 3.71 | 7.86 | 3.71 | 4.76 | 4.95 | 4.55 | 4.1 | 4.03
| k | 1.4227 | 1.4195 | 1.4076 | 1.4215 | 2.2783 | 1.4215 | 1.4209 | 1.4062 | 1.4043 | 1.3993 | 1.4091
120 | Center | 3.76 | 6.09 | 6.42 | 3.37 | 4.45 | 3.37 | 4.33 | 4.54 | 4.18 | 3.78 | 3.67
| CTDIw | 5.27 | 8.53 | 8.92 | 4.73 | 6.18 | 4.73 | 6.06 | 6.3 | 5.84 | 5.29 | 5.13
| k | 1.4016 | 1.4007 | 1.3894 | 1.4036 | 1.3888 | 1.4036 | 1.3995 | 1.3877 | 1.3971 | 1.3995 | 1.3978
130 | Center | 4.73 | 7.65 | 8.08 | 4.24 | 5.6 | 4.24 | 5.44 | 5.71 | 5.26 | 4.75 | 4.65
| CTDIw | 6.56 | 10.6 | 11.1 | 5.88 | 7.71 | 5.88 | 7.55 | 7.86 | 7.27 | 6.59 | 6.39
| k | 1.3869 | 1.3856 | 1.3738 | 1.3868 | 1.3768 | 1.3868 | 1.3879 | 1.3765 | 1.3821 | 1.3874 | 1.3742
140 | Center | 5.81 | 9.41 | 9.96 | 5.21 | 6.9 | 5.21 | 6.69 | 7.04 | 6.48 | 5.84 | 5.72
| CTDIw | 8 | 13 | 13.6 | 7.18 | 9.4 | 7.18 | 9.21 | 9.58 | 8.87 | 8.04 | 7.8
| k | 1.3769 | 1.3815 | 1.3655 | 1.3781 | 1.3623 | 1.3781 | 1.3767 | 1.3608 | 1.3688 | 1.3767 | 1.3636
Sn 100 | Center | 0.48 | 0.78 | 0.82 | 0.43 | 0.57 | 0.43 | 0.55 | 0.58 | 0.54 | 0.48 | 0.47
| CTDIw | 0.63 | 1.02 | 1.08 | 0.57 | 0.75 | 0.57 | 0.73 | 0.76 | 0.7 | 0.64 | 0.62
| k | 1.3125 | 1.3077 | 1.3171 | 1.3256 | 1.3158 | 1.3256 | 1.3273 | 1.3103 | 1.2963 | 1.3333 | 1.3191
Sn 140 | Center | 2.23 | 3.61 | 3.83 | 2 | 2.66 | 2 | 2.57 | 2.71 | 2.49 | 2.24 | 2.2
| CTDIw | 2.9 | 4.7 | 4.94 | 2.6 | 3.43 | 2.6 | 3.34 | 3.49 | 3.22 | 2.92 | 2.84
| k | 1.3004 | 1.3019 | 1.2898 | 1.3 | 1.2895 | 1.3 | 1.2996 | 1.2878 | 1.2932 | 1.3036 | 1.2909
32 cm phantom, Standard shaped filter (Adult Body)
| Collimation | 2x1 | 6x0.6 | 8x0.6 | 1x5 | 10x0.6 | 1x10 | 12x1.2 | 20x0.6 | 32x0.6 | 12x1.2 | 64x0.6
| Total | 2 | 3.6 | 4.8 | 5 | 6 | 10 | 14.4 | 12 | 19.2 | 14.4 | 38.4
| Acquisition | 2x1 | 6x0.6 | 16x0.6 | 1x5 | 20x0.6 | 1x10 | 12x1.2 | 40x0.6 | 64x0.6 | 12x1.2 | 128x0.6
| Shaped filter |  |  | z-sharp |  | z-sharp |  |  | z-sharp | z-sharp
kV
70 | Center | 0.58 | 0.94 | 0.99 | 0.52 | 0.69 | 0.52 | 0.66 | 0.7 | 0.64 | 0.58 | 0.57
| CTDIw | 1.05 | 1.7 | 1.79 | 0.94 | 1.24 | 0.94 | 1.21 | 1.27 | 1.17 | 1.05 | 1.03
| k | 1.8103 | 1.8085 | 1.8081 | 1.8077 | 1.7971 | 1.8077 | 1.8333 | 1.8143 | 1.8281 | 1.8103 | 1.807
80 | Center | 1.04 | 1.69 | 1.79 | 0.94 | 1.24 | 0.94 | 1.2 | 1.26 | 1.16 | 1.05 | 1.03
| CTDIw | 1.79 | 2.9 | 3.04 | 1.6 | 2.11 | 1.6 | 2.06 | 2.15 | 1.99 | 1.8 | 1.75
| k | 1.7212 | 1.716 | 1.6983 | 1.7021 | 1.7016 | 1.7021 | 1.7167 | 1.7063 | 1.7155 | 1.7143 | 1.699
90 | Center | 1.66 | 2.68 | 2.84 | 1.49 | 1.97 | 1.49 | 1.91 | 2 | 1.85 | 1.66 | 1.63
| CTDIw | 2.73 | 4.42 | 4.64 | 2.45 | 3.22 | 2.45 | 3.14 | 3.28 | 3.03 | 2.74 | 2.67
| k | 1.6446 | 1.6493 | 1.6338 | 1.6443 | 1.6345 | 1.6443 | 1.644 | 1.64 | 1.6378 | 1.6506 | 1.638
100 | Center | 2.38 | 3.85 | 4.07 | 2.13 | 2.82 | 2.13 | 2.73 | 2.88 | 2.65 | 2.39 | 2.34
| CTDIw | 3.81 | 6.17 | 6.48 | 3.42 | 4.5 | 3.42 | 4.39 | 4.58 | 4.23 | 3.83 | 3.73
| k | 1.6008 | 1.6026 | 1.5921 | 1.6056 | 1.5957 | 1.6056 | 1.6081 | 1.5903 | 1.5962 | 1.6025 | 1.594
110 | Center | 3.28 | 5.31 | 5.62 | 2.94 | 3.9 | 2.94 | 3.77 | 3.97 | 3.66 | 3.29 | 3.23
| CTDIw | 5.16 | 8.36 | 8.77 | 4.63 | 6.08 | 4.63 | 5.94 | 6.2 | 5.73 | 5.18 | 5.04
| k | 1.5732 | 1.5744 | 1.5605 | 1.5748 | 1.559 | 1.5748 | 1.5756 | 1.5617 | 1.5656 | 1.5745 | 1.5604
120 | Center | 4.22 | 6.84 | 7.23 | 3.79 | 5.01 | 3.79 | 4.86 | 5.11 | 4.7 | 4.24 | 4.16
| CTDIw | 6.53 | 10.6 | 11.1 | 5.86 | 7.7 | 5.86 | 7.52 | 7.85 | 7.25 | 6.56 | 6.39
| k | 1.5474 | 1.5497 | 1.5353 | 1.5462 | 1.5369 | 1.5462 | 1.5473 | 1.5362 | 1.5426 | 1.5472 | 1.5361
130 | Center | 5.29 | 8.57 | 9.07 | 4.75 | 6.29 | 4.75 | 6.09 | 6.41 | 5.9 | 5.31 | 5.22
| CTDIw | 8.1 | 13.1 | 13.8 | 7.26 | 9.55 | 7.26 | 9.32 | 9.74 | 8.99 | 8.13 | 7.92
| k | 1.5312 | 1.5286 | 1.5215 | 1.5284 | 1.5183 | 1.5284 | 1.5304 | 1.5195 | 1.5237 | 1.5311 | 1.5172
140 | Center | 6.49 | 10.5 | 11.1 | 5.82 | 7.72 | 5.82 | 7.47 | 7.86 | 7.24 | 6.52 | 6.4
| CTDIw | 9.84 | 15.9 | 16.7 | 8.83 | 11.6 | 8.83 | 11.3 | 11.8 | 10.9 | 9.89 | 9.63
| k | 1.5162 | 1.5143 | 1.5045 | 1.5172 | 1.5026 | 1.5172 | 1.5127 | 1.5013 | 1.5055 | 1.5169 | 1.5047
Sn 100 | Center | 0.54 | 0.87 | 0.92 | 0.48 | 0.64 | 0.48 | 0.62 | 0.65 | 0.6 | 0.54 | 0.53
| CTDIw | 0.77 | 1.25 | 1.32 | 0.69 | 0.91 | 0.69 | 0.89 | 0.93 | 0.86 | 0.78 | 0.76
| k | 1.4259 | 1.4368 | 1.4348 | 1.4375 | 1.4219 | 1.4375 | 1.4355 | 1.4308 | 1.4333 | 1.4444 | 1.434
Sn 140 | Center | 2.45 | 3.97 | 4.21 | 2.2 | 2.92 | 2.2 | 2.82 | 2.97 | 2.73 | 2.46 | 2.42
| CTDIw | 3.47 | 5.62 | 5.94 | 3.12 | 4.12 | 3.12 | 4 | 4.2 | 3.87 | 3.49 | 3.42
| k | 1.4163 | 1.4156 | 1.4109 | 1.4182 | 1.411 | 1.4182 | 1.4184 | 1.4141 | 1.4176 | 1.4187 | 1.4132

## Somatom Force
16 cm phantom, Standard shaped filter (Routine head)
| Collimation | 2x1 | 6x0.6 | 1x5 | 1x10 | 12x1.2 | 32x0.6 | 32x0.6 | 32x0.6 | 48x1.2 | 64x1.2 | 96x0.6 | 24x1.2 | 48x1.2
| Total | 2 | 3.6 | 5 | 10 | 14.4 | 19.2 | 19.2 | 19.2 | 57.6 | 76.8 | 57.6 | 28.8 | 57.6
| Acquisition | 2x1 | 6x0.6 | 1x5 | 1x10 | 12x1.2 | 32x0.6 | 64x0.6 | 64x0.6 | 96x0.6 | 128x0.6 | 192x0.6 | 24x1.2 | 48*1.2
| Shaped filter | [5 empty] | UHR | sURH z-sharp | z-sharp | z-sharp | z-sharp | z-sharp | z-sharp
kV
70 | Center | 2.53 | 4.21 | 2.44 | 2.44 | 3.48 | 2.75 | 2.96 | 3.36 | 3.05 | 2.9 | 2.72 | 3 | 2.73
| CTDIw | 2.72 | 4.53 | 2.63 | 2.63 | 3.75 | 2.96 | 3.18 | 3.62 | 3.28 | 3.12 | 2.93 | 3.23 | 2.93
| k | 1.0751 | 1.076 | 1.0779 | 1.0779 | 1.0776 | 1.0764 | 1.0743 | 1.0774 | 1.0754 | 1.0759 | 1.0772 | 1.0767 | 1.0733
80 | Center | 4.04 | 6.73 | 3.9 | 3.9 | 5.56 | 4.4 | 4.72 | 5.37 | 4.87 | 4.62 | 4.35 | 4.79 | 4.35
| CTDIw | 4.24 | 7.07 | 4.1 | 4.1 | 5.85 | 4.63 | 4.97 | 5.65 | 5.12 | 4.86 | 4.58 | 5.03 | 4.58
| k | 1.0495 | 1.0505 | 1.0513 | 1.0513 | 1.0522 | 1.0523 | 1.053 | 1.0521 | 1.0513 | 1.0519 | 1.0529 | 1.0501 | 1.0529
90 | Center | 5.93 | 9.88 | 5.73 | 5.73 | 8.17 | 6.46 | 6.94 | 7.89 | 7.16 | 6.79 | 6.39 | 7.03 | 6.4
| CTDIw | 6.08 | 10.13 | 5.87 | 5.87 | 8.37 | 6.62 | 7.11 | 8.08 | 7.33 | 6.96 | 6.55 | 7.2 | 6.55
| k | 1.0253 | 1.0253 | 1.0244 | 1.0244 | 1.0245 | 1.0248 | 1.0245 | 1.0241 | 1.0237 | 1.025 | 1.025 | 1.0242 | 1.0234
100 | Center | 8.08 | 13.47 | 7.81 | 7.81 | 11.14 | 8.81 | 9.45 | 10.75 | 9.76 | 9.26 | 8.71 | 9.58 | 8.72
| CTDIw | 8.23 | 13.72 | 7.95 | 7.95 | 11.34 | 8.97 | 9.63 | 10.95 | 9.94 | 9.43 | 8.87 | 9.76 | 8.88
| k | 1.0186 | 1.0186 | 1.0179 | 1.0179 | 1.018 | 1.0182 | 1.019 | 1.0186 | 1.0184 | 1.0184 | 1.0184 | 1.0188 | 1.0183
110 | Center | 10.45 | 17.42 | 10.1 | 10.1 | 14.41 | 11.39 | 12.23 | 13.91 | 12.62 | 11.98 | 11.27 | 12.4 | 11.28
| CTDIw | 10.61 | 17.68 | 10.25 | 10.25 | 14.62 | 11.56 | 12.41 | 14.11 | 12.81 | 12.16 | 11.44 | 12.58 | 11.45
| k | 1.0153 | 1.0149 | 1.0149 | 1.0149 | 1.0146 | 1.0149 | 1.0147 | 1.0144 | 1.0151 | 1.015 | 1.0151 | 1.0145 | 1.0151
120 | Center | 13.06 | 21.76 | 12.62 | 12.62 | 18 | 14.23 | 15.28 | 17.37 | 15.76 | 14.96 | 14.07 | 15.49 | 14.09
| CTDIw | 13.22 | 22.04 | 12.77 | 12.77 | 18.22 | 14.41 | 15.47 | 17.59 | 15.96 | 15.15 | 14.25 | 15.68 | 14.26
| k | 1.0123 | 1.0129 | 1.0119 | 1.0119 | 1.0122 | 1.0126 | 1.0124 | 1.0127 | 1.0127 | 1.0127 | 1.0128 | 1.0123 | 1.0121
130 | Center | 15.73 | 26.21 | 15.2 | 15.2 | 21.68 | 17.14 | 18.4 | 20.92 | 18.99 | 18.02 | 16.95 | 18.65 | 16.97
| CTDIw | 15.92 | 26.53 | 15.38 | 15.38 | 21.93 | 17.34 | 18.62 | 21.17 | 19.21 | 18.24 | 17.15 | 18.87 | 17.17
| k | 1.0121 | 1.0122 | 1.0118 | 1.0118 | 1.0115 | 1.0117 | 1.012 | 1.012 | 1.0116 | 1.0122 | 1.0118 | 1.0118 | 1.0118
140 | Center | 18.55 | 30.92 | 17.93 | 17.93 | 25.57 | 20.22 | 21.7 | 24.68 | 22.4 | 21.26 | 20 | 22 | 20.01
| CTDIw | 18.75 | 31.25 | 18.12 | 18.12 | 25.84 | 20.43 | 21.94 | 24.94 | 22.64 | 21.49 | 20.21 | 22.24 | 20.23
| k | 1.0108 | 1.0107 | 1.0106 | 1.0106 | 1.0106 | 1.0104 | 1.0111 | 1.0105 | 1.0107 | 1.0108 | 1.0105 | 1.0109 | 1.011
150 | Center | 21.49 | 35.82 | 20.77 | 20.77 | 29.62 | 23.42 | 25.15 | 28.59 | 25.95 | 24.63 | 23.17 | 25.49 | 23.19
| CTDIw | 21.73 | 36.22 | 21 | 21 | 29.95 | 23.68 | 25.42 | 28.91 | 26.24 | 24.9 | 23.42 | 25.77 | 23.44
| k | 1.0112 | 1.0112 | 1.0111 | 1.0111 | 1.0111 | 1.0111 | 1.0107 | 1.0112 | 1.0112 | 1.011 | 1.0108 | 1.011 | 1.0108
Sn 100 | Center | 0.79 | 1.31 | 0.76 | 0.76 | 1.08 | 0.86 | 0.92 | 1.05 | 0.95 | 0.9 | 0.85 | 0.93 | 0.85
| CTDIw | 0.77 | 1.29 | 0.75 | 0.75 | 1.07 | 0.84 | 0.9 | 1.03 | 0.93 | 0.89 | 0.83 | 0.92 | 0.83
| k | 0.97468 | 0.98473 | 0.98684 | 0.98684 | 0.99074 | 0.97674 | 0.97826 | 0.98095 | 0.97895 | 0.98889 | 0.97647 | 0.98925 | 0.97647
Sn 120 | Center | 1.92 | 3.21 | 1.86 | 1.86 | 2.65 | 2.1 | 2.25 | 2.56 | 2.32 | 2.2 | 2.07 | 2.28 | 2.07
| CTDIw | 1.92 | 3.2 | 1.85 | 1.85 | 2.65 | 2.09 | 2.25 | 2.55 | 2.32 | 2.2 | 2.07 | 2.28 | 2.07
| k | 1 | 0.99688 | 0.99462 | 0.99462 | 1 | 0.99524 | 1 | 0.99609 | 1 | 1 | 1 | 1 | 1
Sn 150 | Center | 4.99 | 8.31 | 4.82 | 4.82 | 6.87 | 5.43 | 5.83 | 6.63 | 6.02 | 5.71 | 5.38 | 5.91 | 5.38
| CTDIw | 5.01 | 8.35 | 4.84 | 4.84 | 6.9 | 5.46 | 5.86 | 6.66 | 6.05 | 5.74 | 5.4 | 5.94 | 5.4
| k | 1.004 | 1.0048 | 1.0041 | 1.0041 | 1.0044 | 1.0055 | 1.0051 | 1.0045 | 1.005 | 1.0053 | 1.0037 | 1.0051 | 1.0037
16 cm phantom, Narrow shaped filter (Inner Ear)
| Collimation | 2x1 | 6x0.6 | 1x5 | 1x10 | 12x1.2 | 32x0.6 | 32x0.6 | 32x0.6 | 48x0.6 | 64x0.6 | 96x0.6 | 24x1.2 | 48x1.2
| Total | 2 | 3.6 | 5 | 10 | 14.4 | 19.2 | 19.2 | 19.2 | 28.8 | 38.4 | 57.6 | 28.8 | 57.6
| Acquisition | 2x1 | 6x0.6 | 1x5 | 1x10 | 12x1.2 | 32x0.6 | 64x0.6 | 64x0.6 | 96x0.6 | 128x0.6 | 192x0.6 | 24x1.2 | 48*1.2
| Shaped filter | [5 empty] | UHR | sURH z-sharp | z-sharp | z-sharp | z-sharp | z-sharp | z-sharp
kV
70 | Center | 2.28 | 3.8 | 2.2 | 2.2 | 3.14 | 2.48 | 2.67 | 3.03 | 2.75 | 2.61 | 2.46 | 2.7 | 2.46
| CTDIw | 2.25 | 3.76 | 2.18 | 2.18 | 3.11 | 2.46 | 2.64 | 3 | 2.72 | 2.58 | 2.43 | 2.64 | 2.43
| k | 0.98684 | 0.98947 | 0.99091 | 0.99091 | 0.99045 | 0.99194 | 0.98876 | 0.9901 | 0.98909 | 0.98851 | 0.9878 | 0.97778 | 0.9878
80 | Center | 3.67 | 6.12 | 3.55 | 3.55 | 5.06 | 4 | 4.29 | 4.88 | 4.43 | 4.21 | 3.96 | 4.35 | 3.96
| CTDIw | 3.6 | 5.99 | 3.48 | 3.48 | 4.96 | 3.92 | 4.21 | 4.78 | 4.34 | 4.12 | 3.88 | 4.27 | 3.88
| k | 0.98093 | 0.97876 | 0.98028 | 0.98028 | 0.98024 | 0.98 | 0.98135 | 0.97951 | 0.97968 | 0.97862 | 0.9798 | 0.98161 | 0.9798
90 | Center | 5.43 | 9.05 | 5.25 | 52.25 | 7.49 | 5.92 | 6.35 | 7.23 | 6.56 | 6.22 | 5.86 | 6.44 | 5.86
| CTDIw | 5.29 | 8.82 | 5.11 | 5.11 | 7.29 | 5.76 | 6.19 | 7.04 | 6.39 | 6.06 | 5.7 | 6.27 | 5.71
| k | 0.97422 | 0.97459 | 0.97333 | 0.097799 | 0.9733 | 0.97297 | 0.9748 | 0.97372 | 0.97409 | 0.97428 | 0.9727 | 0.9736 | 0.9744
100 | Center | 7.45 | 12.43 | 7.2 | 7.2 | 10.26 | 8.12 | 8.71 | 9.91 | 8.99 | 8.53 | 8.03 | 8.83 | 8.03
| CTDIw | 7.23 | 12.04 | 6.98 | 6.98 | 9.96 | 7.87 | 8.45 | 9.61 | 8.72 | 8.28 | 7.79 | 8.57 | 7.79
| k | 0.97047 | 0.96862 | 0.96944 | 0.96944 | 0.97076 | 0.96921 | 0.97015 | 0.96973 | 0.96997 | 0.97069 | 0.97011 | 0.97055 | 0.97011
110 | Center | 9.66 | 16.11 | 9.34 | 9.34 | 13.32 | 10.53 | 11.31 | 12.86 | 11.67 | 11.07 | 10.42 | 11.46 | 10.43
| CTDIw | 9.37 | 15.62 | 9.06 | 9.06 | 12.92 | 10.22 | 10.97 | 12.47 | 11.32 | 10.74 | 10.1 | 11.12 | 10.11
| k | 0.96998 | 0.96958 | 0.97002 | 0.97002 | 0.96997 | 0.97056 | 0.96994 | 0.96967 | 0.97001 | 0.97019 | 0.96929 | 0.97033 | 0.96932
120 | Center | 12.1 | 20.16 | 11.69 | 11.69 | 16.67 | 13.18 | 14.15 | 16.09 | 14.6 | 13.86 | 13.04 | 14.34 | 13.05
| CTDIw | 11.74 | 19.57 | 11.34 | 11.34 | 16.18 | 12.79 | 13.73 | 15.62 | 14.17 | 13.45 | 12.65 | 13.92 | 12.66
| k | 0.97025 | 0.97073 | 0.97006 | 0.97006 | 0.97061 | 0.97041 | 0.97032 | 0.97079 | 0.97055 | 0.97042 | 0.97009 | 0.97071 | 0.97011
130 | Center | 16.63 | 24.38 | 14.13 | 14.13 | 20.16 | 15.94 | 17.12 | 19.46 | 17.66 | 16.76 | 15.77 | 17.35 | 15.78
| CTDIw | 14.19 | 23.66 | 13.71 | 13.71 | 19.56 | 15.47 | 16.6 | 18.88 | 17.13 | 16.6 | 15.3 | 16.83 | 15.31
| k | 0.85328 | 0.97047 | 0.97028 | 0.97028 | 0.97024 | 0.97051 | 0.96963 | 0.9702 | 0.96999 | 0.99045 | 0.9702 | 0.97003 | 0.97022
140 | Center | 17.3 | 28.83 | 16.71 | 16.71 | 23.84 | 18.85 | 20.24 | 23.01 | 20.88 | 19.82 | 18.64 | 20.51 | 18.66
| CTDIw | 16.79 | 27.98 | 16.22 | 16.22 | 23.14 | 18.3 | 19.64 | 22.33 | 20.27 | 19.24 | 19.1 | 19.91 | 18.11
| k | 0.97052 | 0.97052 | 0.97068 | 0.97068 | 0.97064 | 0.97082 | 0.97036 | 0.97045 | 0.97079 | 0.97074 | 1.0247 | 0.97075 | 0.97053
150 | Center | 20.1 | 33.49 | 19.42 | 19.42 | 27.7 | 21.9 | 23.51 | 26.73 | 24.26 | 23.03 | 21.66 | 23.83 | 21.68
| CTDIw | 19.51 | 32.51 | 18.85 | 18.85 | 26.88 | 21.26 | 22.82 | 2595 | 23.55 | 22.35 | 21.03 | 23.13 | 21.04
| k | 0.97065 | 0.97074 | 0.97065 | 0.97065 | 0.9704 | 0.97078 | 0.97065 | 97.0819 | 0.97073 | 0.97047 | 0.97091 | 0.97063 | 0.97048
Sn 100 | Center | 0.74 | 1.23 | 0.71 | 0.71 | 1.02 | 1 | 0.86 | 98 | 0.89 | 0.85 | 0.8 | 0.88 | 0.8
| CTDIw | 0.71 | 1.18 | 0.68 | 0.68 | 0.98 | 0.77 | 0.83 | 0.94 | 0.85 | 0.81 | 0.76 | 0.84 | 0.76
| k | 0.95946 | 0.95935 | 0.95775 | 0.95775 | 0.96078 | 0.77 | 0.96512 | 0.0095918 | 0.95506 | 0.95294 | 0.95 | 0.95455 | 0.95
Sn 120 | Center | 1.83 | 3.04 | 1.76 | 1.76 | 2.52 | 1.99 | 2.14 | 2.43 | 2.2 | 2.09 | 1.97 | 2.17 | 1.97
| CTDIw | 1.77 | 2.95 | 1.71 | 1.71 | 2.44 | 1.93 | 2.07 | 2.35 | 2.14 | 2.03 | 1.91 | 2.1 | 1.91
| k | 0.96721 | 0.97039 | 0.97159 | 0.97159 | 0.96825 | 0.96985 | 0.96729 | 0.96708 | 0.97273 | 0.97129 | 0.96954 | 0.96774 | 0.96954
Sn 150 | Center | 4.78 | 7.97 | 4.62 | 4.62 | 6.59 | 5.21 | 5.59 | 6.36 | 5.77 | 5.48 | 5.15 | 567 | 5.16
| CTDIw | 4.66 | 7.77 | 4.5 | 4.5 | 6.42 | 5.08 | 5.45 | 6.2 | 5.63 | 5.34 | 5.02 | 5.53 | 5.03
| k | 0.9749 | 0.97491 | 0.97403 | 0.97403 | 0.9742 | 0.97505 | 0.97496 | 0.97484 | 0.97574 | 0.97445 | 0.97476 | 0.0097531 | 0.97481
32 cm phantom, Standard shaped filter (Adult Body)
| Collimation | 2x1 | 6x0.6 | 1x5 | 1x10 | 12x1.2 | 32x0.6 | 32x0.6 | 32x0.6 | 48x0.6 | 64x0.6 | 96x0.6 | 24x1.2 | 48x1.2
| Total | 2 | 3.6 | 5 | 10 | 14.4 | 19.2 | 19.2 | 19.2 | 28.8 | 38.4 | 57.6 | 28.8 | 57.6
| Acquisition | 2x1 | 6x0.6 | 1x5 | 1x10 | 12x1.2 | 32x0.6 | 64x0.6 | 64x0.6 | 96x0.6 | 128x0.6 | 192x0.6 | 24x1.2 | 48*1.2
| Shaped filter | [5 empty] | UHR | sURH z-sharp | z-sharp | z-sharp | z-sharp | z-sharp | z-sharp
kV
70 | Center | 0.58 | 0.97 | 0.56 | 0.56 | 0.8 | 0.64 | 0.68 | 0.78 | 0.7 | 0.67 | 0.63 | 0.69 | 0.63
| CTDIw | 1.09 | 1.81 | 1.05 | 1.05 | 1.5 | 1.18 | 1.27 | 1.44 | 1.31 | 1.24 | 1.17 | 1.29 | 1.17
| k | 1.8793 | 1.866 | 1.875 | 1.875 | 1.875 | 1.8437 | 1.8676 | 1.8462 | 1.8714 | 1.8507 | 1.8571 | 1.8696 | 1.8571
80 | Center | 1.03 | 1.72 | 1 | 1 | 1.42 | 1.12 | 1.21 | 1.37 | 1.24 | 1.18 | 1.11 | 1.22 | 1.11
| CTDIw | 1.78 | 2.97 | 1.72 | 1.72 | 2.46 | 1.94 | 2.08 | 2.37 | 2.15 | 2.04 | 1.92 | 2.11 | 1.92
| k | 1.7282 | 1.7267 | 1.72 | 1.72 | 1.7324 | 1.7321 | 1.719 | 1.7299 | 1.7339 | 1.7288 | 1.7297 | 1.7295 | 1.7297
90 | Center | 1.62 | 2.69 | 1.56 | 1.56 | 2.23 | 1.76 | 1.89 | 2.15 | 1.95 | 1.85 | 1.7 | 1.92 | 1.74
| CTDIw | 2.68 | 4.46 | 2.59 | 2.59 | 3.69 | 2.92 | 3.13 | 3.56 | 3.23 | 3.07 | 2.89 | 3.18 | 2.89
| k | 1.6543 | 1.658 | 1.6603 | 1.6603 | 1.6547 | 1.6591 | 1.6561 | 1.6558 | 1.6564 | 1.6595 | 1.7 | 1.6563 | 1.6609
100 | Center | 2.32 | 3.86 | 2.24 | 2.24 | 3.19 | 2.52 | 2.71 | 3.08 | 2.8 | 2.65 | 2.5 | 2.75 | 2.5
| CTDIw | 3.72 | 6.2 | 3.59 | 3.59 | 5.13 | 4.05 | 4.35 | 4.95 | 4.49 | 4.26 | 4.01 | 4.41 | 4.01
| k | 1.6034 | 1.6062 | 1.6027 | 1.6027 | 1.6082 | 1.6071 | 1.6052 | 1.6071 | 1.6036 | 1.6075 | 1.604 | 1.6036 | 1.604
110 | Center | 3.11 | 5.19 | 3.01 | 3.01 | 4.29 | 3.39 | 3.64 | 4.14 | 3.76 | 3.57 | 3.36 | 3.69 | 3.36
| CTDIw | 4.84 | 8.07 | 4.68 | 4.68 | 6.67 | 5.28 | 5.66 | 6.44 | 5.85 | 5.55 | 5.22 | 5.74 | 5.22
| k | 1.5563 | 1.5549 | 1.5548 | 1.5548 | 1.5548 | 1.5575 | 1.5549 | 1.5556 | 1.5559 | 1.5546 | 1.5536 | 1.5556 | 1.5536
120 | Center | 4.02 | 6.69 | 3.88 | 3.88 | 5.53 | 4.38 | 4.7 | 5.34 | 4.85 | 4.6 | 4.33 | 4.76 | 4.33
| CTDIw | 6.2 | 10.34 | 5.99 | 5.99 | 8.55 | 6.76 | 7.26 | 8.25 | 7.49 | 7.11 | 6.69 | 7.36 | 6.69
| k | 1.5423 | 1.5456 | 1.5438 | 1.5438 | 1.5461 | 1.5434 | 1.5447 | 1.5449 | 1.5443 | 1.5457 | 1.545 | 1.5462 | 1.545
130 | Center | 4.97 | 8.28 | 4.8 | 4.8 | 6.84 | 5.41 | 5.81 | 6.6 | 5.99 | 5.69 | 5.35 | 5.89 | 5.36
| CTDIw | 7.58 | 12.63 | 7.32 | 7.32 | 10.44 | 8.26 | 8.86 | 10.08 | 9.15 | 8.67 | 8.17 | 8.98 | 8.17
| k | 1.5252 | 1.5254 | 1.525 | 1.525 | 1.5263 | 1.5268 | 1.525 | 1.5273 | 1.5275 | 1.5237 | 1.5271 | 1.5246 | 1.5243
140 | Center | 5.98 | 9.97 | 5.78 | 5.78 | 8.25 | 6.52 | 7 | 7.96 | 7.23 | 6.86 | 6.45 | 7.1 | 6.46
| CTDIw | 9.05 | 15.08 | 8.74 | 8.74 | 12.47 | 9.86 | 10.58 | 12.03 | 10.92 | 10.36 | 9.75 | 10.73 | 9.76
| k | 1.5134 | 1.5125 | 1.5121 | 1.5121 | 1.5115 | 1.5123 | 1.5114 | 1.5113 | 1.5104 | 1.5102 | 1.5116 | 1.5113 | 1.5108
150 | Center | 7.07 | 11.79 | 6.83 | 6.83 | 9.75 | 7.71 | 8.25 | 9.41 | 8.54 | 8.11 | 8.62 | 8.39 | 7.63
| CTDIw | 10.61 | 17.68 | 10.25 | 10.25 | 14.62 | 11.56 | 12.41 | 14.11 | 12.81 | 12.15 | 11.43 | 12.58 | 11.44
| k | 1.5007 | 1.4996 | 1.5007 | 1.5007 | 1.4995 | 1.4994 | 1.5042 | 1.4995 | 1.5 | 1.4982 | 1.326 | 1.4994 | 1.4993
Sn 100 | Center | 0.28 | 0.47 | 0.27 | 0.27 | 0.38 | 0.3 | 0.33 | 0.37 | 0.34 | 0.32 | 0.3 | 0.33 | 0.3
| CTDIw | 0.39 | 0.66 | 0.38 | 0.38 | 0.54 | 0.43 | 0.46 | 0.52 | 0.47 | 0.45 | 0.42 | 0.47 | 0.42
| k | 1.3929 | 1.4043 | 1.4074 | 1.4074 | 1.4211 | 1.4333 | 1.3939 | 1.4054 | 1.3824 | 1.4062 | 1.4 | 1.4242 | 1.4
Sn 120 | Center | 0.71 | 1.19 | 0.69 | 0.69 | 0.99 | 0.78 | 0.84 | 0.95 | 0.86 | 0.82 | 0.77 | 0.85 | 0.77
| CTDIw | 1 | 1.67 | 0.97 | 0.97 | 1.38 | 1.09 | 1.17 | 1.33 | 1.21 | 1.15 | 1.08 | 1.19 | 1.08
| k | 1.4085 | 1.4034 | 1.4058 | 1.4058 | 1.3939 | 1.3974 | 1.3929 | 1.4 | 1.407 | 1.4024 | 1.4026 | 1.4 | 1.4026
Sn 150 | Center | 1.96 | 3.27 | 1.89 | 1.89 | 2.7 | 2.14 | 2.29 | 2.61 | 2.37 | 2.25 | 2.11 | 2.32 | 2.11
| CTDIw | 2.7 | 4.51 | 2.61 | 2.61 | 3.73 | 2.95 | 3.16 | 3.6 | 3.26 | 3.1 | 2.92 | 3.21 | 2.92
| k | 1.3776 | 1.3792 | 1.381 | 1.381 | 1.3815 | 1.3785 | 1.3799 | 1.3793 | 1.3755 | 1.3778 | 1.3839 | 1.3836 | 1.3839
32 cm phantom, Narrow shaped filter (Cardiac, Ped Body, Flash)
| Collimation | 2x1 | 6x0.6 | 1x5 | 1x10 | 12x1.2 | 32x0.6 | 32x0.6 | 32x0.6 | 48x0.6 | 64x0.6 | 96x0.6 | 24x1.2 | 48x1.2
| Total | 2 | 3.6 | 5 | 10 | 14.4 | 19.2 | 19.2 | 19.2 | 28.8 | 38.4 | 57.6 | 28.8 | 57.6
| Acquisition | 2x1 | 6x0.6 | 1x5 | 1x10 | 12x1.2 | 32x0.6 | 64x0.6 | 64x0.6 | 96x0.6 | 128x0.6 | 192x0.6 | 24x1.2 | 48*1.2
| Shaped filter | [5 empty] | UHR | sURH z-sharp | z-sharp | z-sharp | z-sharp | z-sharp | z-sharp
kV
70 | Center | 0.51 | 0.84 | 0.49 | 0.49 | 0.7 | 0.55 | 0.59 | 0.67 | 0.61 | 0.58 | 0.54 | 0.6 | 0.55
| CTDIw | 0.83 | 1.39 | 0.81 | 0.81 | 1.15 | 0.91 | 0.98 | 1.11 | 1.01 | 0.96 | 0.9 | 0.99 | 0.9
| k | 1.6275 | 1.6548 | 1.6531 | 1.6531 | 1.6429 | 1.6545 | 1.661 | 1.6567 | 1.6557 | 1.6552 | 1.6667 | 1.65 | 1.6364
80 | Center | 0.9 | 1.49 | 0.87 | 0.87 | 1.23 | 0.98 | 1.05 | 1.19 | 1.08 | 1.03 | 0.97 | 1.06 | 0.97
| CTDIw | 1.36 | 2.27 | 1.32 | 1.32 | 1.88 | 1.48 | 1.59 | 1.81 | 1.64 | 1.56 | 1.47 | 1.61 | 1.47
| k | 1.5111 | 1.5235 | 1.5172 | 1.5172 | 1.5285 | 1.5102 | 1.5143 | 1.521 | 1.5185 | 1.5146 | 1.5155 | 1.5189 | 1.5155
90 | Center | 1.4 | 2.34 | 1.35 | 1.35 | 1.93 | 1.53 | 1.64 | 1.87 | 1.69 | 1.61 | 1.51 | 1.66 | 1.51
| CTDIw | 2.1 | 3.49 | 2.02 | 2.02 | 2.89 | 2.28 | 2.45 | 2.79 | 2.53 | 2.4 | 2.26 | 2.49 | 2.26
| k | 1.5 | 1.4915 | 1.4963 | 1.4963 | 1.4974 | 1.4902 | 1.4939 | 1.492 | 1.497 | 1.4907 | 1.4967 | 1.5 | 1.4967
100 | Center | 2.04 | 3.41 | 1.98 | 1.98 | 2.82 | 2.23 | 2.39 | 2.72 | 2.47 | 2.34 | 2.2 | 2.42 | 2.21
| CTDIw | 2.94 | 4.9 | 2.84 | 2.84 | 4.05 | 3.2 | 3.44 | 3.91 | 3.55 | 3.37 | 3.17 | 3.48 | 3.17
| k | 1.4412 | 1.437 | 1.4343 | 1.4343 | 1.4362 | 1.435 | 1.4393 | 1.4375 | 1.4372 | 1.4402 | 1.4409 | 1.438 | 1.4344
110 | Center | 2.75 | 4.59 | 2.66 | 2.66 | 3.79 | 3 | 3.22 | 3.66 | 3.32 | 3.5 | 2.97 | 3.26 | 2.97
| CTDIw | 3.89 | 6.48 | 3.76 | 3.76 | 5.36 | 4.24 | 4.55 | 5.17 | 4.69 | 4.45 | 4.19 | 4.61 | 4.19
| k | 1.4145 | 1.4118 | 1.4135 | 1.4135 | 1.4142 | 1.4133 | 1.413 | 1.4126 | 1.4127 | 1.2714 | 1.4108 | 1.4141 | 1.4108
120 | Center | 3.56 | 5.94 | 3.44 | 3.44 | 4.91 | 3.88 | 4.17 | 4.74 | 4.3 | 4.08 | 3.84 | 4.23 | 3.85
| CTDIw | 4.95 | 8.26 | 4.79 | 4.79 | 6.83 | 5.4 | 5.8 | 6.59 | 5.98 | 5.68 | 5.34 | 5.87 | 5.34
| k | 1.3904 | 1.3906 | 1.3924 | 1.3924 | 1.391 | 1.3918 | 1.3909 | 1.3903 | 1.3907 | 1.3922 | 1.3906 | 1.3877 | 1.387
130 | Center | 4.42 | 7.37 | 4.27 | 4.27 | 609 | 4.82 | 5.17 | 5.88 | 5.34 | 5.07 | 4.77 | 5.24 | 4.77
| CTDIw | 6.08 | 10.13 | 5.87 | 5.87 | 8.38 | 6.63 | 7.11 | 8.09 | 7.34 | 6.97 | 6.55 | 7.21 | 6.56
| k | 1.3756 | 1.3745 | 1.3747 | 1.3747 | 0.01376 | 1.3755 | 1.3752 | 1.3759 | 1.3745 | 1.3748 | 1.3732 | 1.376 | 1.3753
140 | Center | 5.34 | 8.9 | 5.16 | 5.16 | 7.36 | 5.82 | 6.25 | 7.1 | 6.45 | 6.12 | 5.76 | 6.33 | 5.76
| CTDIw | 7.28 | 12.14 | 7.04 | 7.04 | 10.04 | 7.94 | 8.52 | 9.69 | 8.79 | 8.35 | 7.85 | 8.64 | 7.86
| k | 1.3633 | 1.364 | 1.3643 | 1.3643 | 1.3641 | 1.3643 | 1.3632 | 1.3648 | 1.3628 | 1.3644 | 1.3628 | 1.3649 | 1.3646
150 | Center | 6.33 | 10.56 | 6.12 | 6.12 | 8.73 | 6.9 | 7.41 | 8.42 | 7.65 | 7.26 | 6.83 | 7.51 | 6.83
| CTDIw | 8.56 | 14.27 | 8.27 | 8.27 | 11.8 | 9.33 | 10.02 | 11.39 | 10.34 | 9.81 | 9.23 | 10.16 | 9.24
| k | 1.3523 | 1.3513 | 1.3513 | 1.3513 | 1.3517 | 1.3522 | 1.3522 | 1.3527 | 1.3516 | 1.3512 | 1.3514 | 1.3529 | 1.3529
Sn 100 | Center | 0.25 | 0.42 | 0.24 | 0.24 | 0.35 | 0.27 | 0.29 | 0.33 | 0.3 | 0.29 | 0.27 | 0.3 | 0.27
| CTDIw | 0.32 | 0.53 | 0.31 | 0.31 | 0.44 | 0.35 | 0.38 | 0.43 | 0.39 | 0.37 | 0.35 | 0.38 | 0.35
| k | 1.28 | 1.2619 | 1.2917 | 1.2917 | 1.2571 | 1.2963 | 1.3103 | 1.303 | 1.3 | 1.2759 | 1.2963 | 1.2667 | 1.2963
Sn 120 | Center | 0.65 | 1.08 | 0.62 | 0.62 | 0.89 | 0.7 | 0.76 | 0.86 | 0.78 | 0.74 | 0.7 | 0.77 | 0.7
| CTDIw | 0.83 | 1.38 | 0.8 | 0.8 | 1.14 | 0.9 | 0.97 | 1.1 | 1 | 0.95 | 0.89 | 0.98 | 0.89
| k | 1.2769 | 1.2778 | 1.2903 | 1.2903 | 1.2809 | 1.2857 | 1.2763 | 1.2791 | 1.2821 | 1.2838 | 1.2714 | 1.2727 | 1.2714
Sn 150 | Center | 1.78 | 2.97 | 1.72 | 1.72 | 2.46 | 1.94 | 2.09 | 2.37 | 2.15 | 2.04 | 1.92 | 2.12 | 1.92
| CTDIw | 2.26 | 3.77 | 2.18 | 2.18 | 3.12 | 2.46 | 2.65 | 3.01 | 2.73 | 2.59 | 2.44 | 2.68 | 2.44
| k | 1.2697 | 1.2694 | 1.2674 | 1.2674 | 1.2683 | 1.268 | 1.2679 | 1.27 | 1.2698 | 1.2696 | 1.2708 | 1.2642 | 1.2708

## Somatom Flash
16 cm phantom, Wide shaped filter (Routine head)
| Collimation | 2x1 | 6x0.6 | 8x0.6 | 1x5 | 10x0.6 | 1x10 | 12x1.2 | 20x0.6 | 32x0.6 | 32x1.2 | 64x0.6
| Total | 2 | 3.6 | 2.4 | 5 | 6 | 10 | 14.4 | 12 | 19.2 | 38.4 | 38.4
| Acquisition | 2x1 | 6x0.6 | 16x0.6 | 1x5 | 20x0.6 | 1x10 | 12x1.2 | 40x0.6 | 64x0.6 | 32x1.2 | 128x0.6
| Shaped filter |  |  | z-sharp |  | z-sharp |  |  | z-sharp | z-sharp |  | z-sharp
kV |  |  | 4.7
70 | Center | 2.9 | 4.7 | 5 | 2.6 | 3.4 | 2.6 | 3.3 | 3.5 | 3.3 | 2.9 | 2.8
| CTDIw | 3.3 | 5.3 | 5.6 | 2.9 | 3.9 | 2.9 | 3.8 | 4 | 3.7 | 3.3 | 3.2
| k | 1.1379 | 1.1277 | 1.12 | 1.1154 | 1.1471 | 1.1154 | 1.1515 | 1.1429 | 1.1212 | 1.1379 | 1.1429
80 | Center | 4.6 | 7.4 | 7.9 | 4.1 | 5.4 | 4.1 | 5.3 | 5.6 | 5.2 | 4.6 | 4.5
| CTDIw | 5 | 8.2 | 8.6 | 4.5 | 6 | 4.5 | 5.8 | 6.1 | 5.7 | 5.1 | 5
| k | 1.087 | 1.1081 | 1.0886 | 1.0976 | 1.1111 | 1.0976 | 1.0943 | 1.0893 | 1.0962 | 1.1087 | 1.1111
100 | Center | 9.2 | 14.9 | 15.8 | 8.2 | 10.9 | 8.2 | 10.6 | 11.2 | 10.4 | 9.2 | 9.1
| CTDIw | 9.8 | 15.9 | 16.8 | 8.8 | 11.6 | 8.8 | 11.3 | 11.9 | 11 | 9.8 | 9.6
| k | 1.0652 | 1.0671 | 1.0633 | 1.0732 | 1.0642 | 1.0732 | 1.066 | 1.0625 | 1.0577 | 1.0652 | 1.0549
120 | Center | 14.8 | 23.9 | 25.4 | 13.3 | 17.6 | 13.3 | 17 | 17.9 | 16.7 | 14.8 | 14.6
| CTDIw | 15.5 | 25.1 | 26.6 | 13.9 | 18.4 | 13.9 | 17.8 | 18.8 | 17.5 | 15.6 | 15.3
| k | 1.0473 | 1.0502 | 1.0472 | 1.0451 | 1.0455 | 1.0451 | 1.0471 | 1.0503 | 1.0479 | 1.0541 | 1.0479
140 | Center | 22 | 35.6 | 37.8 | 19.7 | 26.2 | 19.7 | 25.3 | 26.7 | 24.9 | 22.1 | 21.8
| CTDIw | 22.9 | 37.1 | 39.3 | 20.5 | 27.2 | 20.5 | 26.3 | 27.8 | 25.9 | 23 | 22.6
| k | 1.0409 | 1.0421 | 1.0397 | 1.0406 | 1.0382 | 1.0406 | 1.0395 | 1.0412 | 1.0402 | 1.0407 | 1.0367
Sn 140 (Tube B only) | Center | 6.6 | 10.6 | 11.3 | 5.9 | 7.8 | 5.9 | 7.5 | 8 | 7.4 | 6.6 | 6.5
| CTDIw | 6.6 | 10.7 | 11.4 | 5.9 | 7.9 | 5.9 | 7.6 | 8 | 7.5 | 6.6 | 6.5
| k | 1 | 1.0094 | 1.0088 | 1 | 1.0128 | 1 | 1.0133 | 1 | 1.0135 | 1 | 1
16 cm phantom, Narrow shaped filter (Inner Ear)
| Collimation | 2x1 | 6x0.6 | 8x0.6 | 1x5 | 10x0.6 | 1x10 | 12x1.2 | 20x0.6 | 32x0.6 | 32x1.2 | 64x0.6
| Total | 2 | 3.6 | 2.4 | 5 | 6 | 10 | 14.4 | 12 | 19.2 | 38.4 | 38.4
| Acquisition | 2x1 | 6x0.6 | 16x0.6 | 1x5 | 20x0.6 | 1x10 | 12x1.2 | 40x0.6 | 64x0.6 | 32x1.2 | 128x0.6
| Shaped filter |  |  | z-sharp |  | z-sharp |  |  | z-sharp | z-sharp |  | z-sharp
kV
70 | Center | 2.28 | 3.8 | 2.2 | 2.2 | 3.14 | 2.48 | 2.67 | 3.03 | 2.75 | 2.61 | 2.46
| CTDIw | 2.25 | 3.76 | 2.18 | 2.18 | 3.11 | 2.46 | 2.64 | 3 | 2.72 | 2.58 | 2.43
| k | 0.98684 | 0.98947 | 0.99091 | 0.99091 | 0.99045 | 0.99194 | 0.98876 | 0.9901 | 0.98909 | 0.98851 | 0.9878
80 | Center | 3.67 | 6.12 | 3.55 | 3.55 | 5.06 | 4 | 4.29 | 4.88 | 4.43 | 4.21 | 3.96
| CTDIw | 3.6 | 5.99 | 3.48 | 3.48 | 4.96 | 3.92 | 4.21 | 4.78 | 4.34 | 4.12 | 3.88
| k | 0.98093 | 0.97876 | 0.98028 | 0.98028 | 0.98024 | 0.98 | 0.98135 | 0.97951 | 0.97968 | 0.97862 | 0.9798
100 | Center | 5.43 | 9.05 | 5.25 | 52.25 | 7.49 | 5.92 | 6.35 | 7.23 | 6.56 | 6.22 | 5.86
| CTDIw | 5.29 | 8.82 | 5.11 | 5.11 | 7.29 | 5.76 | 6.19 | 7.04 | 6.39 | 6.06 | 5.7
| k | 0.97422 | 0.97459 | 0.97333 | 0.097799 | 0.9733 | 0.97297 | 0.9748 | 0.97372 | 0.97409 | 0.97428 | 0.9727
120 | Center | 7.45 | 12.43 | 7.2 | 7.2 | 10.26 | 8.12 | 8.71 | 9.91 | 8.99 | 8.53 | 8.03
| CTDIw | 7.23 | 12.04 | 6.98 | 6.98 | 9.96 | 7.87 | 8.45 | 9.61 | 8.72 | 8.28 | 7.79
| k | 0.97047 | 0.96862 | 0.96944 | 0.96944 | 0.97076 | 0.96921 | 0.97015 | 0.96973 | 0.96997 | 0.97069 | 0.97011
140 | Center | 9.66 | 16.11 | 9.34 | 9.34 | 13.32 | 10.53 | 11.31 | 12.86 | 11.67 | 11.07 | 10.42
| CTDIw | 9.37 | 15.62 | 9.06 | 9.06 | 12.92 | 10.22 | 10.97 | 12.47 | 11.32 | 10.74 | 10.1
| k | 0.96998 | 0.96958 | 0.97002 | 0.97002 | 0.96997 | 0.97056 | 0.96994 | 0.96967 | 0.97001 | 0.97019 | 0.96929
Sn 140 (Tube B only) | Center | 12.1 | 20.16 | 11.69 | 11.69 | 16.67 | 13.18 | 14.15 | 16.09 | 14.6 | 13.86 | 13.04
| CTDIw | 11.74 | 19.57 | 11.34 | 11.34 | 16.18 | 12.79 | 13.73 | 15.62 | 14.17 | 13.45 | 12.65
| k | 0.97025 | 0.97073 | 0.97006 | 0.97006 | 0.97061 | 0.97041 | 0.97032 | 0.97079 | 0.97055 | 0.97042 | 0.97009
32 cm phantom, Standard shaped filter (Adult Body) Tube A only
| Collimation | 2x1 | 6x0.6 | 8x0.6 | 1x5 | 10x0.6 | 1x10 | 12x1.2 | 20x0.6 | 32x0.6 | 32x1.2 | 64x0.6
| Total | 2 | 3.6 | 2.4 | 5 | 6 | 10 | 14.4 | 12 | 19.2 | 38.4 | 38.4
| Acquisition | 2x1 | 6x0.6 | 16x0.6 | 1x5 | 20x0.6 | 1x10 | 12x1.2 | 40x0.6 | 64x0.6 | 32x1.2 | 128x0.6
| Shaped filter |  |  | z-sharp |  | z-sharp |  |  | z-sharp | z-sharp |  | z-sharp
kV
70 | Center | 0.6 | 1 | 1.1 | 0.6 | 0.8 | 0.6 | 0.7 | 0.8 | 0.7 | 0.6 | 0.6
| CTDIw | 1.3 | 2.1 | 2.2 | 1.1 | 1.5 | 1.1 | 1.5 | 1.5 | 1.4 | 1.3 | 1.2
| k | 2.1667 | 2.1 | 2 | 1.8333 | 1.875 | 1.8333 | 2.1429 | 1.875 | 2 | 2.1667 | 2
80 | Center | 1.1 | 1.8 | 1.9 | 1 | 1.3 | 1 | 1.3 | 1.3 | 1.2 | 1.1 | 1.1
| CTDIw | 2 | 3.2 | 3.4 | 1.8 | 2.4 | 1.8 | 2.3 | 2.4 | 2.2 | 2 | 2
| k | 1.8182 | 1.7778 | 1.7895 | 1.8 | 1.8462 | 1.8 | 1.7692 | 1.8462 | 1.8333 | 1.8182 | 1.8182
100 | Center | 2.5 | 4.1 | 4.3 | 2.3 | 3 | 2.3 | 2.9 | 3.1 | 2.9 | 2.5 | 2.5
| CTDIw | 4.2 | 6.8 | 7.2 | 3.8 | 5 | 3.8 | 4.8 | 5.1 | 4.7 | 4.2 | 4.1
| k | 1.68 | 1.6585 | 1.6744 | 1.6522 | 1.6667 | 1.6522 | 1.6552 | 1.6452 | 1.6207 | 1.68 | 1.64
120 | Center | 4.4 | 7.1 | 7.5 | 3.9 | 5.2 | 3.9 | 5 | 5.3 | 4.9 | 4.4 | 4.3
| CTDIw | 6.9 | 11.1 | 11.7 | 6.2 | 8.1 | 6.2 | 7.9 | 8.3 | 7.7 | 6.9 | 6.7
| k | 1.5682 | 1.5634 | 1.56 | 1.5897 | 1.5577 | 1.5897 | 1.58 | 1.566 | 1.5714 | 1.5682 | 1.5581
140 | Center | 6.5 | 10.6 | 11.2 | 5.9 | 7.8 | 5.9 | 7.5 | 7.9 | 7.4 | 6.6 | 6.5
| CTDIw | 10 | 16.2 | 17.1 | 9 | 11.8 | 9 | 11.5 | 12.1 | 11.2 | 10 | 9.8
| k | 1.5385 | 1.5283 | 1.5268 | 1.5254 | 1.5128 | 1.5254 | 1.5333 | 1.5316 | 1.5135 | 1.5152 | 1.5077
32 cm phantom, Wide shaped filter (Adult Body) Tube B only
| Collimation | 2x1 | 6x0.6 | 8x0.6 | 1x5 | 10x0.6 | 1x10 | 12x1.2 | 20x0.6 | 32x0.6 | 32x1.2 | 64x0.6
| Total | 2 | 3.6 | 2.4 | 5 | 6 | 10 | 14.4 | 12 | 19.2 | 38.4 | 38.4
| Acquisition | 2x1 | 6x0.6 | 16x0.6 | 1x5 | 20x0.6 | 1x10 | 12x1.2 | 40x0.6 | 64x0.6 | 32x1.2 | 128x0.6
| Shaped filter |  |  | z-sharp |  | z-sharp |  |  | z-sharp | z-sharp |  | z-sharp
kV
70 | Center | 0.7 | 1.2 | 1.2 | 0.6 | 0.9 | 0.6 | 0.8 | 0.9 | 0.8 | 0.7 | 0.7
| CTDIw | 1.7 | 2.7 | 2.9 | 1.5 | 2 | 1.5 | 2 | 2.1 | 1.9 | 1.7 | 1.7
| k | 2.4286 | 2.25 | 2.4167 | 2.5 | 2.2222 | 2.5 | 2.5 | 2.3333 | 2.375 | 2.4286 | 2.4286
80 | Center | 1.2 | 2 | 2.1 | 1.1 | 1.5 | 1.1 | 1.4 | 1.5 | 1.4 | 1.2 | 1.2
| CTDIw | 2.6 | 4.3 | 4.5 | 2.4 | 3.1 | 2.4 | 3 | 3.2 | 2.9 | 2.6 | 2.6
| k | 2.1667 | 2.15 | 2.1429 | 2.1818 | 2.0667 | 2.1818 | 2.1429 | 2.1333 | 2.0714 | 2.1667 | 2.1667
100 | Center | 2.7 | 4.4 | 4.7 | 2.5 | 3.3 | 2.5 | 3.2 | 3.3 | 3.1 | 2.8 | 2.7
| CTDIw | 5.2 | 8.4 | 8.9 | 4.6 | 6.1 | 4.6 | 6 | 6.3 | 5.8 | 5.2 | 5.1
| k | 1.9259 | 1.9091 | 1.8936 | 1.84 | 1.8485 | 1.84 | 1.875 | 1.9091 | 1.871 | 1.8571 | 1.8889
120 | Center | 4.7 | 7.6 | 8 | 4.2 | 5.6 | 4.2 | 5.4 | 5.7 | 5.3 | 4.7 | 4.6
| CTDIw | 8.3 | 13.4 | 14.2 | 7.4 | 9.9 | 7.4 | 9.6 | 10 | 9.4 | 8.3 | 8.2
| k | 1.766 | 1.7632 | 1.775 | 1.7619 | 1.7679 | 1.7619 | 1.7778 | 1.7544 | 1.7736 | 1.766 | 1.7826
140 | Center | 7.3 | 11.8 | 12.6 | 6.6 | 8.7 | 6.6 | 8.4 | 8.9 | 8.3 | 7.3 | 7.2
| CTDIw | 12.4 | 20.1 | 21.2 | 11.1 | 14.7 | 11.1 | 14.3 | 15 | 14 | 12.5 | 12.2
| k | 1.6986 | 1.7034 | 1.6825 | 1.6818 | 1.6897 | 1.6818 | 1.7024 | 1.6854 | 1.6867 | 1.7123 | 1.6944
Sn 140 (Tube B only) | Center | 2.5 | 4.1 | 4.3 | 2.3 | 3 | 2.3 | 2.9 | 3.1 | 2.8 | 2.5 | 2.5
| CTDIw | 3.7 | 6 | 6.3 | 3.3 | 4.4 | 3.3 | 4.2 | 4.5 | 4.1 | 3.7 | 3.6
| k | 1.48 | 1.4634 | 1.4651 | 1.4348 | 1.4667 | 1.4348 | 1.4483 | 1.4516 | 1.4643 | 1.48 | 1.44
32 cm phantom, Narrow shaped filter (Cardiac, Ped Body, Flash)
| Collimation | 2x1 | 6x0.6 | 8x0.6 | 1x5 | 10x0.6 | 1x10 | 12x1.2 | 20x0.6 | 32x0.6 | 32x1.2 | 64x0.6
| Total | 2 | 3.6 | 2.4 | 5 | 6 | 10 | 14.4 | 12 | 19.2 | 38.4 | 38.4
| Acquisition | 2x1 | 6x0.6 | 16x0.6 | 1x5 | 20x0.6 | 1x10 | 12x1.2 | 40x0.6 | 64x0.6 | 32x1.2 | 128x0.6
| Shaped filter |  |  | z-sharp |  | z-sharp |  |  | z-sharp | z-sharp |  | z-sharp
kV
70 | Center | 0.6 | 0.9 | 1 | 0.5 | 0.7 | 0.5 | 0.6 | 0.7 | 0.6 | 0.6 | 0.6
| CTDIw | 1 | 1.6 | 1.7 | 0.9 | 1.2 | 0.9 | 1.1 | 1.2 | 1.1 | 1 | 1
| k | 1.6667 | 1.7778 | 1.7 | 1.8 | 1.7143 | 1.8 | 1.8333 | 1.7143 | 1.8333 | 1.6667 | 1.6667
80 | Center | 1 | 1.6 | 1.7 | 0.9 | 1.2 | 0.9 | 1.1 | 1.2 | 1.1 | 1 | 1
| CTDIw | 1.6 | 2.6 | 2.7 | 1.4 | 1.9 | 1.4 | 1.9 | 1.9 | 1.8 | 1.6 | 1.6
| k | 1.6 | 1.625 | 1.5882 | 1.5556 | 1.5833 | 1.5556 | 1.7273 | 1.5833 | 1.6364 | 1.6 | 1.6
100 | Center | 2.2 | 3.6 | 3.8 | 2 | 2.6 | 2 | 2.6 | 2.7 | 2.5 | 2.2 | 2.2
| CTDIw | 3.4 | 5.5 | 5.8 | 3 | 4 | 3 | 3.9 | 4.1 | 3.8 | 3.4 | 3.3
| k | 1.5455 | 1.5278 | 1.5263 | 1.5 | 1.5385 | 1.5 | 1.5 | 1.5185 | 1.52 | 1.5455 | 1.5
120 | Center | 3.9 | 6.3 | 6.7 | 3.5 | 4.6 | 3.5 | 4.5 | 4.7 | 4.4 | 3.9 | 3.8
| CTDIw | 5.6 | 9.1 | 9.6 | 5.1 | 6.7 | 5.1 | 6.5 | 6.8 | 6.3 | 5.7 | 5.5
| k | 1.4359 | 1.4444 | 1.4328 | 1.4571 | 1.4565 | 1.4571 | 1.4444 | 1.4468 | 1.4318 | 1.4615 | 1.4474
140 | Center | 6.1 | 9.9 | 10.5 | 5.5 | 7.3 | 5.5 | 7.1 | 7.4 | 6.9 | 6.2 | 6
| CTDIw | 8.7 | 14.1 | 14.8 | 7.8 | 10.3 | 7.8 | 10 | 10.5 | 9.7 | 8.7 | 8.5
| k | 1.4262 | 1.4242 | 1.4095 | 1.4182 | 1.411 | 1.4182 | 1.4085 | 1.4189 | 1.4058 | 1.4032 | 1.4167
Sn 140 (Tube B only) | Center | 2.2 | 3.6 | 3.8 | 2 | 2.6 | 2 | 2.5 | 2.7 | 2.5 | 2.2 | 2.2
| CTDIw | 2.9 | 4.6 | 4.9 | 2.6 | 3.4 | 2.6 | 3.3 | 3.5 | 3.2 | 2.9 | 2.8
| k | 1.3182 | 1.2778 | 1.2895 | 1.3 | 1.3077 | 1.3 | 1.32 | 1.2963 | 1.28 | 1.3182 | 1.2727

## Definition AS 128
16 cm phantom, Standard shaped filter (Routine head)
| Collimation | 2x1 | 6x0.6 | 8x0.6 | 8x0.6 | 1x5 | 10x0.6 | 6x1.2 | 1x10 | 20x0.6 | 12x1.2 | 32x0.6 | 64x0.6
| Total | 2 | 3.6 | 4.8 | 4.8 | 5 | 6 | 7.2 | 10 | 12 | 14.4 | 38.4 | 38.4
| Acquisition | 2x1 | 6x0.6 | 8x0.6 | 16x0.6 | 1x5 | 20x0.6 | 6x1.2 | 1x10 | 40x0.6 | 12x1.2 | 32x1.2 | 128x0.6
|  |  | Topogram | UHR | UHR, z-sharp |  | z-sharp |  |  | z-sharp |  |  | z-sharp
kV
70 | Center | 2.6 | 4.22 | 3.88 | 4.58 | 2.34 | 3.17 | 3.3 | 2.34 | 3.23 | 3 | 3.01 | 2.87
| CTDIw | 2.74 | 4.43 | 4.07 | 4.78 | 2.46 | 3.31 | 3.35 | 2.46 | 3.38 | 3.15 | 3.15 | 3.01
| k | 1.0538 | 1.0498 | 1.049 | 1.0437 | 1.0513 | 1.0442 | 1.0152 | 1.0513 | 1.0464 | 1.05 | 1.0465 | 1.0488
80 | Center | 4.18 | 6.76 | 6.22 | 7.24 | 3.75 | 5.02 | 5.39 | 3.75 | 5.12 | 4.81 | 4.77 | 4.6
| CTDIw | 4.36 | 7.06 | 6.49 | 7.46 | 3.91 | 5.17 | 5.62 | 3.91 | 5.27 | 5.02 | 4.91 | 4.8
| k | 1.0431 | 1.0444 | 1.0434 | 1.0304 | 1.0427 | 1.0299 | 1.0427 | 1.0427 | 1.0293 | 1.0437 | 1.0294 | 1.0435
100 | Center | 8.44 | 13.67 | 12.56 | 14.69 | 7.57 | 10.19 | 10.89 | 7.57 | 10.38 | 9.71 | 8.48 | 8.57
| CTDIw | 8.65 | 14.01 | 12.88 | 14.93 | 7.76 | 10.35 | 11.16 | 7.76 | 10.55 | 9.96 | 8.69 | 8.71
| k | 1.0249 | 1.0249 | 1.0255 | 1.0163 | 1.0251 | 1.0157 | 1.0248 | 1.0251 | 1.0164 | 1.0257 | 1.0248 | 1.0163
120 | Center | 13.85 | 22.42 | 20.6 | 24.33 | 12.42 | 16.87 | 17.86 | 12.42 | 17.19 | 15.93 | 13.91 | 14.19
| CTDIw | 14.11 | 22.85 | 21 | 24.63 | 12.66 | 17.08 | 18.2 | 12.66 | 17.41 | 16.24 | 14.17 | 14.37
| k | 1.0188 | 1.0192 | 1.0194 | 1.0123 | 1.0193 | 1.0124 | 1.019 | 1.0193 | 1.0128 | 1.0195 | 1.0187 | 1.0127
140 | Center | 20.24 | 32.77 | 30.12 | 36.16 | 18.16 | 25.07 | 26.11 | 18.16 | 25.55 | 23.3 | 20.33 | 21.09
| CTDIw | 20.6 | 33.35 | 30.65 | 36.6 | 18.48 | 25.4 | 26.56 | 18.48 | 25.89 | 23.7 | 20.69 | 21.37
| k | 1.0178 | 1.0177 | 1.0176 | 1.0122 | 1.0176 | 1.0132 | 1.0172 | 1.0176 | 1.0133 | 1.0172 | 1.0177 | 1.0133
100Sn | Center
CTDIw
k
120Sn | Center
CTDIw
k
140Sn | Center
CTDIw
k
16 cm phantom, Narrow shaped filter (Pediatric body)
| Collimation | 2x1 | 6x0.6 | 8x0.6 | 8x0.6 | 1x5 | 10x0.6 | 6x1.2 | 1x10 | 20x0.6 | 12x1.2 | 32x0.6 | 64x0.6
| Total | 2 | 3.6 | 4.8 | 4.8 | 5 | 6 | 7.2 | 10 | 12 | 14.4 | 38.4 | 38.4
| Acquisition | 2x1 | 6x0.6 | 8x0.6 | 16x0.6 | 1x5 | 20x0.6 | 6x1.2 | 1x10 | 40x0.6 | 12x1.2 | 32x1.2 | 28.x.6
| Shaped filter |  |  |  | z-sharp |  | z-sharp |  |  | z-sharp |  |  | z-sharp
kV
70 | Center | 2.36 | 3.83 | 3.52 | 4.12 | 2.12 | 2.86 | 3.05 | 2.12 | 2.91 | 2.72 | 2.38 | 2.4
| CTDIw | 2.36 | 3.81 | 3.51 | 4.07 | 2.11 | 2.82 | 3.04 | 2.11 | 2.88 | 2.71 | 2.37 | 2.37
| k | 1 | 0.99478 | 0.99716 | 0.98786 | 0.99528 | 0.98601 | 0.99672 | 0.99528 | 0.98969 | 0.99632 | 0.9958 | 0.9875
80 | Center | 3.81 | 6.17 | 5.67 | 6.58 | 3.42 | 4.56 | 4.92 | 3.42 | 4.65 | 4.39 | 3.83 | 3.84
| CTDIw | 3.78 | 6.12 | 5.63 | 6.44 | 3.39 | 4.47 | 4.88 | 3.39 | 4.55 | 4.35 | 3.8 | 3.76
| k | 0.99213 | 0.9919 | 0.99295 | 0.97872 | 0.99123 | 0.98026 | 0.99187 | 0.99123 | 0.97849 | 0.99089 | 0.99217 | 0.97917
100 | Center | 7.81 | 12.65 | 11.63 | 13.53 | 7.01 | 9.38 | 10.08 | 7.01 | 9.56 | 8.99 | 7.85 | 7.89
| CTDIw | 7.69 | 12.45 | 11.44 | 13.14 | 6.9 | 9.11 | 9.92 | 6.9 | 9.29 | 8.85 | 7.72 | 7.67
| k | 0.98464 | 0.98419 | 0.98366 | 0.97118 | 0.98431 | 0.97122 | 0.98413 | 0.98431 | 0.97176 | 0.98443 | 0.98344 | 0.97212
120 | Center | 12.93 | 20.93 | 19.23 | 22.53 | 11.6 | 15.62 | 16.67 | 11.6 | 15.92 | 14.87 | 12.98 | 13.14
| CTDIw | 12.68 | 20.52 | 18.86 | 21.91 | 11.37 | 15.19 | 16.35 | 11.37 | 15.48 | 14.59 | 12.73 | 12.78
| k | 0.98067 | 0.98041 | 0.98076 | 0.97248 | 0.98017 | 0.97247 | 0.9808 | 0.98017 | 0.97236 | 0.98117 | 0.98074 | 0.9726
140 | Center | 18.83 | 30.49 | 28.02 | 33.71 | 16.89 | 23.38 | 24.29 | 16.89 | 23.82 | 21.67 | 18.91 | 19.67
| CTDIw | 18.5 | 29.95 | 27.53 | 32.85 | 16.6 | 22.77 | 23.86 | 16.6 | 23.21 | 21.29 | 18.58 | 19.16
| k | 0.98247 | 0.98229 | 0.98251 | 0.97449 | 0.98283 | 0.97391 | 0.9823 | 0.98283 | 0.97439 | 0.98246 | 0.98255 | 0.97407
100Sn | Center
CTDIw
k
120Sn | Center
CTDIw
k
140Sn | Center
CTDIw
k
32 cm phantom, Standard shaped filter (Adult Body)
| Collimation | 2x1 | 6x0.6 | 8x0.6 | 8x0.6 | 1x5 | 10x0.6 | 6x1.2 | 1x10 | 20x0.6 | 12x1.2 | 32x0.6 | 64x0.6
| Total | 2 | 3.6 | 4.8 | 4.8 | 5 | 6 | 7.2 | 10 | 12 | 14.4 | 38.4 | 38.4
| Acquisition | 2x1 | 6x0.6 | 8x0.6 | 16x0.6 | 1x5 | 20x0.6 | 6x1.2 | 1x10 | 40x0.6 | 12x1.2 | 32x1.2 | 28.x.6
| Shaped filter |  |  |  | z-sharp |  | z-sharp |  |  | z-sharp |  |  | z-sharp
kV
70 | Center | 0.61 | 0.99 | 0.91 | 1.07 | 0.55 | 0.74 | 0.79 | 0.55 | 0.75 | 0.7 | 0.61 | 0.62
| CTDIw | 1.13 | 1.83 | 1.68 | 1.97 | 1.01 | 1.37 | 1.46 | 1.01 | 1.4 | 1.3 | 1.14 | 1.15
| k | 1.8525 | 1.8485 | 1.8462 | 1.8411 | 1.8364 | 1.8514 | 1.8481 | 1.8364 | 1.8667 | 1.8571 | 1.8689 | 1.8548
80 | Center | 1.07 | 1.73 | 1.59 | 1.83 | 0.96 | 1.27 | 1.38 | 0.96 | 1.29 | 1.23 | 1.08 | 1.07
| CTDIw | 1.86 | 3.01 | 2.77 | 3.18 | 1.67 | 2.21 | 2.4 | 1.67 | 2.25 | 2.14 | 1.87 | 1.86
| k | 1.7383 | 1.7399 | 1.7421 | 1.7377 | 1.7396 | 1.7402 | 1.7391 | 1.7396 | 1.7442 | 1.7398 | 1.7315 | 1.7383
100 | Center | 2.43 | 3.94 | 3.62 | 4.19 | 2.18 | 2.91 | 3.14 | 2.18 | 2.96 | 2.8 | 2.44 | 2.45
| CTDIw | 3.9 | 6.32 | 5.81 | 6.75 | 3.5 | 4.68 | 5.03 | 3.5 | 4.77 | 4.49 | 3.92 | 3.94
| k | 1.6049 | 1.6041 | 1.605 | 1.611 | 1.6055 | 1.6082 | 1.6019 | 1.6055 | 1.6115 | 1.6036 | 1.6066 | 1.6082
120 | Center | 4.28 | 6.93 | 6.37 | 7.5 | 3.84 | 5 | 5.52 | 3.84 | 5.3 | 4.93 | 4.3 | 4.37
| CTDIw | 6.61 | 10.7 | 9.84 | 11.55 | 5.93 | 8.01 | 8.53 | 5.93 | 8.16 | 7.61 | 6.64 | 6.74
| k | 1.5444 | 1.544 | 1.5447 | 1.54 | 1.5443 | 1.602 | 1.5453 | 1.5443 | 1.5396 | 1.5436 | 1.5442 | 1.5423
140 | Center | 6.57 | 10.64 | 9.78 | 11.59 | 5.89 | 8.03 | 8.47 | 5.89 | 8.19 | 7.56 | 6.6 | 6.76
| CTDIw | 9.94 | 16.09 | 14.79 | 17.62 | 8.91 | 12.22 | 12.81 | 8.91 | 12.45 | 11.43 | 9.98 | 10.28
| k | 1.5129 | 1.5122 | 1.5123 | 1.5203 | 1.5127 | 1.5218 | 1.5124 | 1.5127 | 1.5201 | 1.5119 | 1.5121 | 1.5207
100Sn | Center
CTDIw
k
120Sn | Center
CTDIw
k
140Sn | Center
CTDIw
k
32 cm phantom, Narrow shaped filter (Cardiac, Ped Body)
| Collimation | 2x1 | 6x0.6 | 8x0.6 | 8x0.6 | 1x5 | 10x0.6 | 6x1.2 | 1x10 | 20x0.6 | 12x1.2 | 32x0.6 | 64x0.6
| Total | 2 | 3.6 | 4.8 | 4.8 | 5 | 6 | 7.2 | 10 | 12 | 14.4 | 38.4 | 38.4
| Acquisition | 2x1 | 6x0.6 | 8x0.6 | 16x0.6 | 1x5 | 20x0.6 | 6x1.2 | 1x10 | 40x0.6 | 12x1.2 | 32x1.2 | 28.x.6
| Shaped filter |  |  |  | z-sharp |  | z-sharp |  |  | z-sharp |  |  | z-sharp
kV
70 | Center | 0.54 | 0.87 | 0.8 | 0.93 | 0.48 | 0.65 | 0.69 | 0.48 | 0.66 | 0.62 | 0.54 | 0.55
| CTDIw | 0.9 | 1.45 | 1.33 | 1.54 | 0.8 | 1.07 | 1.16 | 0.8 | 1.09 | 1.03 | 0.9 | 0.9
| k | 1.6667 | 1.6667 | 1.6625 | 1.6559 | 1.6667 | 1.6462 | 1.6812 | 1.6667 | 1.6515 | 1.6613 | 1.6667 | 1.6364
80 | Center | 0.94 | 1.52 | 1.39 | 1.62 | 0.84 | 1.12 | 1.21 | 0.84 | 1.15 | 1.08 | 0.94 | 0.95
| CTDIw | 1.46 | 2.36 | 2.17 | 2.48 | 1.31 | 1.72 | 1.88 | 1.31 | 1.76 | 1.68 | 1.47 | 1.45
| k | 1.5532 | 1.5526 | 1.5612 | 1.5309 | 1.5595 | 1.5357 | 1.5537 | 1.5595 | 1.5304 | 1.5556 | 1.5638 | 1.5263
100 | Center | 2.17 | 3.51 | 3.22 | 3.75 | 1.94 | 2.6 | 2.79 | 1.94 | 2.65 | 2.49 | 2.17 | 2.19
| CTDIw | 3.12 | 5.06 | 4.65 | 5.34 | 2.8 | 3.7 | 4.03 | 2.8 | 3.78 | 3.59 | 3.14 | 3.12
| k | 1.4378 | 1.4416 | 1.4441 | 1.424 | 1.4433 | 1.4231 | 1.4444 | 1.4433 | 1.4264 | 1.4418 | 1.447 | 1.4247
120 | Center | 3.84 | 6.22 | 5.72 | 6.67 | 3.45 | 4.62 | 4.95 | 3.45 | 4.71 | 4.42 | 3.86 | 3.89
| CTDIw | 5.36 | 8.68 | 7.98 | 9.21 | 4.81 | 6.39 | 6.91 | 4.81 | 6.51 | 6.17 | 5.39 | 5.38
| k | 1.3958 | 1.3955 | 1.3951 | 1.3808 | 1.3942 | 1.3831 | 1.396 | 1.3942 | 1.3822 | 1.3959 | 1.3964 | 1.383
140 | Center | 5.89 | 9.53 | 8.76 | 10.45 | 5.28 | 7.25 | 7.59 | 5.28 | 7.38 | 6.77 | 5.91 | 6.1
| CTDIw | 8.03 | 13 | 11.95 | 14.2 | 7.2 | 9.84 | 10.36 | 7.2 | 10.03 | 9.24 | 8.07 | 8.28
| k | 1.3633 | 1.3641 | 1.3642 | 1.3589 | 1.3636 | 1.3572 | 1.365 | 1.3636 | 1.3591 | 1.3648 | 1.3655 | 1.3574
100Sn | Center
CTDIw
k
120Sn | Center
CTDIw
k
140Sn | Center
CTDIw
k
16 cm phantom Standard shaped filter TwinBeam Dual Energy
| Collimation | 1x10 | 64x0.6
| Total | 10 | 38.4
kV
120 | Center | 3.68 | 4.28
| CTDIw | 3.7 | 4.29
| k | 1.0054 | 1.0023
16 cm phantom Narrow shaped filter TwinBeam Dual Energy
| Collimation | 1x10 | 64x0.6
| Total | 10 | 38.4
kV
120 | Center | 3.44 | 3.99
| CTDIw | 3.35 | 3.89
| k | 0.97384 | 0.97494
32 cm phantom Standard shaped filter TwinBeam Dual Energy
| Collimation | 1x10 | 64x0.6
| Total | 10 | 38.4
kV
120 | Center | 1.25 | 1.45
| CTDIw | 1.84 | 2.13
| k | 1.472 | 1.469
32 cm phantom Narrow shaped filter TwinBeam Dual Energy
| Collimation | 1x10 | 64x0.6
| Total | 10 | 38.4
kV
120 | Center | 1.12 | 1.3
| CTDIw | 1.5 | 1.74
| k | 1.3393 | 1.3385

## Definition AS 64
16 cm phantom, Standard shaped filter (Routine head)
| Collimation | 2x1 | 6x0.6 | 8x0.6 | 8x0.6 | 1x5 | 10x0.6 | 6x1.2 | 1x10 | 20x0.6 | 12x1.2 | 32x0.6 | 64x0.6
| Total | 2 | 3.6 | 4.8 | 4.8 | 5 | 6 | 7.2 | 10 | 12 | 14.4 | 38.4 | 38.4
| Acquisition | 2x1 | 6x0.6 | 8x0.6 | 16x0.6 | 1x5 | 20x0.6 | 6x1.2 | 1x10 | 40x0.6 | 12x1.2 | 32x1.2 | 128x0.6
|  |  | Topogram | UHR | UHR, z-sharp |  | z-sharp |  |  | z-sharp |  |  | z-sharp
kV
70 | Center | 2.6 | 4.22 | 3.88 | 4.58 | 2.34 | 3.17 | 3.3 | 2.34 | 3.23 | 3 | 3.01 | 2.87
| CTDIw | 2.74 | 4.43 | 4.07 | 4.78 | 2.46 | 3.31 | 3.35 | 2.46 | 3.38 | 3.15 | 3.15 | 3.01
| k | 1.0538 | 1.0498 | 1.049 | 1.0437 | 1.0513 | 1.0442 | 1.0152 | 1.0513 | 1.0464 | 1.05 | 1.0465 | 1.0488
80 | Center | 4.18 | 6.76 | 6.22 | 7.24 | 3.75 | 5.02 | 5.39 | 3.75 | 5.12 | 4.81 | 4.77 | 4.6
| CTDIw | 4.36 | 7.06 | 6.49 | 7.46 | 3.91 | 5.17 | 5.62 | 3.91 | 5.27 | 5.02 | 4.91 | 4.8
| k | 1.0431 | 1.0444 | 1.0434 | 1.0304 | 1.0427 | 1.0299 | 1.0427 | 1.0427 | 1.0293 | 1.0437 | 1.0294 | 1.0435
100 | Center | 8.44 | 13.67 | 12.56 | 14.69 | 7.57 | 10.19 | 10.89 | 7.57 | 10.38 | 9.71 | 8.48 | 8.57
| CTDIw | 8.65 | 14.01 | 12.88 | 14.93 | 7.76 | 10.35 | 11.16 | 7.76 | 10.55 | 9.96 | 8.69 | 8.71
| k | 1.0249 | 1.0249 | 1.0255 | 1.0163 | 1.0251 | 1.0157 | 1.0248 | 1.0251 | 1.0164 | 1.0257 | 1.0248 | 1.0163
120 | Center | 13.85 | 22.42 | 20.6 | 24.33 | 12.42 | 16.87 | 17.86 | 12.42 | 17.19 | 15.93 | 13.91 | 14.19
| CTDIw | 14.11 | 22.85 | 21 | 24.63 | 12.66 | 17.08 | 18.2 | 12.66 | 17.41 | 16.24 | 14.17 | 14.37
| k | 1.0188 | 1.0192 | 1.0194 | 1.0123 | 1.0193 | 1.0124 | 1.019 | 1.0193 | 1.0128 | 1.0195 | 1.0187 | 1.0127
140 | Center | 20.24 | 32.77 | 30.12 | 36.16 | 18.16 | 25.07 | 26.11 | 18.16 | 25.55 | 23.3 | 20.33 | 21.09
| CTDIw | 20.6 | 33.35 | 30.65 | 36.6 | 18.48 | 25.4 | 26.56 | 18.48 | 25.89 | 23.7 | 20.69 | 21.37
| k | 1.0178 | 1.0177 | 1.0176 | 1.0122 | 1.0176 | 1.0132 | 1.0172 | 1.0176 | 1.0133 | 1.0172 | 1.0177 | 1.0133
100Sn | Center
CTDIw
k
120Sn | Center
CTDIw
k
140Sn | Center
CTDIw
k
16 cm phantom, Narrow shaped filter (Pediatric body)
| Collimation | 2x1 | 6x0.6 | 8x0.6 | 8x0.6 | 1x5 | 10x0.6 | 6x1.2 | 1x10 | 20x0.6 | 12x1.2 | 32x0.6 | 64x0.6
| Total | 2 | 3.6 | 4.8 | 4.8 | 5 | 6 | 7.2 | 10 | 12 | 14.4 | 38.4 | 38.4
| Acquisition | 2x1 | 6x0.6 | 8x0.6 | 16x0.6 | 1x5 | 20x0.6 | 6x1.2 | 1x10 | 40x0.6 | 12x1.2 | 32x1.2 | 28.x.6
| Shaped filter |  |  |  | z-sharp |  | z-sharp |  |  | z-sharp |  |  | z-sharp
kV
70 | Center | 2.36 | 3.83 | 3.52 | 4.12 | 2.12 | 2.86 | 3.05 | 2.12 | 2.91 | 2.72 | 2.38 | 2.4
| CTDIw | 2.36 | 3.81 | 3.51 | 4.07 | 2.11 | 2.82 | 3.04 | 2.11 | 2.88 | 2.71 | 2.37 | 2.37
| k | 1 | 0.99478 | 0.99716 | 0.98786 | 0.99528 | 0.98601 | 0.99672 | 0.99528 | 0.98969 | 0.99632 | 0.9958 | 0.9875
80 | Center | 3.81 | 6.17 | 5.67 | 6.58 | 3.42 | 4.56 | 4.92 | 3.42 | 4.65 | 4.39 | 3.83 | 3.84
| CTDIw | 3.78 | 6.12 | 5.63 | 6.44 | 3.39 | 4.47 | 4.88 | 3.39 | 4.55 | 4.35 | 3.8 | 3.76
| k | 0.99213 | 0.9919 | 0.99295 | 0.97872 | 0.99123 | 0.98026 | 0.99187 | 0.99123 | 0.97849 | 0.99089 | 0.99217 | 0.97917
100 | Center | 7.81 | 12.65 | 11.63 | 13.53 | 7.01 | 9.38 | 10.08 | 7.01 | 9.56 | 8.99 | 7.85 | 7.89
| CTDIw | 7.69 | 12.45 | 11.44 | 13.14 | 6.9 | 9.11 | 9.92 | 6.9 | 9.29 | 8.85 | 7.72 | 7.67
| k | 0.98464 | 0.98419 | 0.98366 | 0.97118 | 0.98431 | 0.97122 | 0.98413 | 0.98431 | 0.97176 | 0.98443 | 0.98344 | 0.97212
120 | Center | 12.93 | 20.93 | 19.23 | 22.53 | 11.6 | 15.62 | 16.67 | 11.6 | 15.92 | 14.87 | 12.98 | 13.14
| CTDIw | 12.68 | 20.52 | 18.86 | 21.91 | 11.37 | 15.19 | 16.35 | 11.37 | 15.48 | 14.59 | 12.73 | 12.78
| k | 0.98067 | 0.98041 | 0.98076 | 0.97248 | 0.98017 | 0.97247 | 0.9808 | 0.98017 | 0.97236 | 0.98117 | 0.98074 | 0.9726
140 | Center | 18.83 | 30.49 | 28.02 | 33.71 | 16.89 | 23.38 | 24.29 | 16.89 | 23.82 | 21.67 | 18.91 | 19.67
| CTDIw | 18.5 | 29.95 | 27.53 | 32.85 | 16.6 | 22.77 | 23.86 | 16.6 | 23.21 | 21.29 | 18.58 | 19.16
| k | 0.98247 | 0.98229 | 0.98251 | 0.97449 | 0.98283 | 0.97391 | 0.9823 | 0.98283 | 0.97439 | 0.98246 | 0.98255 | 0.97407
100Sn | Center
CTDIw
k
120Sn | Center
CTDIw
k
140Sn | Center
CTDIw
k
32 cm phantom, Standard shaped filter (Adult Body)
| Collimation | 2x1 | 6x0.6 | 8x0.6 | 8x0.6 | 1x5 | 10x0.6 | 6x1.2 | 1x10 | 20x0.6 | 12x1.2 | 32x0.6 | 64x0.6
| Total | 2 | 3.6 | 4.8 | 4.8 | 5 | 6 | 7.2 | 10 | 12 | 14.4 | 38.4 | 38.4
| Acquisition | 2x1 | 6x0.6 | 8x0.6 | 16x0.6 | 1x5 | 20x0.6 | 6x1.2 | 1x10 | 40x0.6 | 12x1.2 | 32x1.2 | 28.x.6
| Shaped filter |  |  |  | z-sharp |  | z-sharp |  |  | z-sharp |  |  | z-sharp
kV
70 | Center | 0.61 | 0.99 | 0.91 | 1.07 | 0.55 | 0.74 | 0.79 | 0.55 | 0.75 | 0.7 | 0.61 | 0.62
| CTDIw | 1.13 | 1.83 | 1.68 | 1.97 | 1.01 | 1.37 | 1.46 | 1.01 | 1.4 | 1.3 | 1.14 | 1.15
| k | 1.8525 | 1.8485 | 1.8462 | 1.8411 | 1.8364 | 1.8514 | 1.8481 | 1.8364 | 1.8667 | 1.8571 | 1.8689 | 1.8548
80 | Center | 1.07 | 1.73 | 1.59 | 1.83 | 0.96 | 1.27 | 1.38 | 0.96 | 1.29 | 1.23 | 1.08 | 1.07
| CTDIw | 1.86 | 3.01 | 2.77 | 3.18 | 1.67 | 2.21 | 2.4 | 1.67 | 2.25 | 2.14 | 1.87 | 1.86
| k | 1.7383 | 1.7399 | 1.7421 | 1.7377 | 1.7396 | 1.7402 | 1.7391 | 1.7396 | 1.7442 | 1.7398 | 1.7315 | 1.7383
100 | Center | 2.43 | 3.94 | 3.62 | 4.19 | 2.18 | 2.91 | 3.14 | 2.18 | 2.96 | 2.8 | 2.44 | 2.45
| CTDIw | 3.9 | 6.32 | 5.81 | 6.75 | 3.5 | 4.68 | 5.03 | 3.5 | 4.77 | 4.49 | 3.92 | 3.94
| k | 1.6049 | 1.6041 | 1.605 | 1.611 | 1.6055 | 1.6082 | 1.6019 | 1.6055 | 1.6115 | 1.6036 | 1.6066 | 1.6082
120 | Center | 4.28 | 6.93 | 6.37 | 7.5 | 3.84 | 5 | 5.52 | 3.84 | 5.3 | 4.93 | 4.3 | 4.37
| CTDIw | 6.61 | 10.7 | 9.84 | 11.55 | 5.93 | 8.01 | 8.53 | 5.93 | 8.16 | 7.61 | 6.64 | 6.74
| k | 1.5444 | 1.544 | 1.5447 | 1.54 | 1.5443 | 1.602 | 1.5453 | 1.5443 | 1.5396 | 1.5436 | 1.5442 | 1.5423
140 | Center | 6.57 | 10.64 | 9.78 | 11.59 | 5.89 | 8.03 | 8.47 | 5.89 | 8.19 | 7.56 | 6.6 | 6.76
| CTDIw | 9.94 | 16.09 | 14.79 | 17.62 | 8.91 | 12.22 | 12.81 | 8.91 | 12.45 | 11.43 | 9.98 | 10.28
| k | 1.5129 | 1.5122 | 1.5123 | 1.5203 | 1.5127 | 1.5218 | 1.5124 | 1.5127 | 1.5201 | 1.5119 | 1.5121 | 1.5207
100Sn | Center
CTDIw
k
120Sn | Center
CTDIw
k
140Sn | Center
CTDIw
k
32 cm phantom, Narrow shaped filter (Cardiac, Ped Body)
| Collimation | 2x1 | 6x0.6 | 8x0.6 | 8x0.6 | 1x5 | 10x0.6 | 6x1.2 | 1x10 | 20x0.6 | 12x1.2 | 32x0.6 | 64x0.6
| Total | 2 | 3.6 | 4.8 | 4.8 | 5 | 6 | 7.2 | 10 | 12 | 14.4 | 38.4 | 38.4
| Acquisition | 2x1 | 6x0.6 | 8x0.6 | 16x0.6 | 1x5 | 20x0.6 | 6x1.2 | 1x10 | 40x0.6 | 12x1.2 | 32x1.2 | 28.x.6
| Shaped filter |  |  |  | z-sharp |  | z-sharp |  |  | z-sharp |  |  | z-sharp
kV
70 | Center | 0.54 | 0.87 | 0.8 | 0.93 | 0.48 | 0.65 | 0.69 | 0.48 | 0.66 | 0.62 | 0.54 | 0.55
| CTDIw | 0.9 | 1.45 | 1.33 | 1.54 | 0.8 | 1.07 | 1.16 | 0.8 | 1.09 | 1.03 | 0.9 | 0.9
| k | 1.6667 | 1.6667 | 1.6625 | 1.6559 | 1.6667 | 1.6462 | 1.6812 | 1.6667 | 1.6515 | 1.6613 | 1.6667 | 1.6364
80 | Center | 0.94 | 1.52 | 1.39 | 1.62 | 0.84 | 1.12 | 1.21 | 0.84 | 1.15 | 1.08 | 0.94 | 0.95
| CTDIw | 1.46 | 2.36 | 2.17 | 2.48 | 1.31 | 1.72 | 1.88 | 1.31 | 1.76 | 1.68 | 1.47 | 1.45
| k | 1.5532 | 1.5526 | 1.5612 | 1.5309 | 1.5595 | 1.5357 | 1.5537 | 1.5595 | 1.5304 | 1.5556 | 1.5638 | 1.5263
100 | Center | 2.17 | 3.51 | 3.22 | 3.75 | 1.94 | 2.6 | 2.79 | 1.94 | 2.65 | 2.49 | 2.17 | 2.19
| CTDIw | 3.12 | 5.06 | 4.65 | 5.34 | 2.8 | 3.7 | 4.03 | 2.8 | 3.78 | 3.59 | 3.14 | 3.12
| k | 1.4378 | 1.4416 | 1.4441 | 1.424 | 1.4433 | 1.4231 | 1.4444 | 1.4433 | 1.4264 | 1.4418 | 1.447 | 1.4247
120 | Center | 3.84 | 6.22 | 5.72 | 6.67 | 3.45 | 4.62 | 4.95 | 3.45 | 4.71 | 4.42 | 3.86 | 3.89
| CTDIw | 5.36 | 8.68 | 7.98 | 9.21 | 4.81 | 6.39 | 6.91 | 4.81 | 6.51 | 6.17 | 5.39 | 5.38
| k | 1.3958 | 1.3955 | 1.3951 | 1.3808 | 1.3942 | 1.3831 | 1.396 | 1.3942 | 1.3822 | 1.3959 | 1.3964 | 1.383
140 | Center | 5.89 | 9.53 | 8.76 | 10.45 | 5.28 | 7.25 | 7.59 | 5.28 | 7.38 | 6.77 | 5.91 | 6.1
| CTDIw | 8.03 | 13 | 11.95 | 14.2 | 7.2 | 9.84 | 10.36 | 7.2 | 10.03 | 9.24 | 8.07 | 8.28
| k | 1.3633 | 1.3641 | 1.3642 | 1.3589 | 1.3636 | 1.3572 | 1.365 | 1.3636 | 1.3591 | 1.3648 | 1.3655 | 1.3574
100Sn | Center
CTDIw
k
120Sn | Center
CTDIw
k
140Sn | Center
CTDIw
k

## Definition AS Open 64
16 cm phantom, Standard shaped filter (Routine head)
| Collimation | 2x1 | 6x0.6 | 8x0.6 | 8x0.6 | 1x5 | 10x0.6 | 6x1.2 | 1x10 | 20x0.6 | 12x1.2 | 32x0.6 | 64x0.6
| Total | 2 | 3.6 | 4.8 | 4.8 | 5 | 6 | 7.2 | 10 | 12 | 14.4 | 38.4 | 38.4
| Acquisition | 2x1 | 6x0.6 | 8x0.6 | 16x0.6 | 1x5 | 20x0.6 | 6x1.2 | 1x10 | 40x0.6 | 12x1.2 | 32x1.2 | 128x0.6
|  |  | Topogram | UHR | UHR, z-sharp |  | z-sharp |  |  | z-sharp |  |  | z-sharp
kV
70 | Center | 2.6 | 4.22 | 3.88 | 4.58 | 2.34 | 3.17 | 3.3 | 2.34 | 3.23 | 3 | 3.01 | 2.87
| CTDIw | 2.74 | 4.43 | 4.07 | 4.78 | 2.46 | 3.31 | 3.35 | 2.46 | 3.38 | 3.15 | 3.15 | 3.01
| k | 1.0538 | 1.0498 | 1.049 | 1.0437 | 1.0513 | 1.0442 | 1.0152 | 1.0513 | 1.0464 | 1.05 | 1.0465 | 1.0488
80 | Center | 4.18 | 6.76 | 6.22 | 7.24 | 3.75 | 5.02 | 5.39 | 3.75 | 5.12 | 4.81 | 4.77 | 4.6
| CTDIw | 4.36 | 7.06 | 6.49 | 7.46 | 3.91 | 5.17 | 5.62 | 3.91 | 5.27 | 5.02 | 4.91 | 4.8
| k | 1.0431 | 1.0444 | 1.0434 | 1.0304 | 1.0427 | 1.0299 | 1.0427 | 1.0427 | 1.0293 | 1.0437 | 1.0294 | 1.0435
100 | Center | 8.44 | 13.67 | 12.56 | 14.69 | 7.57 | 10.19 | 10.89 | 7.57 | 10.38 | 9.71 | 8.48 | 8.57
| CTDIw | 8.65 | 14.01 | 12.88 | 14.93 | 7.76 | 10.35 | 11.16 | 7.76 | 10.55 | 9.96 | 8.69 | 8.71
| k | 1.0249 | 1.0249 | 1.0255 | 1.0163 | 1.0251 | 1.0157 | 1.0248 | 1.0251 | 1.0164 | 1.0257 | 1.0248 | 1.0163
120 | Center | 13.85 | 22.42 | 20.6 | 24.33 | 12.42 | 16.87 | 17.86 | 12.42 | 17.19 | 15.93 | 13.91 | 14.19
| CTDIw | 14.11 | 22.85 | 21 | 24.63 | 12.66 | 17.08 | 18.2 | 12.66 | 17.41 | 16.24 | 14.17 | 14.37
| k | 1.0188 | 1.0192 | 1.0194 | 1.0123 | 1.0193 | 1.0124 | 1.019 | 1.0193 | 1.0128 | 1.0195 | 1.0187 | 1.0127
140 | Center | 20.24 | 32.77 | 30.12 | 36.16 | 18.16 | 25.07 | 26.11 | 18.16 | 25.55 | 23.3 | 20.33 | 21.09
| CTDIw | 20.6 | 33.35 | 30.65 | 36.6 | 18.48 | 25.4 | 26.56 | 18.48 | 25.89 | 23.7 | 20.69 | 21.37
| k | 1.0178 | 1.0177 | 1.0176 | 1.0122 | 1.0176 | 1.0132 | 1.0172 | 1.0176 | 1.0133 | 1.0172 | 1.0177 | 1.0133
100Sn | Center
CTDIw
k
120Sn | Center
CTDIw
k
140Sn | Center
CTDIw
k
16 cm phantom, Narrow shaped filter (Pediatric body)
| Collimation | 2x1 | 6x0.6 | 8x0.6 | 8x0.6 | 1x5 | 10x0.6 | 6x1.2 | 1x10 | 20x0.6 | 12x1.2 | 32x0.6 | 64x0.6
| Total | 2 | 3.6 | 4.8 | 4.8 | 5 | 6 | 7.2 | 10 | 12 | 14.4 | 38.4 | 38.4
| Acquisition | 2x1 | 6x0.6 | 8x0.6 | 16x0.6 | 1x5 | 20x0.6 | 6x1.2 | 1x10 | 40x0.6 | 12x1.2 | 32x1.2 | 28.x.6
| Shaped filter |  |  |  | z-sharp |  | z-sharp |  |  | z-sharp |  |  | z-sharp
kV
70 | Center | 2.36 | 3.83 | 3.52 | 4.12 | 2.12 | 2.86 | 3.05 | 2.12 | 2.91 | 2.72 | 2.38 | 2.4
| CTDIw | 2.36 | 3.81 | 3.51 | 4.07 | 2.11 | 2.82 | 3.04 | 2.11 | 2.88 | 2.71 | 2.37 | 2.37
| k | 1 | 0.99478 | 0.99716 | 0.98786 | 0.99528 | 0.98601 | 0.99672 | 0.99528 | 0.98969 | 0.99632 | 0.9958 | 0.9875
80 | Center | 3.81 | 6.17 | 5.67 | 6.58 | 3.42 | 4.56 | 4.92 | 3.42 | 4.65 | 4.39 | 3.83 | 3.84
| CTDIw | 3.78 | 6.12 | 5.63 | 6.44 | 3.39 | 4.47 | 4.88 | 3.39 | 4.55 | 4.35 | 3.8 | 3.76
| k | 0.99213 | 0.9919 | 0.99295 | 0.97872 | 0.99123 | 0.98026 | 0.99187 | 0.99123 | 0.97849 | 0.99089 | 0.99217 | 0.97917
100 | Center | 7.81 | 12.65 | 11.63 | 13.53 | 7.01 | 9.38 | 10.08 | 7.01 | 9.56 | 8.99 | 7.85 | 7.89
| CTDIw | 7.69 | 12.45 | 11.44 | 13.14 | 6.9 | 9.11 | 9.92 | 6.9 | 9.29 | 8.85 | 7.72 | 7.67
| k | 0.98464 | 0.98419 | 0.98366 | 0.97118 | 0.98431 | 0.97122 | 0.98413 | 0.98431 | 0.97176 | 0.98443 | 0.98344 | 0.97212
120 | Center | 12.93 | 20.93 | 19.23 | 22.53 | 11.6 | 15.62 | 16.67 | 11.6 | 15.92 | 14.87 | 12.98 | 13.14
| CTDIw | 12.68 | 20.52 | 18.86 | 21.91 | 11.37 | 15.19 | 16.35 | 11.37 | 15.48 | 14.59 | 12.73 | 12.78
| k | 0.98067 | 0.98041 | 0.98076 | 0.97248 | 0.98017 | 0.97247 | 0.9808 | 0.98017 | 0.97236 | 0.98117 | 0.98074 | 0.9726
140 | Center | 18.83 | 30.49 | 28.02 | 33.71 | 16.89 | 23.38 | 24.29 | 16.89 | 23.82 | 21.67 | 18.91 | 19.67
| CTDIw | 18.5 | 29.95 | 27.53 | 32.85 | 16.6 | 22.77 | 23.86 | 16.6 | 23.21 | 21.29 | 18.58 | 19.16
| k | 0.98247 | 0.98229 | 0.98251 | 0.97449 | 0.98283 | 0.97391 | 0.9823 | 0.98283 | 0.97439 | 0.98246 | 0.98255 | 0.97407
100Sn | Center
CTDIw
k
120Sn | Center
CTDIw
k
140Sn | Center
CTDIw
k
32 cm phantom, Standard shaped filter (Adult Body)
| Collimation | 2x1 | 6x0.6 | 8x0.6 | 8x0.6 | 1x5 | 10x0.6 | 6x1.2 | 1x10 | 20x0.6 | 12x1.2 | 32x0.6 | 64x0.6
| Total | 2 | 3.6 | 4.8 | 4.8 | 5 | 6 | 7.2 | 10 | 12 | 14.4 | 38.4 | 38.4
| Acquisition | 2x1 | 6x0.6 | 8x0.6 | 16x0.6 | 1x5 | 20x0.6 | 6x1.2 | 1x10 | 40x0.6 | 12x1.2 | 32x1.2 | 28.x.6
| Shaped filter |  |  |  | z-sharp |  | z-sharp |  |  | z-sharp |  |  | z-sharp
kV
70 | Center | 0.61 | 0.99 | 0.91 | 1.07 | 0.55 | 0.74 | 0.79 | 0.55 | 0.75 | 0.7 | 0.61 | 0.62
| CTDIw | 1.13 | 1.83 | 1.68 | 1.97 | 1.01 | 1.37 | 1.46 | 1.01 | 1.4 | 1.3 | 1.14 | 1.15
| k | 1.8525 | 1.8485 | 1.8462 | 1.8411 | 1.8364 | 1.8514 | 1.8481 | 1.8364 | 1.8667 | 1.8571 | 1.8689 | 1.8548
80 | Center | 1.07 | 1.73 | 1.59 | 1.83 | 0.96 | 1.27 | 1.38 | 0.96 | 1.29 | 1.23 | 1.08 | 1.07
| CTDIw | 1.86 | 3.01 | 2.77 | 3.18 | 1.67 | 2.21 | 2.4 | 1.67 | 2.25 | 2.14 | 1.87 | 1.86
| k | 1.7383 | 1.7399 | 1.7421 | 1.7377 | 1.7396 | 1.7402 | 1.7391 | 1.7396 | 1.7442 | 1.7398 | 1.7315 | 1.7383
100 | Center | 2.43 | 3.94 | 3.62 | 4.19 | 2.18 | 2.91 | 3.14 | 2.18 | 2.96 | 2.8 | 2.44 | 2.45
| CTDIw | 3.9 | 6.32 | 5.81 | 6.75 | 3.5 | 4.68 | 5.03 | 3.5 | 4.77 | 4.49 | 3.92 | 3.94
| k | 1.6049 | 1.6041 | 1.605 | 1.611 | 1.6055 | 1.6082 | 1.6019 | 1.6055 | 1.6115 | 1.6036 | 1.6066 | 1.6082
120 | Center | 4.28 | 6.93 | 6.37 | 7.5 | 3.84 | 5 | 5.52 | 3.84 | 5.3 | 4.93 | 4.3 | 4.37
| CTDIw | 6.61 | 10.7 | 9.84 | 11.55 | 5.93 | 8.01 | 8.53 | 5.93 | 8.16 | 7.61 | 6.64 | 6.74
| k | 1.5444 | 1.544 | 1.5447 | 1.54 | 1.5443 | 1.602 | 1.5453 | 1.5443 | 1.5396 | 1.5436 | 1.5442 | 1.5423
140 | Center | 6.57 | 10.64 | 9.78 | 11.59 | 5.89 | 8.03 | 8.47 | 5.89 | 8.19 | 7.56 | 6.6 | 6.76
| CTDIw | 9.94 | 16.09 | 14.79 | 17.62 | 8.91 | 12.22 | 12.81 | 8.91 | 12.45 | 11.43 | 9.98 | 10.28
| k | 1.5129 | 1.5122 | 1.5123 | 1.5203 | 1.5127 | 1.5218 | 1.5124 | 1.5127 | 1.5201 | 1.5119 | 1.5121 | 1.5207
100Sn | Center
CTDIw
k
120Sn | Center
CTDIw
k
140Sn | Center
CTDIw
k
32 cm phantom, Narrow shaped filter (Cardiac, Ped Body)
| Collimation | 2x1 | 6x0.6 | 8x0.6 | 8x0.6 | 1x5 | 10x0.6 | 6x1.2 | 1x10 | 20x0.6 | 12x1.2 | 32x0.6 | 64x0.6
| Total | 2 | 3.6 | 4.8 | 4.8 | 5 | 6 | 7.2 | 10 | 12 | 14.4 | 38.4 | 38.4
| Acquisition | 2x1 | 6x0.6 | 8x0.6 | 16x0.6 | 1x5 | 20x0.6 | 6x1.2 | 1x10 | 40x0.6 | 12x1.2 | 32x1.2 | 28.x.6
| Shaped filter |  |  |  | z-sharp |  | z-sharp |  |  | z-sharp |  |  | z-sharp
kV
70 | Center | 0.54 | 0.87 | 0.8 | 0.93 | 0.48 | 0.65 | 0.69 | 0.48 | 0.66 | 0.62 | 0.54 | 0.55
| CTDIw | 0.9 | 1.45 | 1.33 | 1.54 | 0.8 | 1.07 | 1.16 | 0.8 | 1.09 | 1.03 | 0.9 | 0.9
| k | 1.6667 | 1.6667 | 1.6625 | 1.6559 | 1.6667 | 1.6462 | 1.6812 | 1.6667 | 1.6515 | 1.6613 | 1.6667 | 1.6364
80 | Center | 0.94 | 1.52 | 1.39 | 1.62 | 0.84 | 1.12 | 1.21 | 0.84 | 1.15 | 1.08 | 0.94 | 0.95
| CTDIw | 1.46 | 2.36 | 2.17 | 2.48 | 1.31 | 1.72 | 1.88 | 1.31 | 1.76 | 1.68 | 1.47 | 1.45
| k | 1.5532 | 1.5526 | 1.5612 | 1.5309 | 1.5595 | 1.5357 | 1.5537 | 1.5595 | 1.5304 | 1.5556 | 1.5638 | 1.5263
100 | Center | 2.17 | 3.51 | 3.22 | 3.75 | 1.94 | 2.6 | 2.79 | 1.94 | 2.65 | 2.49 | 2.17 | 2.19
| CTDIw | 3.12 | 5.06 | 4.65 | 5.34 | 2.8 | 3.7 | 4.03 | 2.8 | 3.78 | 3.59 | 3.14 | 3.12
| k | 1.4378 | 1.4416 | 1.4441 | 1.424 | 1.4433 | 1.4231 | 1.4444 | 1.4433 | 1.4264 | 1.4418 | 1.447 | 1.4247
120 | Center | 3.84 | 6.22 | 5.72 | 6.67 | 3.45 | 4.62 | 4.95 | 3.45 | 4.71 | 4.42 | 3.86 | 3.89
| CTDIw | 5.36 | 8.68 | 7.98 | 9.21 | 4.81 | 6.39 | 6.91 | 4.81 | 6.51 | 6.17 | 5.39 | 5.38
| k | 1.3958 | 1.3955 | 1.3951 | 1.3808 | 1.3942 | 1.3831 | 1.396 | 1.3942 | 1.3822 | 1.3959 | 1.3964 | 1.383
140 | Center | 5.89 | 9.53 | 8.76 | 10.45 | 5.28 | 7.25 | 7.59 | 5.28 | 7.38 | 6.77 | 5.91 | 6.1
| CTDIw | 8.03 | 13 | 11.95 | 14.2 | 7.2 | 9.84 | 10.36 | 7.2 | 10.03 | 9.24 | 8.07 | 8.28
| k | 1.3633 | 1.3641 | 1.3642 | 1.3589 | 1.3636 | 1.3572 | 1.365 | 1.3636 | 1.3591 | 1.3648 | 1.3655 | 1.3574
100Sn | Center
CTDIw
k
120Sn | Center
CTDIw
k
140Sn | Center
CTDIw
k
16 cm phantom Standard shaped filter TwinBeam Dual Energy
| Collimation | 1x10 | 64x0.6
| Total | 10 | 38.4
kV
120 | Center | 3.68 | 4.28
| CTDIw | 3.7 | 4.29
| k | 1.0054 | 1.0023
16 cm phantom Narrow shaped filter TwinBeam Dual Energy
| Collimation | 1x10 | 64x0.6
| Total | 10 | 38.4
kV
120 | Center | 3.44 | 3.99
| CTDIw | 3.35 | 3.89
| k | 0.97384 | 0.97494
32 cm phantom Standard shaped filter TwinBeam Dual Energy
| Collimation | 1x10 | 64x0.6
| Total | 10 | 38.4
kV
120 | Center | 1.25 | 1.45
| CTDIw | 1.84 | 2.13
| k | 1.472 | 1.469
32 cm phantom Narrow shaped filter TwinBeam Dual Energy
| Collimation | 1x10 | 64x0.6
| Total | 10 | 38.4
kV
120 | Center | 1.12 | 1.3
| CTDIw | 1.5 | 1.74
| k | 1.3393 | 1.3385

## Definition AS 40
16 cm phantom, Standard shaped filter (Routine head)
| Collimation | 2x1 | 6x0.6 | 8x0.6 | 8x0.6 | 1x5 | 10x0.6 | 6x1.2 | 1x10 | 20x0.6 | 12x1.2 | 32x0.6 | 64x0.6
| Total | 2 | 3.6 | 4.8 | 4.8 | 5 | 6 | 7.2 | 10 | 12 | 14.4 | 38.4 | 38.4
| Acquisition | 2x1 | 6x0.6 | 8x0.6 | 16x0.6 | 1x5 | 20x0.6 | 6x1.2 | 1x10 | 40x0.6 | 12x1.2 | 32x1.2 | 128x0.6
|  |  | Topogram | UHR | UHR, z-sharp |  | z-sharp |  |  | z-sharp |  |  | z-sharp
kV
70 | Center | 2.6 | 4.22 | 3.88 | 4.58 | 2.34 | 3.17 | 3.3 | 2.34 | 3.23 | 3 | 3.01 | 2.87
| CTDIw | 2.74 | 4.43 | 4.07 | 4.78 | 2.46 | 3.31 | 3.35 | 2.46 | 3.38 | 3.15 | 3.15 | 3.01
| k | 1.0538 | 1.0498 | 1.049 | 1.0437 | 1.0513 | 1.0442 | 1.0152 | 1.0513 | 1.0464 | 1.05 | 1.0465 | 1.0488
80 | Center | 4.18 | 6.76 | 6.22 | 7.24 | 3.75 | 5.02 | 5.39 | 3.75 | 5.12 | 4.81 | 4.77 | 4.6
| CTDIw | 4.36 | 7.06 | 6.49 | 7.46 | 3.91 | 5.17 | 5.62 | 3.91 | 5.27 | 5.02 | 4.91 | 4.8
| k | 1.0431 | 1.0444 | 1.0434 | 1.0304 | 1.0427 | 1.0299 | 1.0427 | 1.0427 | 1.0293 | 1.0437 | 1.0294 | 1.0435
100 | Center | 8.44 | 13.67 | 12.56 | 14.69 | 7.57 | 10.19 | 10.89 | 7.57 | 10.38 | 9.71 | 8.48 | 8.57
| CTDIw | 8.65 | 14.01 | 12.88 | 14.93 | 7.76 | 10.35 | 11.16 | 7.76 | 10.55 | 9.96 | 8.69 | 8.71
| k | 1.0249 | 1.0249 | 1.0255 | 1.0163 | 1.0251 | 1.0157 | 1.0248 | 1.0251 | 1.0164 | 1.0257 | 1.0248 | 1.0163
120 | Center | 13.85 | 22.42 | 20.6 | 24.33 | 12.42 | 16.87 | 17.86 | 12.42 | 17.19 | 15.93 | 13.91 | 14.19
| CTDIw | 14.11 | 22.85 | 21 | 24.63 | 12.66 | 17.08 | 18.2 | 12.66 | 17.41 | 16.24 | 14.17 | 14.37
| k | 1.0188 | 1.0192 | 1.0194 | 1.0123 | 1.0193 | 1.0124 | 1.019 | 1.0193 | 1.0128 | 1.0195 | 1.0187 | 1.0127
140 | Center | 20.24 | 32.77 | 30.12 | 36.16 | 18.16 | 25.07 | 26.11 | 18.16 | 25.55 | 23.3 | 20.33 | 21.09
| CTDIw | 20.6 | 33.35 | 30.65 | 36.6 | 18.48 | 25.4 | 26.56 | 18.48 | 25.89 | 23.7 | 20.69 | 21.37
| k | 1.0178 | 1.0177 | 1.0176 | 1.0122 | 1.0176 | 1.0132 | 1.0172 | 1.0176 | 1.0133 | 1.0172 | 1.0177 | 1.0133
100Sn | Center
CTDIw
k
120Sn | Center
CTDIw
k
140Sn | Center
CTDIw
k
16 cm phantom, Narrow shaped filter (Pediatric body)
| Collimation | 2x1 | 6x0.6 | 8x0.6 | 8x0.6 | 1x5 | 10x0.6 | 6x1.2 | 1x10 | 20x0.6 | 12x1.2 | 32x0.6 | 64x0.6
| Total | 2 | 3.6 | 4.8 | 4.8 | 5 | 6 | 7.2 | 10 | 12 | 14.4 | 38.4 | 38.4
| Acquisition | 2x1 | 6x0.6 | 8x0.6 | 16x0.6 | 1x5 | 20x0.6 | 6x1.2 | 1x10 | 40x0.6 | 12x1.2 | 32x1.2 | 28.x.6
| Shaped filter |  |  |  | z-sharp |  | z-sharp |  |  | z-sharp |  |  | z-sharp
kV
70 | Center | 2.36 | 3.83 | 3.52 | 4.12 | 2.12 | 2.86 | 3.05 | 2.12 | 2.91 | 2.72 | 2.38 | 2.4
| CTDIw | 2.36 | 3.81 | 3.51 | 4.07 | 2.11 | 2.82 | 3.04 | 2.11 | 2.88 | 2.71 | 2.37 | 2.37
| k | 1 | 0.99478 | 0.99716 | 0.98786 | 0.99528 | 0.98601 | 0.99672 | 0.99528 | 0.98969 | 0.99632 | 0.9958 | 0.9875
80 | Center | 3.81 | 6.17 | 5.67 | 6.58 | 3.42 | 4.56 | 4.92 | 3.42 | 4.65 | 4.39 | 3.83 | 3.84
| CTDIw | 3.78 | 6.12 | 5.63 | 6.44 | 3.39 | 4.47 | 4.88 | 3.39 | 4.55 | 4.35 | 3.8 | 3.76
| k | 0.99213 | 0.9919 | 0.99295 | 0.97872 | 0.99123 | 0.98026 | 0.99187 | 0.99123 | 0.97849 | 0.99089 | 0.99217 | 0.97917
100 | Center | 7.81 | 12.65 | 11.63 | 13.53 | 7.01 | 9.38 | 10.08 | 7.01 | 9.56 | 8.99 | 7.85 | 7.89
| CTDIw | 7.69 | 12.45 | 11.44 | 13.14 | 6.9 | 9.11 | 9.92 | 6.9 | 9.29 | 8.85 | 7.72 | 7.67
| k | 0.98464 | 0.98419 | 0.98366 | 0.97118 | 0.98431 | 0.97122 | 0.98413 | 0.98431 | 0.97176 | 0.98443 | 0.98344 | 0.97212
120 | Center | 12.93 | 20.93 | 19.23 | 22.53 | 11.6 | 15.62 | 16.67 | 11.6 | 15.92 | 14.87 | 12.98 | 13.14
| CTDIw | 12.68 | 20.52 | 18.86 | 21.91 | 11.37 | 15.19 | 16.35 | 11.37 | 15.48 | 14.59 | 12.73 | 12.78
| k | 0.98067 | 0.98041 | 0.98076 | 0.97248 | 0.98017 | 0.97247 | 0.9808 | 0.98017 | 0.97236 | 0.98117 | 0.98074 | 0.9726
140 | Center | 18.83 | 30.49 | 28.02 | 33.71 | 16.89 | 23.38 | 24.29 | 16.89 | 23.82 | 21.67 | 18.91 | 19.67
| CTDIw | 18.5 | 29.95 | 27.53 | 32.85 | 16.6 | 22.77 | 23.86 | 16.6 | 23.21 | 21.29 | 18.58 | 19.16
| k | 0.98247 | 0.98229 | 0.98251 | 0.97449 | 0.98283 | 0.97391 | 0.9823 | 0.98283 | 0.97439 | 0.98246 | 0.98255 | 0.97407
100Sn | Center
CTDIw
k
120Sn | Center
CTDIw
k
140Sn | Center
CTDIw
k
32 cm phantom, Standard shaped filter (Adult Body)
| Collimation | 2x1 | 6x0.6 | 8x0.6 | 8x0.6 | 1x5 | 10x0.6 | 6x1.2 | 1x10 | 20x0.6 | 12x1.2 | 32x0.6 | 64x0.6
| Total | 2 | 3.6 | 4.8 | 4.8 | 5 | 6 | 7.2 | 10 | 12 | 14.4 | 38.4 | 38.4
| Acquisition | 2x1 | 6x0.6 | 8x0.6 | 16x0.6 | 1x5 | 20x0.6 | 6x1.2 | 1x10 | 40x0.6 | 12x1.2 | 32x1.2 | 28.x.6
| Shaped filter |  |  |  | z-sharp |  | z-sharp |  |  | z-sharp |  |  | z-sharp
kV
70 | Center | 0.61 | 0.99 | 0.91 | 1.07 | 0.55 | 0.74 | 0.79 | 0.55 | 0.75 | 0.7 | 0.61 | 0.62
| CTDIw | 1.13 | 1.83 | 1.68 | 1.97 | 1.01 | 1.37 | 1.46 | 1.01 | 1.4 | 1.3 | 1.14 | 1.15
| k | 1.8525 | 1.8485 | 1.8462 | 1.8411 | 1.8364 | 1.8514 | 1.8481 | 1.8364 | 1.8667 | 1.8571 | 1.8689 | 1.8548
80 | Center | 1.07 | 1.73 | 1.59 | 1.83 | 0.96 | 1.27 | 1.38 | 0.96 | 1.29 | 1.23 | 1.08 | 1.07
| CTDIw | 1.86 | 3.01 | 2.77 | 3.18 | 1.67 | 2.21 | 2.4 | 1.67 | 2.25 | 2.14 | 1.87 | 1.86
| k | 1.7383 | 1.7399 | 1.7421 | 1.7377 | 1.7396 | 1.7402 | 1.7391 | 1.7396 | 1.7442 | 1.7398 | 1.7315 | 1.7383
100 | Center | 2.43 | 3.94 | 3.62 | 4.19 | 2.18 | 2.91 | 3.14 | 2.18 | 2.96 | 2.8 | 2.44 | 2.45
| CTDIw | 3.9 | 6.32 | 5.81 | 6.75 | 3.5 | 4.68 | 5.03 | 3.5 | 4.77 | 4.49 | 3.92 | 3.94
| k | 1.6049 | 1.6041 | 1.605 | 1.611 | 1.6055 | 1.6082 | 1.6019 | 1.6055 | 1.6115 | 1.6036 | 1.6066 | 1.6082
120 | Center | 4.28 | 6.93 | 6.37 | 7.5 | 3.84 | 5 | 5.52 | 3.84 | 5.3 | 4.93 | 4.3 | 4.37
| CTDIw | 6.61 | 10.7 | 9.84 | 11.55 | 5.93 | 8.01 | 8.53 | 5.93 | 8.16 | 7.61 | 6.64 | 6.74
| k | 1.5444 | 1.544 | 1.5447 | 1.54 | 1.5443 | 1.602 | 1.5453 | 1.5443 | 1.5396 | 1.5436 | 1.5442 | 1.5423
140 | Center | 6.57 | 10.64 | 9.78 | 11.59 | 5.89 | 8.03 | 8.47 | 5.89 | 8.19 | 7.56 | 6.6 | 6.76
| CTDIw | 9.94 | 16.09 | 14.79 | 17.62 | 8.91 | 12.22 | 12.81 | 8.91 | 12.45 | 11.43 | 9.98 | 10.28
| k | 1.5129 | 1.5122 | 1.5123 | 1.5203 | 1.5127 | 1.5218 | 1.5124 | 1.5127 | 1.5201 | 1.5119 | 1.5121 | 1.5207
100Sn | Center
CTDIw
k
120Sn | Center
CTDIw
k
140Sn | Center
CTDIw
k
32 cm phantom, Narrow shaped filter (Cardiac, Ped Body)
| Collimation | 2x1 | 6x0.6 | 8x0.6 | 8x0.6 | 1x5 | 10x0.6 | 6x1.2 | 1x10 | 20x0.6 | 12x1.2 | 32x0.6 | 64x0.6
| Total | 2 | 3.6 | 4.8 | 4.8 | 5 | 6 | 7.2 | 10 | 12 | 14.4 | 38.4 | 38.4
| Acquisition | 2x1 | 6x0.6 | 8x0.6 | 16x0.6 | 1x5 | 20x0.6 | 6x1.2 | 1x10 | 40x0.6 | 12x1.2 | 32x1.2 | 28.x.6
| Shaped filter |  |  |  | z-sharp |  | z-sharp |  |  | z-sharp |  |  | z-sharp
kV
70 | Center | 0.54 | 0.87 | 0.8 | 0.93 | 0.48 | 0.65 | 0.69 | 0.48 | 0.66 | 0.62 | 0.54 | 0.55
| CTDIw | 0.9 | 1.45 | 1.33 | 1.54 | 0.8 | 1.07 | 1.16 | 0.8 | 1.09 | 1.03 | 0.9 | 0.9
| k | 1.6667 | 1.6667 | 1.6625 | 1.6559 | 1.6667 | 1.6462 | 1.6812 | 1.6667 | 1.6515 | 1.6613 | 1.6667 | 1.6364
80 | Center | 0.94 | 1.52 | 1.39 | 1.62 | 0.84 | 1.12 | 1.21 | 0.84 | 1.15 | 1.08 | 0.94 | 0.95
| CTDIw | 1.46 | 2.36 | 2.17 | 2.48 | 1.31 | 1.72 | 1.88 | 1.31 | 1.76 | 1.68 | 1.47 | 1.45
| k | 1.5532 | 1.5526 | 1.5612 | 1.5309 | 1.5595 | 1.5357 | 1.5537 | 1.5595 | 1.5304 | 1.5556 | 1.5638 | 1.5263
100 | Center | 2.17 | 3.51 | 3.22 | 3.75 | 1.94 | 2.6 | 2.79 | 1.94 | 2.65 | 2.49 | 2.17 | 2.19
| CTDIw | 3.12 | 5.06 | 4.65 | 5.34 | 2.8 | 3.7 | 4.03 | 2.8 | 3.78 | 3.59 | 3.14 | 3.12
| k | 1.4378 | 1.4416 | 1.4441 | 1.424 | 1.4433 | 1.4231 | 1.4444 | 1.4433 | 1.4264 | 1.4418 | 1.447 | 1.4247
120 | Center | 3.84 | 6.22 | 5.72 | 6.67 | 3.45 | 4.62 | 4.95 | 3.45 | 4.71 | 4.42 | 3.86 | 3.89
| CTDIw | 5.36 | 8.68 | 7.98 | 9.21 | 4.81 | 6.39 | 6.91 | 4.81 | 6.51 | 6.17 | 5.39 | 5.38
| k | 1.3958 | 1.3955 | 1.3951 | 1.3808 | 1.3942 | 1.3831 | 1.396 | 1.3942 | 1.3822 | 1.3959 | 1.3964 | 1.383
140 | Center | 5.89 | 9.53 | 8.76 | 10.45 | 5.28 | 7.25 | 7.59 | 5.28 | 7.38 | 6.77 | 5.91 | 6.1
| CTDIw | 8.03 | 13 | 11.95 | 14.2 | 7.2 | 9.84 | 10.36 | 7.2 | 10.03 | 9.24 | 8.07 | 8.28
| k | 1.3633 | 1.3641 | 1.3642 | 1.3589 | 1.3636 | 1.3572 | 1.365 | 1.3636 | 1.3591 | 1.3648 | 1.3655 | 1.3574
100Sn | Center
CTDIw
k
120Sn | Center
CTDIw
k
140Sn | Center
CTDIw
k
16 cm phantom Standard shaped filter TwinBeam Dual Energy
| Collimation | 1x10 | 64x0.6
| Total | 10 | 38.4
kV
120 | Center | 3.68 | 4.28
| CTDIw | 3.7 | 4.29
| k | 1.0054 | 1.0023
16 cm phantom Narrow shaped filter TwinBeam Dual Energy
| Collimation | 1x10 | 64x0.6
| Total | 10 | 38.4
kV
120 | Center | 3.44 | 3.99
| CTDIw | 3.35 | 3.89
| k | 0.97384 | 0.97494
32 cm phantom Standard shaped filter TwinBeam Dual Energy
| Collimation | 1x10 | 64x0.6
| Total | 10 | 38.4
kV
120 | Center | 1.25 | 1.45
| CTDIw | 1.84 | 2.13
| k | 1.472 | 1.469
32 cm phantom Narrow shaped filter TwinBeam Dual Energy
| Collimation | 1x10 | 64x0.6
| Total | 10 | 38.4
kV
120 | Center | 1.12 | 1.3
| CTDIw | 1.5 | 1.74
| k | 1.3393 | 1.3385

## Definition AS 20
16 cm phantom, Standard shaped filter (Routine head)
| Collimation | 2x1 | 6x0.6 | 8x0.6 | 8x0.6 | 1x5 | 10x0.6 | 6x1.2 | 1x10 | 20x0.6 | 12x1.2 | 32x0.6 | 64x0.6
| Total | 2 | 3.6 | 4.8 | 4.8 | 5 | 6 | 7.2 | 10 | 12 | 14.4 | 38.4 | 38.4
| Acquisition | 2x1 | 6x0.6 | 8x0.6 | 16x0.6 | 1x5 | 20x0.6 | 6x1.2 | 1x10 | 40x0.6 | 12x1.2 | 32x1.2 | 128x0.6
|  |  | Topogram | UHR | UHR, z-sharp |  | z-sharp |  |  | z-sharp |  |  | z-sharp
kV
70 | Center | 2.6 | 4.22 | 3.88 | 4.58 | 2.34 | 3.17 | 3.3 | 2.34 | 3.23 | 3 | 3.01 | 2.87
| CTDIw | 2.74 | 4.43 | 4.07 | 4.78 | 2.46 | 3.31 | 3.35 | 2.46 | 3.38 | 3.15 | 3.15 | 3.01
| k | 1.0538 | 1.0498 | 1.049 | 1.0437 | 1.0513 | 1.0442 | 1.0152 | 1.0513 | 1.0464 | 1.05 | 1.0465 | 1.0488
80 | Center | 4.18 | 6.76 | 6.22 | 7.24 | 3.75 | 5.02 | 5.39 | 3.75 | 5.12 | 4.81 | 4.77 | 4.6
| CTDIw | 4.36 | 7.06 | 6.49 | 7.46 | 3.91 | 5.17 | 5.62 | 3.91 | 5.27 | 5.02 | 4.91 | 4.8
| k | 1.0431 | 1.0444 | 1.0434 | 1.0304 | 1.0427 | 1.0299 | 1.0427 | 1.0427 | 1.0293 | 1.0437 | 1.0294 | 1.0435
100 | Center | 8.44 | 13.67 | 12.56 | 14.69 | 7.57 | 10.19 | 10.89 | 7.57 | 10.38 | 9.71 | 8.48 | 8.57
| CTDIw | 8.65 | 14.01 | 12.88 | 14.93 | 7.76 | 10.35 | 11.16 | 7.76 | 10.55 | 9.96 | 8.69 | 8.71
| k | 1.0249 | 1.0249 | 1.0255 | 1.0163 | 1.0251 | 1.0157 | 1.0248 | 1.0251 | 1.0164 | 1.0257 | 1.0248 | 1.0163
120 | Center | 13.85 | 22.42 | 20.6 | 24.33 | 12.42 | 16.87 | 17.86 | 12.42 | 17.19 | 15.93 | 13.91 | 14.19
| CTDIw | 14.11 | 22.85 | 21 | 24.63 | 12.66 | 17.08 | 18.2 | 12.66 | 17.41 | 16.24 | 14.17 | 14.37
| k | 1.0188 | 1.0192 | 1.0194 | 1.0123 | 1.0193 | 1.0124 | 1.019 | 1.0193 | 1.0128 | 1.0195 | 1.0187 | 1.0127
140 | Center | 20.24 | 32.77 | 30.12 | 36.16 | 18.16 | 25.07 | 26.11 | 18.16 | 25.55 | 23.3 | 20.33 | 21.09
| CTDIw | 20.6 | 33.35 | 30.65 | 36.6 | 18.48 | 25.4 | 26.56 | 18.48 | 25.89 | 23.7 | 20.69 | 21.37
| k | 1.0178 | 1.0177 | 1.0176 | 1.0122 | 1.0176 | 1.0132 | 1.0172 | 1.0176 | 1.0133 | 1.0172 | 1.0177 | 1.0133
100Sn | Center
CTDIw
k
120Sn | Center
CTDIw
k
140Sn | Center
CTDIw
k
16 cm phantom, Narrow shaped filter (Pediatric body)
| Collimation | 2x1 | 6x0.6 | 8x0.6 | 8x0.6 | 1x5 | 10x0.6 | 6x1.2 | 1x10 | 20x0.6 | 12x1.2 | 32x0.6 | 64x0.6
| Total | 2 | 3.6 | 4.8 | 4.8 | 5 | 6 | 7.2 | 10 | 12 | 14.4 | 38.4 | 38.4
| Acquisition | 2x1 | 6x0.6 | 8x0.6 | 16x0.6 | 1x5 | 20x0.6 | 6x1.2 | 1x10 | 40x0.6 | 12x1.2 | 32x1.2 | 28.x.6
| Shaped filter |  |  |  | z-sharp |  | z-sharp |  |  | z-sharp |  |  | z-sharp
kV
70 | Center | 2.36 | 3.83 | 3.52 | 4.12 | 2.12 | 2.86 | 3.05 | 2.12 | 2.91 | 2.72 | 2.38 | 2.4
| CTDIw | 2.36 | 3.81 | 3.51 | 4.07 | 2.11 | 2.82 | 3.04 | 2.11 | 2.88 | 2.71 | 2.37 | 2.37
| k | 1 | 0.99478 | 0.99716 | 0.98786 | 0.99528 | 0.98601 | 0.99672 | 0.99528 | 0.98969 | 0.99632 | 0.9958 | 0.9875
80 | Center | 3.81 | 6.17 | 5.67 | 6.58 | 3.42 | 4.56 | 4.92 | 3.42 | 4.65 | 4.39 | 3.83 | 3.84
| CTDIw | 3.78 | 6.12 | 5.63 | 6.44 | 3.39 | 4.47 | 4.88 | 3.39 | 4.55 | 4.35 | 3.8 | 3.76
| k | 0.99213 | 0.9919 | 0.99295 | 0.97872 | 0.99123 | 0.98026 | 0.99187 | 0.99123 | 0.97849 | 0.99089 | 0.99217 | 0.97917
100 | Center | 7.81 | 12.65 | 11.63 | 13.53 | 7.01 | 9.38 | 10.08 | 7.01 | 9.56 | 8.99 | 7.85 | 7.89
| CTDIw | 7.69 | 12.45 | 11.44 | 13.14 | 6.9 | 9.11 | 9.92 | 6.9 | 9.29 | 8.85 | 7.72 | 7.67
| k | 0.98464 | 0.98419 | 0.98366 | 0.97118 | 0.98431 | 0.97122 | 0.98413 | 0.98431 | 0.97176 | 0.98443 | 0.98344 | 0.97212
120 | Center | 12.93 | 20.93 | 19.23 | 22.53 | 11.6 | 15.62 | 16.67 | 11.6 | 15.92 | 14.87 | 12.98 | 13.14
| CTDIw | 12.68 | 20.52 | 18.86 | 21.91 | 11.37 | 15.19 | 16.35 | 11.37 | 15.48 | 14.59 | 12.73 | 12.78
| k | 0.98067 | 0.98041 | 0.98076 | 0.97248 | 0.98017 | 0.97247 | 0.9808 | 0.98017 | 0.97236 | 0.98117 | 0.98074 | 0.9726
140 | Center | 18.83 | 30.49 | 28.02 | 33.71 | 16.89 | 23.38 | 24.29 | 16.89 | 23.82 | 21.67 | 18.91 | 19.67
| CTDIw | 18.5 | 29.95 | 27.53 | 32.85 | 16.6 | 22.77 | 23.86 | 16.6 | 23.21 | 21.29 | 18.58 | 19.16
| k | 0.98247 | 0.98229 | 0.98251 | 0.97449 | 0.98283 | 0.97391 | 0.9823 | 0.98283 | 0.97439 | 0.98246 | 0.98255 | 0.97407
100Sn | Center
CTDIw
k
120Sn | Center
CTDIw
k
140Sn | Center
CTDIw
k
32 cm phantom, Standard shaped filter (Adult Body)
| Collimation | 2x1 | 6x0.6 | 8x0.6 | 8x0.6 | 1x5 | 10x0.6 | 6x1.2 | 1x10 | 20x0.6 | 12x1.2 | 32x0.6 | 64x0.6
| Total | 2 | 3.6 | 4.8 | 4.8 | 5 | 6 | 7.2 | 10 | 12 | 14.4 | 38.4 | 38.4
| Acquisition | 2x1 | 6x0.6 | 8x0.6 | 16x0.6 | 1x5 | 20x0.6 | 6x1.2 | 1x10 | 40x0.6 | 12x1.2 | 32x1.2 | 28.x.6
| Shaped filter |  |  |  | z-sharp |  | z-sharp |  |  | z-sharp |  |  | z-sharp
kV
70 | Center | 0.61 | 0.99 | 0.91 | 1.07 | 0.55 | 0.74 | 0.79 | 0.55 | 0.75 | 0.7 | 0.61 | 0.62
| CTDIw | 1.13 | 1.83 | 1.68 | 1.97 | 1.01 | 1.37 | 1.46 | 1.01 | 1.4 | 1.3 | 1.14 | 1.15
| k | 1.8525 | 1.8485 | 1.8462 | 1.8411 | 1.8364 | 1.8514 | 1.8481 | 1.8364 | 1.8667 | 1.8571 | 1.8689 | 1.8548
80 | Center | 1.07 | 1.73 | 1.59 | 1.83 | 0.96 | 1.27 | 1.38 | 0.96 | 1.29 | 1.23 | 1.08 | 1.07
| CTDIw | 1.86 | 3.01 | 2.77 | 3.18 | 1.67 | 2.21 | 2.4 | 1.67 | 2.25 | 2.14 | 1.87 | 1.86
| k | 1.7383 | 1.7399 | 1.7421 | 1.7377 | 1.7396 | 1.7402 | 1.7391 | 1.7396 | 1.7442 | 1.7398 | 1.7315 | 1.7383
100 | Center | 2.43 | 3.94 | 3.62 | 4.19 | 2.18 | 2.91 | 3.14 | 2.18 | 2.96 | 2.8 | 2.44 | 2.45
| CTDIw | 3.9 | 6.32 | 5.81 | 6.75 | 3.5 | 4.68 | 5.03 | 3.5 | 4.77 | 4.49 | 3.92 | 3.94
| k | 1.6049 | 1.6041 | 1.605 | 1.611 | 1.6055 | 1.6082 | 1.6019 | 1.6055 | 1.6115 | 1.6036 | 1.6066 | 1.6082
120 | Center | 4.28 | 6.93 | 6.37 | 7.5 | 3.84 | 5 | 5.52 | 3.84 | 5.3 | 4.93 | 4.3 | 4.37
| CTDIw | 6.61 | 10.7 | 9.84 | 11.55 | 5.93 | 8.01 | 8.53 | 5.93 | 8.16 | 7.61 | 6.64 | 6.74
| k | 1.5444 | 1.544 | 1.5447 | 1.54 | 1.5443 | 1.602 | 1.5453 | 1.5443 | 1.5396 | 1.5436 | 1.5442 | 1.5423
140 | Center | 6.57 | 10.64 | 9.78 | 11.59 | 5.89 | 8.03 | 8.47 | 5.89 | 8.19 | 7.56 | 6.6 | 6.76
| CTDIw | 9.94 | 16.09 | 14.79 | 17.62 | 8.91 | 12.22 | 12.81 | 8.91 | 12.45 | 11.43 | 9.98 | 10.28
| k | 1.5129 | 1.5122 | 1.5123 | 1.5203 | 1.5127 | 1.5218 | 1.5124 | 1.5127 | 1.5201 | 1.5119 | 1.5121 | 1.5207
100Sn | Center
CTDIw
k
120Sn | Center
CTDIw
k
140Sn | Center
CTDIw
k
32 cm phantom, Narrow shaped filter (Cardiac, Ped Body)
| Collimation | 2x1 | 6x0.6 | 8x0.6 | 8x0.6 | 1x5 | 10x0.6 | 6x1.2 | 1x10 | 20x0.6 | 12x1.2 | 32x0.6 | 64x0.6
| Total | 2 | 3.6 | 4.8 | 4.8 | 5 | 6 | 7.2 | 10 | 12 | 14.4 | 38.4 | 38.4
| Acquisition | 2x1 | 6x0.6 | 8x0.6 | 16x0.6 | 1x5 | 20x0.6 | 6x1.2 | 1x10 | 40x0.6 | 12x1.2 | 32x1.2 | 28.x.6
| Shaped filter |  |  |  | z-sharp |  | z-sharp |  |  | z-sharp |  |  | z-sharp
kV
70 | Center | 0.54 | 0.87 | 0.8 | 0.93 | 0.48 | 0.65 | 0.69 | 0.48 | 0.66 | 0.62 | 0.54 | 0.55
| CTDIw | 0.9 | 1.45 | 1.33 | 1.54 | 0.8 | 1.07 | 1.16 | 0.8 | 1.09 | 1.03 | 0.9 | 0.9
| k | 1.6667 | 1.6667 | 1.6625 | 1.6559 | 1.6667 | 1.6462 | 1.6812 | 1.6667 | 1.6515 | 1.6613 | 1.6667 | 1.6364
80 | Center | 0.94 | 1.52 | 1.39 | 1.62 | 0.84 | 1.12 | 1.21 | 0.84 | 1.15 | 1.08 | 0.94 | 0.95
| CTDIw | 1.46 | 2.36 | 2.17 | 2.48 | 1.31 | 1.72 | 1.88 | 1.31 | 1.76 | 1.68 | 1.47 | 1.45
| k | 1.5532 | 1.5526 | 1.5612 | 1.5309 | 1.5595 | 1.5357 | 1.5537 | 1.5595 | 1.5304 | 1.5556 | 1.5638 | 1.5263
100 | Center | 2.17 | 3.51 | 3.22 | 3.75 | 1.94 | 2.6 | 2.79 | 1.94 | 2.65 | 2.49 | 2.17 | 2.19
| CTDIw | 3.12 | 5.06 | 4.65 | 5.34 | 2.8 | 3.7 | 4.03 | 2.8 | 3.78 | 3.59 | 3.14 | 3.12
| k | 1.4378 | 1.4416 | 1.4441 | 1.424 | 1.4433 | 1.4231 | 1.4444 | 1.4433 | 1.4264 | 1.4418 | 1.447 | 1.4247
120 | Center | 3.84 | 6.22 | 5.72 | 6.67 | 3.45 | 4.62 | 4.95 | 3.45 | 4.71 | 4.42 | 3.86 | 3.89
| CTDIw | 5.36 | 8.68 | 7.98 | 9.21 | 4.81 | 6.39 | 6.91 | 4.81 | 6.51 | 6.17 | 5.39 | 5.38
| k | 1.3958 | 1.3955 | 1.3951 | 1.3808 | 1.3942 | 1.3831 | 1.396 | 1.3942 | 1.3822 | 1.3959 | 1.3964 | 1.383
140 | Center | 5.89 | 9.53 | 8.76 | 10.45 | 5.28 | 7.25 | 7.59 | 5.28 | 7.38 | 6.77 | 5.91 | 6.1
| CTDIw | 8.03 | 13 | 11.95 | 14.2 | 7.2 | 9.84 | 10.36 | 7.2 | 10.03 | 9.24 | 8.07 | 8.28
| k | 1.3633 | 1.3641 | 1.3642 | 1.3589 | 1.3636 | 1.3572 | 1.365 | 1.3636 | 1.3591 | 1.3648 | 1.3655 | 1.3574
100Sn | Center
CTDIw
k
120Sn | Center
CTDIw
k
140Sn | Center
CTDIw
k
16 cm phantom Standard shaped filter TwinBeam Dual Energy
| Collimation | 1x10 | 64x0.6
| Total | 10 | 38.4
kV
120 | Center | 3.68 | 4.28
| CTDIw | 3.7 | 4.29
| k | 1.0054 | 1.0023
16 cm phantom Narrow shaped filter TwinBeam Dual Energy
| Collimation | 1x10 | 64x0.6
| Total | 10 | 38.4
kV
120 | Center | 3.44 | 3.99
| CTDIw | 3.35 | 3.89
| k | 0.97384 | 0.97494
32 cm phantom Standard shaped filter TwinBeam Dual Energy
| Collimation | 1x10 | 64x0.6
| Total | 10 | 38.4
kV
120 | Center | 1.25 | 1.45
| CTDIw | 1.84 | 2.13
| k | 1.472 | 1.469
32 cm phantom Narrow shaped filter TwinBeam Dual Energy
| Collimation | 1x10 | 64x0.6
| Total | 10 | 38.4
kV
120 | Center | 1.12 | 1.3
| CTDIw | 1.5 | 1.74
| k | 1.3393 | 1.3385

## Definition AS Open 20
16 cm phantom, Standard shaped filter (Routine head)
| Collimation | 2x1 | 6x0.6 | 8x0.6 | 8x0.6 | 1x5 | 10x0.6 | 6x1.2 | 1x10 | 20x0.6 | 12x1.2 | 32x0.6 | 64x0.6
| Total | 2 | 3.6 | 4.8 | 4.8 | 5 | 6 | 7.2 | 10 | 12 | 14.4 | 38.4 | 38.4
| Acquisition | 2x1 | 6x0.6 | 8x0.6 | 16x0.6 | 1x5 | 20x0.6 | 6x1.2 | 1x10 | 40x0.6 | 12x1.2 | 32x1.2 | 128x0.6
|  |  | Topogram | UHR | UHR, z-sharp |  | z-sharp |  |  | z-sharp |  |  | z-sharp
kV
70 | Center | 2.6 | 4.22 | 3.88 | 4.58 | 2.34 | 3.17 | 3.3 | 2.34 | 3.23 | 3 | 3.01 | 2.87
| CTDIw | 2.74 | 4.43 | 4.07 | 4.78 | 2.46 | 3.31 | 3.35 | 2.46 | 3.38 | 3.15 | 3.15 | 3.01
| k | 1.0538 | 1.0498 | 1.049 | 1.0437 | 1.0513 | 1.0442 | 1.0152 | 1.0513 | 1.0464 | 1.05 | 1.0465 | 1.0488
80 | Center | 4.18 | 6.76 | 6.22 | 7.24 | 3.75 | 5.02 | 5.39 | 3.75 | 5.12 | 4.81 | 4.77 | 4.6
| CTDIw | 4.36 | 7.06 | 6.49 | 7.46 | 3.91 | 5.17 | 5.62 | 3.91 | 5.27 | 5.02 | 4.91 | 4.8
| k | 1.0431 | 1.0444 | 1.0434 | 1.0304 | 1.0427 | 1.0299 | 1.0427 | 1.0427 | 1.0293 | 1.0437 | 1.0294 | 1.0435
100 | Center | 8.44 | 13.67 | 12.56 | 14.69 | 7.57 | 10.19 | 10.89 | 7.57 | 10.38 | 9.71 | 8.48 | 8.57
| CTDIw | 8.65 | 14.01 | 12.88 | 14.93 | 7.76 | 10.35 | 11.16 | 7.76 | 10.55 | 9.96 | 8.69 | 8.71
| k | 1.0249 | 1.0249 | 1.0255 | 1.0163 | 1.0251 | 1.0157 | 1.0248 | 1.0251 | 1.0164 | 1.0257 | 1.0248 | 1.0163
120 | Center | 13.85 | 22.42 | 20.6 | 24.33 | 12.42 | 16.87 | 17.86 | 12.42 | 17.19 | 15.93 | 13.91 | 14.19
| CTDIw | 14.11 | 22.85 | 21 | 24.63 | 12.66 | 17.08 | 18.2 | 12.66 | 17.41 | 16.24 | 14.17 | 14.37
| k | 1.0188 | 1.0192 | 1.0194 | 1.0123 | 1.0193 | 1.0124 | 1.019 | 1.0193 | 1.0128 | 1.0195 | 1.0187 | 1.0127
140 | Center | 20.24 | 32.77 | 30.12 | 36.16 | 18.16 | 25.07 | 26.11 | 18.16 | 25.55 | 23.3 | 20.33 | 21.09
| CTDIw | 20.6 | 33.35 | 30.65 | 36.6 | 18.48 | 25.4 | 26.56 | 18.48 | 25.89 | 23.7 | 20.69 | 21.37
| k | 1.0178 | 1.0177 | 1.0176 | 1.0122 | 1.0176 | 1.0132 | 1.0172 | 1.0176 | 1.0133 | 1.0172 | 1.0177 | 1.0133
100Sn | Center
CTDIw
k
120Sn | Center
CTDIw
k
140Sn | Center
CTDIw
k
16 cm phantom, Narrow shaped filter (Pediatric body)
| Collimation | 2x1 | 6x0.6 | 8x0.6 | 8x0.6 | 1x5 | 10x0.6 | 6x1.2 | 1x10 | 20x0.6 | 12x1.2 | 32x0.6 | 64x0.6
| Total | 2 | 3.6 | 4.8 | 4.8 | 5 | 6 | 7.2 | 10 | 12 | 14.4 | 38.4 | 38.4
| Acquisition | 2x1 | 6x0.6 | 8x0.6 | 16x0.6 | 1x5 | 20x0.6 | 6x1.2 | 1x10 | 40x0.6 | 12x1.2 | 32x1.2 | 28.x.6
| Shaped filter |  |  |  | z-sharp |  | z-sharp |  |  | z-sharp |  |  | z-sharp
kV
70 | Center | 2.36 | 3.83 | 3.52 | 4.12 | 2.12 | 2.86 | 3.05 | 2.12 | 2.91 | 2.72 | 2.38 | 2.4
| CTDIw | 2.36 | 3.81 | 3.51 | 4.07 | 2.11 | 2.82 | 3.04 | 2.11 | 2.88 | 2.71 | 2.37 | 2.37
| k | 1 | 0.99478 | 0.99716 | 0.98786 | 0.99528 | 0.98601 | 0.99672 | 0.99528 | 0.98969 | 0.99632 | 0.9958 | 0.9875
80 | Center | 3.81 | 6.17 | 5.67 | 6.58 | 3.42 | 4.56 | 4.92 | 3.42 | 4.65 | 4.39 | 3.83 | 3.84
| CTDIw | 3.78 | 6.12 | 5.63 | 6.44 | 3.39 | 4.47 | 4.88 | 3.39 | 4.55 | 4.35 | 3.8 | 3.76
| k | 0.99213 | 0.9919 | 0.99295 | 0.97872 | 0.99123 | 0.98026 | 0.99187 | 0.99123 | 0.97849 | 0.99089 | 0.99217 | 0.97917
100 | Center | 7.81 | 12.65 | 11.63 | 13.53 | 7.01 | 9.38 | 10.08 | 7.01 | 9.56 | 8.99 | 7.85 | 7.89
| CTDIw | 7.69 | 12.45 | 11.44 | 13.14 | 6.9 | 9.11 | 9.92 | 6.9 | 9.29 | 8.85 | 7.72 | 7.67
| k | 0.98464 | 0.98419 | 0.98366 | 0.97118 | 0.98431 | 0.97122 | 0.98413 | 0.98431 | 0.97176 | 0.98443 | 0.98344 | 0.97212
120 | Center | 12.93 | 20.93 | 19.23 | 22.53 | 11.6 | 15.62 | 16.67 | 11.6 | 15.92 | 14.87 | 12.98 | 13.14
| CTDIw | 12.68 | 20.52 | 18.86 | 21.91 | 11.37 | 15.19 | 16.35 | 11.37 | 15.48 | 14.59 | 12.73 | 12.78
| k | 0.98067 | 0.98041 | 0.98076 | 0.97248 | 0.98017 | 0.97247 | 0.9808 | 0.98017 | 0.97236 | 0.98117 | 0.98074 | 0.9726
140 | Center | 18.83 | 30.49 | 28.02 | 33.71 | 16.89 | 23.38 | 24.29 | 16.89 | 23.82 | 21.67 | 18.91 | 19.67
| CTDIw | 18.5 | 29.95 | 27.53 | 32.85 | 16.6 | 22.77 | 23.86 | 16.6 | 23.21 | 21.29 | 18.58 | 19.16
| k | 0.98247 | 0.98229 | 0.98251 | 0.97449 | 0.98283 | 0.97391 | 0.9823 | 0.98283 | 0.97439 | 0.98246 | 0.98255 | 0.97407
100Sn | Center
CTDIw
k
120Sn | Center
CTDIw
k
140Sn | Center
CTDIw
k
32 cm phantom, Standard shaped filter (Adult Body)
| Collimation | 2x1 | 6x0.6 | 8x0.6 | 8x0.6 | 1x5 | 10x0.6 | 6x1.2 | 1x10 | 20x0.6 | 12x1.2 | 32x0.6 | 64x0.6
| Total | 2 | 3.6 | 4.8 | 4.8 | 5 | 6 | 7.2 | 10 | 12 | 14.4 | 38.4 | 38.4
| Acquisition | 2x1 | 6x0.6 | 8x0.6 | 16x0.6 | 1x5 | 20x0.6 | 6x1.2 | 1x10 | 40x0.6 | 12x1.2 | 32x1.2 | 28.x.6
| Shaped filter |  |  |  | z-sharp |  | z-sharp |  |  | z-sharp |  |  | z-sharp
kV
70 | Center | 0.61 | 0.99 | 0.91 | 1.07 | 0.55 | 0.74 | 0.79 | 0.55 | 0.75 | 0.7 | 0.61 | 0.62
| CTDIw | 1.13 | 1.83 | 1.68 | 1.97 | 1.01 | 1.37 | 1.46 | 1.01 | 1.4 | 1.3 | 1.14 | 1.15
| k | 1.8525 | 1.8485 | 1.8462 | 1.8411 | 1.8364 | 1.8514 | 1.8481 | 1.8364 | 1.8667 | 1.8571 | 1.8689 | 1.8548
80 | Center | 1.07 | 1.73 | 1.59 | 1.83 | 0.96 | 1.27 | 1.38 | 0.96 | 1.29 | 1.23 | 1.08 | 1.07
| CTDIw | 1.86 | 3.01 | 2.77 | 3.18 | 1.67 | 2.21 | 2.4 | 1.67 | 2.25 | 2.14 | 1.87 | 1.86
| k | 1.7383 | 1.7399 | 1.7421 | 1.7377 | 1.7396 | 1.7402 | 1.7391 | 1.7396 | 1.7442 | 1.7398 | 1.7315 | 1.7383
100 | Center | 2.43 | 3.94 | 3.62 | 4.19 | 2.18 | 2.91 | 3.14 | 2.18 | 2.96 | 2.8 | 2.44 | 2.45
| CTDIw | 3.9 | 6.32 | 5.81 | 6.75 | 3.5 | 4.68 | 5.03 | 3.5 | 4.77 | 4.49 | 3.92 | 3.94
| k | 1.6049 | 1.6041 | 1.605 | 1.611 | 1.6055 | 1.6082 | 1.6019 | 1.6055 | 1.6115 | 1.6036 | 1.6066 | 1.6082
120 | Center | 4.28 | 6.93 | 6.37 | 7.5 | 3.84 | 5 | 5.52 | 3.84 | 5.3 | 4.93 | 4.3 | 4.37
| CTDIw | 6.61 | 10.7 | 9.84 | 11.55 | 5.93 | 8.01 | 8.53 | 5.93 | 8.16 | 7.61 | 6.64 | 6.74
| k | 1.5444 | 1.544 | 1.5447 | 1.54 | 1.5443 | 1.602 | 1.5453 | 1.5443 | 1.5396 | 1.5436 | 1.5442 | 1.5423
140 | Center | 6.57 | 10.64 | 9.78 | 11.59 | 5.89 | 8.03 | 8.47 | 5.89 | 8.19 | 7.56 | 6.6 | 6.76
| CTDIw | 9.94 | 16.09 | 14.79 | 17.62 | 8.91 | 12.22 | 12.81 | 8.91 | 12.45 | 11.43 | 9.98 | 10.28
| k | 1.5129 | 1.5122 | 1.5123 | 1.5203 | 1.5127 | 1.5218 | 1.5124 | 1.5127 | 1.5201 | 1.5119 | 1.5121 | 1.5207
100Sn | Center
CTDIw
k
120Sn | Center
CTDIw
k
140Sn | Center
CTDIw
k
32 cm phantom, Narrow shaped filter (Cardiac, Ped Body)
| Collimation | 2x1 | 6x0.6 | 8x0.6 | 8x0.6 | 1x5 | 10x0.6 | 6x1.2 | 1x10 | 20x0.6 | 12x1.2 | 32x0.6 | 64x0.6
| Total | 2 | 3.6 | 4.8 | 4.8 | 5 | 6 | 7.2 | 10 | 12 | 14.4 | 38.4 | 38.4
| Acquisition | 2x1 | 6x0.6 | 8x0.6 | 16x0.6 | 1x5 | 20x0.6 | 6x1.2 | 1x10 | 40x0.6 | 12x1.2 | 32x1.2 | 28.x.6
| Shaped filter |  |  |  | z-sharp |  | z-sharp |  |  | z-sharp |  |  | z-sharp
kV
70 | Center | 0.54 | 0.87 | 0.8 | 0.93 | 0.48 | 0.65 | 0.69 | 0.48 | 0.66 | 0.62 | 0.54 | 0.55
| CTDIw | 0.9 | 1.45 | 1.33 | 1.54 | 0.8 | 1.07 | 1.16 | 0.8 | 1.09 | 1.03 | 0.9 | 0.9
| k | 1.6667 | 1.6667 | 1.6625 | 1.6559 | 1.6667 | 1.6462 | 1.6812 | 1.6667 | 1.6515 | 1.6613 | 1.6667 | 1.6364
80 | Center | 0.94 | 1.52 | 1.39 | 1.62 | 0.84 | 1.12 | 1.21 | 0.84 | 1.15 | 1.08 | 0.94 | 0.95
| CTDIw | 1.46 | 2.36 | 2.17 | 2.48 | 1.31 | 1.72 | 1.88 | 1.31 | 1.76 | 1.68 | 1.47 | 1.45
| k | 1.5532 | 1.5526 | 1.5612 | 1.5309 | 1.5595 | 1.5357 | 1.5537 | 1.5595 | 1.5304 | 1.5556 | 1.5638 | 1.5263
100 | Center | 2.17 | 3.51 | 3.22 | 3.75 | 1.94 | 2.6 | 2.79 | 1.94 | 2.65 | 2.49 | 2.17 | 2.19
| CTDIw | 3.12 | 5.06 | 4.65 | 5.34 | 2.8 | 3.7 | 4.03 | 2.8 | 3.78 | 3.59 | 3.14 | 3.12
| k | 1.4378 | 1.4416 | 1.4441 | 1.424 | 1.4433 | 1.4231 | 1.4444 | 1.4433 | 1.4264 | 1.4418 | 1.447 | 1.4247
120 | Center | 3.84 | 6.22 | 5.72 | 6.67 | 3.45 | 4.62 | 4.95 | 3.45 | 4.71 | 4.42 | 3.86 | 3.89
| CTDIw | 5.36 | 8.68 | 7.98 | 9.21 | 4.81 | 6.39 | 6.91 | 4.81 | 6.51 | 6.17 | 5.39 | 5.38
| k | 1.3958 | 1.3955 | 1.3951 | 1.3808 | 1.3942 | 1.3831 | 1.396 | 1.3942 | 1.3822 | 1.3959 | 1.3964 | 1.383
140 | Center | 5.89 | 9.53 | 8.76 | 10.45 | 5.28 | 7.25 | 7.59 | 5.28 | 7.38 | 6.77 | 5.91 | 6.1
| CTDIw | 8.03 | 13 | 11.95 | 14.2 | 7.2 | 9.84 | 10.36 | 7.2 | 10.03 | 9.24 | 8.07 | 8.28
| k | 1.3633 | 1.3641 | 1.3642 | 1.3589 | 1.3636 | 1.3572 | 1.365 | 1.3636 | 1.3591 | 1.3648 | 1.3655 | 1.3574
100Sn | Center
CTDIw
k
120Sn | Center
CTDIw
k
140Sn | Center
CTDIw
k
16 cm phantom Standard shaped filter TwinBeam Dual Energy
| Collimation | 1x10 | 64x0.6
| Total | 10 | 38.4
kV
120 | Center | 3.68 | 4.28
| CTDIw | 3.7 | 4.29
| k | 1.0054 | 1.0023
16 cm phantom Narrow shaped filter TwinBeam Dual Energy
| Collimation | 1x10 | 64x0.6
| Total | 10 | 38.4
kV
120 | Center | 3.44 | 3.99
| CTDIw | 3.35 | 3.89
| k | 0.97384 | 0.97494
32 cm phantom Standard shaped filter TwinBeam Dual Energy
| Collimation | 1x10 | 64x0.6
| Total | 10 | 38.4
kV
120 | Center | 1.25 | 1.45
| CTDIw | 1.84 | 2.13
| k | 1.472 | 1.469
32 cm phantom Narrow shaped filter TwinBeam Dual Energy
| Collimation | 1x10 | 64x0.6
| Total | 10 | 38.4
kV
120 | Center | 1.12 | 1.3
| CTDIw | 1.5 | 1.74
| k | 1.3393 | 1.3385

## Naeotom Alpha
16 cm phantom, Standard shaped filter (Routine head)
|  | Collimation | 1x5 | 36x0.2 | 3x3 | 3x5 | 36x0.4 | 48x0.4 | 96x0.2 | 120x0.2 | 72x0.4 | 96x0.4 | 144x0.4
|  | Total | 5 | 7.2 | 9 | 15 | 14.4 | 19.2 | 19.2 | 24 | 28.8 | 38.4 | 57.6
|  |  | 6.5 | 7.5 | 14.2 | 20.2 |  | 29 | 23.3 | 29.2 |  | 49 | 69
Spectral filter |  | Acquisition | 1x5 | 36x0.2 | 3x3 | 3x5 | 36x0.4 | 48x0.4 | 96x0.2 | 120x0.2 | 72x0.4 | 96x0.4 | 144x0.4
kV
None | 70 | Center | 3.21 | 2.69 | 4.08 | 3.88 |  | 4.09 | 3.64 | 3.56 | 3.93 | 3.72 | 3.51
|  | CTDIw | 3.45 | 2.9 | 4.39 | 4.17 |  | 4.4 | 3.92 | 3.83 | 4.23 | 4 | 3.77
|  | k | 1.0748 | 1.0781 | 1.076 | 1.0747 |  | 1.0758 | 1.0769 | 1.0758 | 1.0763 | 1.0753 | 1.0741
| 90 | Center | 7.13 | 5.99 | 9.08 | 8.63 | 9.85 | 9.11 | 8.1 | 7.92 | 8.73 | 8.27 | 7.8
|  | CTDIw | 7.44 | 6.25 | 9.48 | 9 | 10.28 | 9.5 | 8.45 | 8.26 | 9.12 | 8.63 | 8.14
|  | k | 1.0435 | 1.0434 | 1.0441 | 1.0429 | 1.0437 | 1.0428 | 1.0432 | 1.0429 | 1.0447 | 1.0435 | 1.0436
| 120 | Center | 15.04 | 12.64 | 19.16 | 18.2 | 20.79 | 19.21 | 17.09 | 16.7 | 18.42 | 17.43 | 16.45
|  | CTDIw | 15.46 | 12.99 | 19.69 | 18.7 | 21.36 | 19.74 | 17.56 | 17.17 | 18.94 | 17.92 | 16.9
|  | k | 1.0279 | 1.0277 | 1.0277 | 1.0275 | 1.0274 | 1.0276 | 1.0275 | 1.0281 | 1.0282 | 1.0281 | 1.0274
| 140 | Center | 21.41 | 17.99 | 27.27 | 25.91 | 29.59 | 27.43 | 24.33 | 23.78 | 26.23 | 24.82 | 23.41
|  | CTDIw | 21.88 | 18.39 | 27.87 | 26.47 | 30.24 | 27.94 | 24.86 | 24.3 | 26.8 | 25.36 | 23.92
|  | k | 1.022 | 1.0222 | 1.022 | 1.0216 | 1.022 | 1.0186 | 1.0218 | 1.0219 | 1.0217 | 1.0218 | 1.0218
Sn LD | 100 | Center | 1.54 | 1.29 | 1.96 | 1.86 | 2.12 | 1.96 | 1.75 | 1.71 | 1.88 | 1.78 | 1.68
|  | CTDIw | 1.54 | 1.29 | 1.96 | 1.86 | 2.14 | 1.96 | 1.75 | 1.71 | 1.88 | 1.78 | 1.68
|  | k | 1 | 1 | 1 | 1 | 1.0094 | 1 | 1 | 1 | 1 | 1 | 1
| 140 | Center | 6.03 | 5.07 | 7.68 | 7.3 | 8.34 | 7.7 | 6.85 | 6.7 | 7.39 | 6.99 | 6.59
|  | CTDIw | 6.09 | 5.12 | 7.76 | 7.37 | 8.42 | 7.78 | 6.92 | 6.77 | 7.46 | 7.06 | 6.66
|  | k | 1.01 | 1.0099 | 1.0104 | 1.0096 | 1.0096 | 1.0104 | 1.0102 | 1.0104 | 1.0095 | 1.01 | 1.0106
150 | 150 | Center | [6 empty] | 5.09
|  | CTDIw | [6 empty] | 5.18
|  | k | [6 empty] | 1.0177
16 cm phantom, Narrow shaped filter (Pediatric body)
|  | Collimation | 1x5 | 36x0.2 | 3x3 | 3x5 | 36x0.4 | 48x0.4 | 96x0.2 | 120x0.2 | 72x0.4 | 96x0.4 | 144x0.4
|  | Total | 5 | 7.2 | 9 | 15 | 14.4 | 19.2 | 19.2 | 24 | 28.8 | 38.4 | 57.6
|  |  | 6.5 | 7.5 | 14.2 | 20.2 |  | 29 | 23.3 | 29.2 |  | 49 | 69
Spectral filter |  | Acquisition | 1x5 | 36x0.2 | 3x3 | 3x5 | 36x0.4 | 48x0.4 | 96x0.2 | 120x0.2 | 72x0.4 | 96x0.4 | 144x0.4
kV
None | 70 | Center | 2.85 | 2.4 | 3.64 | 3.64 |  | 3.64 | 3.64 | 3.17 | 3.5 | 3.31 | 3.12
|  | CTDIw | 2.89 | 2.43 | 3.68 | 3.69 |  | 3.69 | 3.69 | 3.21 | 3.54 | 3.35 | 3.16
|  | k | 1.014 | 1.0125 | 1.011 | 1.0137 |  | 1.0137 | 1.0137 | 1.0126 | 1.0114 | 1.0121 | 1.0128
| 90 | Center | 6.46 | 5.43 | 8.23 | 8.26 | 8.93 | 8.26 | 8.26 | 7.18 | 7.92 | 7.49 | 7.07
|  | CTDIw | 6.41 | 5.38 | 8.16 | 8.18 | 8.85 | 8.18 | 8.18 | 7.11 | 7.85 | 7.43 | 7.01
|  | k | 0.99226 | 0.99079 | 0.99149 | 0.99031 | 0.99104 | 0.99031 | 0.99031 | 0.99025 | 0.99116 | 0.99199 | 0.99151
| 120 | Center | 13.85 | 11.64 | 17.65 | 17.7 | 19.15 | 17.7 | 17.7 | 15.39 | 16.98 | 16.06 | 15.15
|  | CTDIw | 13.62 | 11.45 | 17.35 | 17.4 | 18.83 | 17.4 | 17.4 | 15.13 | 16.69 | 15.8 | 14.9
|  | k | 0.98339 | 0.98368 | 0.983 | 0.98305 | 0.98329 | 0.98305 | 0.98305 | 0.98311 | 0.98292 | 0.98381 | 0.9835
| 140 | Center | 19.8 | 16.64 | 25.22 | 25.29 | 27.37 | 25.29 | 25.29 | 21.99 | 24.26 | 22.95 | 21.65
|  | CTDIw | 19.45 | 16.34 | 24.77 | 24.84 | 26.88 | 24.84 | 24.84 | 21.6 | 23.83 | 22.55 | 21.27
|  | k | 0.98232 | 0.98197 | 0.98216 | 0.98221 | 0.9821 | 0.98221 | 0.98221 | 0.98226 | 0.98228 | 0.98257 | 0.98245
Sn LD | 100 | Center | 1.44 | 1.21 | 1.84 | 1.84 | 1.99 | 1.84 | 1.84 | 1.6 | 1.77 | 1.67 | 1.58
|  | CTDIw | 1.4 | 1.18 | 1.78 | 1.79 | 1.93 | 1.79 | 1.79 | 1.55 | 1.71 | 1.62 | 1.53
|  | k | 0.97222 | 0.97521 | 0.96739 | 0.97283 | 0.96985 | 0.97283 | 0.97283 | 0.96875 | 0.9661 | 0.97006 | 0.96835
| 140 | Center | 5.72 | 4.8 | 7.28 | 7.3 | 7.9 | 7.3 | 7.3 | 6.35 | 7 | 6.63 | 6.25
|  | CTDIw | 5.62 | 4.72 | 7.16 | 7.18 | 7.77 | 7.18 | 7.18 | 6.24 | 6.89 | 6.52 | 6.15
|  | k | 0.98252 | 0.98333 | 0.98352 | 0.98356 | 0.98354 | 0.98356 | 0.98356 | 0.98268 | 0.98429 | 0.98341 | 0.984
Sn DE | 150 | Center | [6 empty] | 4.85
|  | CTDIw | [6 empty] | 4.81
|  | k | [6 empty] | 0.99175
32 cm phantom, Standard shaped filter (Adult body)
|  | Collimation | 1x5 | 36x0.2 | 3x3 | 3x5 | 36x0.4 | 48x0.4 | 96x0.2 | 120x0.2 | 72x0.4 | 96x0.4 | 144x0.4
|  | Total | 5 | 7.2 | 9 | 15 | 14.4 | 19.2 | 19.2 | 24 | 28.8 | 38.4 | 57.6
|  |  | 6.5 | 7.5 | 14.2 | 20.2 |  | 29 | 23.3 | 29.2 |  | 49 | 69
Spectral filter |  | Acquisition | 1x5 | 36x0.2 | 3x3 | 3x5 | 36x0.4 | 48x0.4 | 96x0.2 | 120x0.2 | 72x0.4 | 96x0.4 | 144x0.4
kV
None | 70 | Center | 0.71 | 0.59 | 0.9 | 0.86 |  | 0.9 | 0.9 | 0.79 | 0.87 | 0.82 | 0.77
|  | CTDIw | 1.41 | 1.18 | 1.8 | 1.71 |  | 1.8 | 1.8 | 1.57 | 1.73 | 1.63 | 1.54
|  | k | 1.9859 | 2 | 2 | 1.9884 |  | 2 | 2 | 1.9873 | 1.9885 | 1.9878 | 2
| 90 | Center | 1.86 | 1.57 | 2.37 | 2.26 | 2.58 | 2.38 | 2.38 | 2.07 | 2.28 | 2.16 | 2.04
|  | CTDIw | 3.26 | 2.74 | 4.16 | 3.95 | 4.51 | 4.17 | 4.17 | 3.62 | 4 | 3.78 | 3.57
|  | k | 1.7527 | 1.7452 | 1.7553 | 1.7478 | 1.7481 | 1.7521 | 1.7521 | 1.7488 | 1.7544 | 1.75 | 1.75
| 120 | Center | 4.47 | 3.76 | 5.7 | 4.51 | 6.18 | 5.71 | 5.71 | 4.97 | 5.48 | 5.18 | 4.89
|  | CTDIw | 7.23 | 6.07 | 9.2 | 8.74 | 9.99 | 9.23 | 9.23 | 8.03 | 8.85 | 8.38 | 7.9
|  | k | 1.6174 | 1.6144 | 1.614 | 1.9379 | 1.6165 | 1.6165 | 1.6165 | 1.6157 | 1.615 | 1.6178 | 1.6155
| 140 | Center | 6.69 | 5.63 | 8.53 | 8.1 | 9.25 | 8.55 | 8.55 | 7.44 | 8.2 | 7.76 | 7.32
|  | CTDIw | 10.51 | 8.84 | 13.39 | 12.73 | 14.53 | 13.43 | 13.43 | 11.68 | 12.88 | 12.19 | 11.5
|  | k | 1.571 | 1.5702 | 1.5698 | 1.5716 | 1.5708 | 1.5708 | 1.5708 | 1.5699 | 1.5707 | 1.5709 | 1.571
Sn LD | 100 | Center | 0.53 | 0.44 | 0.67 | 0.64 | 0.73 | 0.67 | 0.67 | 0.59 | 0.65 | 0.61 | 0.58
|  | CTDIw | 0.78 | 0.66 | 0.99 | 0.94 | 1.08 | 1 | 1 | 0.87 | 0.96 | 0.9 | 0.85
|  | k | 1.4717 | 1.5 | 1.4776 | 1.4687 | 1.4795 | 1.4925 | 1.4925 | 1.4746 | 1.4769 | 1.4754 | 1.4655
| 140 | Center | 2.26 | 1.9 | 2.88 | 2.74 | 3.12 | 2.89 | 2.89 | 2.51 | 2.77 | 2.62 | 2.47
|  | CTDIw | 3.26 | 2.74 | 4.15 | 3.94 | 4.5 | 4.16 | 4.16 | 3.62 | 3.99 | 3.78 | 3.56
|  | k | 1.4425 | 1.4421 | 1.441 | 1.438 | 1.4423 | 1.4394 | 1.4394 | 1.4422 | 1.4404 | 1.4427 | 1.4413
Sn DE | 150 | Center | [6 empty] | 2.02
|  | CTDIw | [6 empty] | 2.87
|  | k | [6 empty] | 1.4208
32 cm phantom, Narrow shaped filter (Cardiac)
|  | Collimation | 1x5 | 36x0.2 | 3x3 | 3x5 | 36x0.4 | 48x0.4 | 96x0.2 | 120x0.2 | 72x0.4 | 96x0.4 | 144x0.4
|  | Total | 5 | 7.2 | 9 | 15 | 14.4 | 19.2 | 19.2 | 24 | 28.8 | 38.4 | 57.6
|  |  | 6.5 | 7.5 | 14.2 | 20.2 |  | 29 | 23.3 | 29.2 |  | 49 | 69
Spectral filter |  | Acquisition | 1x5 | 36x0.2 | 3x3 | 3x5 | 36x0.4 | 48x0.4 | 96x0.2 | 120x0.2 | 72x0.4 | 96x0.4 | 144x0.4
kV
None | 70 | Center | 0.61 | 0.51 | 0.78 | 0.78 |  | 0.78 | 0.78 | 0.68 | 0.75 | 0.71 | 0.67
|  | CTDIw | 1.07 | 0.9 | 1.37 | 1.37 |  | 1.37 | 1.37 | 1.19 | 1.32 | 1.25 | 1.18
|  | k | 1.7541 | 1.7647 | 1.7564 | 1.7564 |  | 1.7564 | 1.7564 | 1.75 | 1.76 | 1.7606 | 1.7612
| 90 | Center | 1.63 | 1.37 | 2.08 | 2.08 | 2.26 | 2.08 | 2.08 | 1.81 | 2 | 1.89 | 1.79
|  | CTDIw | 2.55 | 2.14 | 3.24 | 3.25 | 3.52 | 3.25 | 3.25 | 2.83 | 3.12 | 2.95 | 2.78
|  | k | 1.5644 | 1.562 | 1.5577 | 1.5625 | 1.5575 | 1.5625 | 1.5625 | 1.5635 | 1.56 | 1.5608 | 1.5531
| 120 | Center | 3.97 | 3.34 | 5.06 | 5.07 | 5.49 | 5.07 | 5.07 | 4.41 | 4.86 | 4.6 | 4.34
|  | CTDIw | 5.75 | 4.84 | 7.33 | 7.35 | 7.95 | 7.35 | 7.35 | 6.39 | 7.05 | 6.67 | 6.29
|  | k | 1.4484 | 1.4491 | 1.4486 | 1.4497 | 1.4481 | 1.4497 | 1.4497 | 1.449 | 1.4506 | 1.45 | 1.4493
| 140 | Center | 5.97 | 5.02 | 7.61 | 7.63 | 8.25 | 7.63 | 7.63 | 6.63 | 7.31 | 6.92 | 6.53
|  | CTDIw | 8.44 | 7.1 | 10.76 | 10.78 | 11.67 | 10.78 | 10.78 | 9.38 | 10.34 | 7.97 | 9.23
|  | k | 1.4137 | 1.4143 | 1.4139 | 1.4128 | 1.4145 | 1.4128 | 1.4128 | 1.4148 | 1.4145 | 1.1517 | 1.4135
Sn LD | 100 | Center | 0.47 | 0.4 | 0.6 | 0.6 | 0.65 | 0.6 | 0.6 | 0.53 | 0.58 | 0.55 | 0.52
|  | CTDIw | 0.64 | 0.53 | 0.81 | 0.81 | 0.88 | 0.81 | 0.81 | 0.71 | 0.78 | 0.74 | 0.7
|  | k | 1.3617 | 1.325 | 1.35 | 1.35 | 1.3538 | 1.35 | 1.35 | 1.3396 | 1.3448 | 1.3455 | 1.3462
| 140 | Center | 2.06 | 1.73 | 2.62 | 2.63 | 2.85 | 2.63 | 2.63 | 2.29 | 2.52 | 2.39 | 2.25
|  | CTDIw | 2.71 | 2.28 | 3.45 | 3.46 | 3.75 | 3.46 | 3.46 | 3.01 | 3.32 | 3.14 | 2.96
|  | k | 1.3155 | 1.3179 | 1.3168 | 1.3156 | 1.3158 | 1.3156 | 1.3156 | 1.3144 | 1.3175 | 1.3138 | 1.3156
Sn DE | 150 | Center | [6 empty] | 1.86
|  | CTDIw | [6 empty] | 2.41
|  | k | [6 empty] | 1.2957